Satlantic MBARI ISUS
1. Device initialization performed by instrument service

Sets to TRIGGERED mode. Sets sampling duration. Sets frame format to FULL ASCII. Turns off status messages.
2. Device metadata returned in SIAM datastream
show

Configuration parameters:

Instrument type:    ISUS

Serial No.:           90

Baudrate:          38400

Deployment Cntr:      21

disk

Persistor CF5578

BIOS 3.0 ((C) JHG/PII, Apr 26 2004)

PicoDOS 3.0 - 1048576 bytes of RAM

Flash Disk Card: Size=256499712 Free=149204992
(NOTE: does not always appear in datastream – this is a bug)

3. Instrument-specific service properties
/** Frames per sample; note that every 11th frame is dark. */

int sampleFrames = 4;

4. Initial service property values for this deployment
powerPolicy=ALWAYS 

commPowerPolicy=SAMPLING 

powerOnDelaySec=0 

sampleSchedule=R */0 */0 */1 */0 * * * * GMT * 

nominalCurrentMa=1000 

currentLimitMa=8000 

sampleFrames=4 

UUID=5ff18e39-a3f2-11da-bb39-00306e389969 

timeSynch=false 

maxSampleTries=1 

isiID=1514 

sampleTimeoutMsec=5000 

serviceStatus=2 

startDelayMsec=20000

serviceName=Satlantic MBARI ISUS s/n 090 

instrumentName=UNKNOWN 

className=moos.devices.isus2.ISUS2$Attributes 

parentID=1481 

extendedVersion=UNK 

frameworkVersion=$Name: march13_2006-2 $ 

5. Results of acceptance test procedures
OK
6. Known-service bugs/deficiencies
Sampling errors from listPorts < 5%
Device metadata not appearing in SIAM datastream

P.I. desires to ensure data collection in event of mooring controller failure – should use ISUS “scheduled” mode







