Satlantic HyperOCR 
(Secondary Node(1481), Port 0, ISI-ID = 1517
1. Instrument device settings initialized by SIAM service. 
None.
2. Device metadata returned in SIAM datastream
parentID=1481, recordType=0

devid=1517, t=1142565883061, seqNo=0, mdref=0

 cause=just checking

MASTER NODE:

ID:

HyperOCR - Hyperspectral Ocean Colour Radiometer

Copyright (C) 2004, Satlantic Inc. All rights reserved.

Firmware version: 1.2B(P), 1.0A(S) - SatNet Type A

Instrument: SATHSE

S/N: 0239
STATE:

show all

Telemetry Baud Rate: 57600 bps

Maximum Frame Rate: AUTO
Initialize Silent Mode: off
Initialize Power Down: off
Initialize Automatic Telemetry: on
Network Mode: off
Network Address: 150
Network Baud Rate: 38400 bps
Network Master Mode: off
Master Controlled Telemetry: off
Master Network Bias: on
Network Reset Delay: 5
Minimum Integration Time: 8 ms
Maximum Integration Time: 2048 ms
Starting Integration Time: 128 ms
Increase Factor: 2.000000
Decrease Factor: 0.500000
Adaptive Gain: on
Upper Threshold: 42000
Lower Threshold: 10000
Dark Frames: 5

3. Instrument-specific service properties
sampleTime=30000 **(amt of time data collected from OCR-3000)
4. Initial service property values for this deployment
powerOnDelaySec=0

instrumentName=UNKNOWN

nominalCurrentMa=1000

className=moos.devices.ocr3000.Ocr3000$Attributes

serviceName=Satlantic HyperOCR s/n 239

parentID=1481

currentLimitMa=8000

sampleSchedule=A */0 */10 */0 */0 * * * * GMT *

extendedVersion=UNK

frameworkVersion=$Name: march14_2006 $

UUID=bd01ceec-a3f8-11da-bb39-00306e389969

powerPolicy=SAMPLING

timeSynch=false

maxSampleTries=3

isiID=1517

sampleTimeoutMsec=1000

commPowerPolicy=SAMPLING

serviceStatus=0

startDelayMsec=10000

5. Results of acceptance test procedures
All Okay.
6. Known-service bugs/deficiencies
unable to access XML file







