WHOI ASIMET LWR
(Secondary Node(1481), Port 4, ISI-ID = 1289)
1. Manually-configured instrument settings not handled by service
None.

2. Device metadata returned in SIAM datastream
parentID=1481, recordType=0

devid=1289, t=1142566263963, seqNo=0, mdref=0

 cause=just checking

Requesting cal constants - OK

LWR01
LWR308

VOSLWR53 v3.5
2.4576 Mhz
02AUG05
2006/03/17  03:36:16
Set1: -5.74036e+05  1.75140e+01  0.00000e+00  0.00000e+00
Set2:  1.02652e-03  2.39674e-04  1.54012e-07  0.00000e+00
Set3: -5.74083e+05  1.75154e+01  0.00000e+00  0.00000e+00
Set4:  1.03175e-03  2.38963e-04  1.56149e-07  0.00000e+00
Set5: -2.00893e+04  6.10967e-01  0.00000e+00  0.00000e+00
Set6:  3.93300e+02  2.66000e+00  0.00000e+00  0.00000e+00
No PCMCIA card installed
MODADR: LWR01
MODMFG: STARENG.
MODMOD: VOSLWR
MODSER: LWR308
MODDAT: 24OCT03
SENMFG: EPPLEY LAB
SENMOD: PIR
SENSER: 33681F3
SENDAT: -
SFTMFG: WHOI/GEOFF
SFTNAM: VOSLWRF
SFTREV: 1.6
SFTDAT: 26APR04
CALFAC: WHOI
CALPER: J.C.SMITH
CALDAT: 02AUG05
DATFRM: -
DATDES: -
DATUNI: -
RAWFRM: -
RAWDES: -

RAWUNI: -

3. Instrument-specific service properties
moduleAddress=01 **
sensorType=LWR **
4. Initial service property values for this deployment
powerOnDelaySec=0

instrumentName=UNKNOWN

nominalCurrentMa=1000

className=moos.devices.asimet.WhoiAsimet$Attributes

serviceName=ASIMET LWR s/n LWR308

parentID=1481

currentLimitMa=2000

sampleSchedule=R */0 */10 */0 */0 * * * * GMT *

extendedVersion=UNK

frameworkVersion=$Name: march14_2006 $

UUID=3dcc5cbb-47ca-11da-950b-00306e389969

powerPolicy=ALWAYS

timeSynch=false

maxSampleTries=3

isiID=1289

sampleTimeoutMsec=3000

commPowerPolicy=SAMPLING

serviceStatus=0

startDelayMsec=4000

5. Results of acceptance test procedures
All Okay.

6. Known-service bugs/deficiencies
unable to access XML file







