MOOS Mooring 12/11/06 Status Notes

1) Current Status of MSE datastreams:

	MSE Surface Node - October 2006 · log Data streams current

	      Wind Sensor Sensor damaged, returned to ASIMET for repair. Spare unit also needs to be sent in    ASIMET WND (asimetwnd) – 1291 
	 
	 

	      Air Pressure Sensor Sensor damaged, still on buoy.
        ASIMET BPR – 1443 
	 
	 

	      Controller-Power Data stream current.
        Power Can – 1547 
	DataStream 
	 

	      CO2 Monitor Data stream current.
        SOON/PCO2 - 1448
	DataStream 
	 

	      Meteorology Package Sensor damaged, still on buoy.
        ASIMET HRH (asimethrh) - 1290
	DataStream · netCDF · log 
	full deployment | last 7 days 

	      Radiometer-Shortwave Data stream current.
        ASIMET SWR (asimetswr) - 1557
	DataStream · netCDF · log 
	full deployment | last 7 days 

	      ADCP Data stream current.
        Workhorse Long Ranger 75KHz ADCP (adcp) - 1294
	DataStream · netCDF · log 
	full deployment | last 7 days 

	      GPS Data stream current.
        GPS (gps) (gps) - 1566
	 
	 

	
	DataStream · netCDF · log 
	full deployment | last 7 

	      Wave Sensor  Sensor destroyed, replacement unit from MTM3 installation pending. 
        Triaxys Directional Wave Sensor - 1286
	 
	 

	MSE Secondary Surface Node - June 2006 · log Major telemetry backlog

	      Profiler Package Installation planned April 2007.
        SeaHorse Profiler - 1405
	 
	 

	      Radiometer-Hyperspectral Sensor OK, last download: 2006-11-27 17:50:53 -0800       HyperOCR-I - 1517
	DataStream 
	 

	      CTD Sensor OK, last download: 2006-11-27 20:10:30 -0800    CTD (ctd) - 1277
	DataStream · netCDF · log 
	full deployment 

	      Nitrogen Sensor-ISUS Sensor OK, last download: 2006-11-27 20:00:22 –0800 Satlantic ISUS - 1514
	DataStream 
	 

	      GPS Sensor OK, last download: 2006-11-27 20:12:49 -0800        GPS (gps) (gps) - 1504
	 
	 

	
	DataStream · netCDF · log 
	full deployment 

	      Radiometer-Longwave Sensor OK, last download: 2006-11-27 20:13:26 –0800 ASIMET LWR (asimetlwr) - 1289
	DataStream · netCDF · log 
	full deployment 

	Fluorometer Sensor OK, last download: 2006-11-27 20:13:53 -0800  Combination Turbidity and Fluorescence Sensor (ecoflntu) – 1530
DataStream · netCDF · log 

full deployment 

      Oxygen Sensor Sensor not returning data since installation.        Aanderaa Oxygen Optode - 1460 
	DataStream · netCDF · log 
	full deployment 

	MSE 20m Snubber Node - June 2006 · log Major telemetry backlog

	      Radiometer-Hyperspectral Sensor OK, last download: 2006-12-01 15:11:59 –0800 Satlantic HyperOCR - 1516
	DataStream 
	 

	      CTD @ 20.0 m Sensor OK, last download: 2006-12-01 11:51:49 -0800     Seabird 37SMP - 1508
	DataStream · netCDF · log 
	full deployment | last 7 days 

	      Radiometer-Hyperspectral Sensor OK, last download: 2006-12-05 06:11:06 –0800 Satlantic HyperOCR - 1515
	DataStream 
	 

	MSE Shelf BIN Node - July 2006 · log Major telemetry backlog

	      Backscatterometer Sensor OK, last download: 2006-12-05 10:35:03 -0800        WETLabs Triplet - 1535
	Summary DataStream · netCDF · log 
	full deployment | last 7 days 

	
	 
	 

	   Currents Sensor-ADCP Summarizer broken, fix in progress. Last download: 2006-12-05 10:35:39 -800 Aquadopp Current Meter - 1474
	Summary DataStream · netCDF · log 
	full deployment 

	
	DataStream 
	 

	      Oxygen Sensor  Sensor OK, last download: 006-12-05 09:46:05 -0800        Aanderaa Oxygen Optode - 1444
	 
	 

	
	Summary DataStream · netCDF · log 
	full deployment | last 7 days 

	      CTD Sensor OK, last download: 2006-12-05 09:39:40 -0800 Seabird SBE-16+ - 1533
	 
	 

	
	Summary DataStream · log 
	full deployment | last 7 days 

	      Currents Sensor-ADCP Data suspect, gradual decay in range returns, last download: 2006-12-05 07:17:16 -0800  Workhorse ADCP - 1525
	Summary DataStream · netCDF · log 
	full deployment | last 7 days 

	
	 
	 

	MSE Axis BIN Node - November 2006 · log Some telemetry backlog

	      Currents Sensor-ADCP Sensor OK, last download: 2006-12-08 07:19:38 –0800 Workhorse Sentinal 300KHz ADCP - 1549
	 
	 

	
	Summary DataStream · netCDF · log 
	full deployment 

	      Backscatterometer Sensor instrument service lost contact and timed out 2006-12-02 01:08:49 –0800 Spikes in the data – suspect cable getting in instrument measurement area.
        WETLabs Triplet (ecotriplet) (ecotriplet) - 1537
	 Summary DataStream · netCDF · log
	 full deployment

	      Oxygen Sensor shows data spikes, last download: 2006-12-05 20:10:16 –0800 Aanderraa Oxygen Optode - 1445
	 
	 

	
	Summary DataStream 
	 full deployment | last 7 days

	      ADCP No data plot in SSDS, Sensor instrument service lost contact and timed out: 2006-12-09 11:19:49 -0800
        Aquadopp Current Meter - 1473
	Summary DataStream · netCDF · log 
	full deployment | last 7 days 

	
	 
	 

	      CTD  Sensor OK except suspect cable interference with transmissometer light path, spikes in data, last download: 2006-12-11 11:53:55 -0800    SBE16+ CTD – 1546 
	Summary DataStream · netCDF · log 
	full deployment | last 7 days

	
	
	


2) Data telemetry delays and dropped data.

We are experiencing a serious problem returning all of the MSE data in real-time as the data back-logs shown above illustrate. As soon as Kent and Tom are available we’ll work on improving the efficiency of the telemetry link to eliminate the backlog.

3) Other technical data being monitored.


a) 400V GF. Ground fault worsened significantly when Dec. 9-10, 2006 storm approached.





C) SNUBBER HUMIDITY. ..has risen about 5% since deployment and gradually been approaching 50%.





C) SURFACE PRESSURE (HUMIDITY OK). Main Surface housing appears to have lost pressure. Mike K. reports hand tight fitting during field deployment. Humidity remaining low so we’re probably OK.



 


