Aanderaa O2 Optode (ID 1445) - AXIS
1. Commands issued to instrument by service at initialization 

"Set_Interval(0)"  ("0" Sets system to POLL mode)

"Save"   
   (Save changes)

"Set_Protect(1)"  ("1" allows user to change special write-protected            

                            properties)

"Set_OutPut(0)"   ("0" Sets system to simple RS-232C data output mode.)

"Save"                 (Save changes)

2. Commands issued by service for safe mode, if applicable

No “SAFE” mode – cannot log internally.

3. Device metadata returned in SIAM datastream

"Get_All" (Output all property values)
Protect   
3830     605       0

PhaseCoef 
3830     605  -1.470497E+00    1.122573E+00    0.000000E+00 0.000000E+00

TempCoef  
3830     605   2.741038E+01   -3.115655E-02    2.898775E-06 -4.338472E-09

FoilNo    
3830     605    4104

C0Coef    
3830     605   3.199840E+03   -1.119634E+02    2.408296E+00 -2.24874  0E-02

C1Coef    
3830     605  -1.744936E+02    5.462500E+00   -1.244084E-01    1.23915  3E-03

C2Coef    
3830     605   3.941711E+00   -1.086677E-01    2.719839E-03   -2.90634  3E-05

C3Coef    
3830     605  -4.220910E-02    1.018155E-03   -2.905609E-05    3.30661  0E-07

C4Coef    
3830     605   1.738870E-04   -3.637668E-06    1.227403E-07   -1.46839  9E-09

Salinity         
3830     605   0.000000E+00

CalAirPhase      3830     605   2.557295E+01

CalAirTemp       3830     605   2.086279E+01

CalAirPressure  3830     605   9.944000E+02

CalZeroPhase   3830     605   6.136143E+01

CalZeroTemp    3830     605   2.151919E+01

Interval         
3830     605       0

AnCoef    
3830     605   0.000000E+00    1.000000E+00

Output    
3830     605       0

SoftwareVersion3830     605       3

SoftwareBuild    3830     605       7
4. Instrument-specific service properties
None.
5. Initial service property values for this deployment
powerOnDelaySec=0

nominalCurrentMa=1000

summaryVars=Oxygen 

serviceName=Aanderaa Optode s/n 605

packetSetSize=10

sampleSchedule=R */10 */0 */0 */0 * * * * GMT *

errCacheLimit=5

UUID=29607c68-2937-11da-80c8-00306e389969

timeSynch=false

maxSampleTries=3

peakCurrentMa=5000

summaryTriggerCount=360

sampleTimeoutMsec=5000

serviceStatus=0

startDelayMsec=1500

isiID = 1445

powerPolicy = sampling

commPowerPolicy = sampling

sampleSchedule = 10

summaryTriggerCount = 360

maxSummarySamples = 360

currentLimitMa = 3000
defaultSkipInterval=0

startDelayMsec=1500

6. Results of acceptance test procedures
     All OK.

7. Summarization

Not applicable.

8. Instrument clock synchronization

Not implemented.

9. Known-service bugs/deficiencies
None.
