Aanderaa O2 Optode (ID 1524) - Shelf
1. Commands issued to instrument by service at initialization 

"Set_Interval(0)"  ("0" Sets system to POLL mode)

"Save"   
   (Save changes)

"Set_Protect(1)"  ("1" allows user to change special write-protected            

                            properties)

"Set_OutPut(0)"   ("0" Sets system to simple RS-232C data output mode.)

"Save"                 (Save changes)

2. Commands issued by service for safe mode, if applicable
No “SAFE” mode – cannot log internally.

3. Device metadata returned in SIAM datastream
"Get_All" (Output all property values)

Protect

3830
677     0

PhaseCoef
3830
677  -6.735069E+00       1.181544E+00   0.000000E+00
 0.000000E+00

TempCoef
3830
677   2.249353E+01
-3.107436E-02  2.934541E-06
-4.165899E-09

FoilNo

3830
677    4804

C0Coef

3830
677   3.172420E+03
-1.072609E+02  2.133159E+00
-1.792340E-02

C1Coef

3830
677  -1.739810E+02
 5.102620E+00 -9.857580E-02
 8.022400E-04

C2Coef

3830
677   3.956000E+00
-9.810400E-02  1.853460E-03
-1.425360E-05

C3Coef

3830
677  -4.263370E-02
 8.788200E-04 -1.644089E-05
 1.137090E-07

C4Coef

3830
677   1.768690E-04
-2.955020E-06  5.550390E-08
-3.084930E-10

Salinity

3830
677   0.000000E+00

CalAirPhase
3830
677   2.842749E+01

CalAirTemp
3830
677   1.984647E+01

CalAirPressure
3830
677   1.025520E+03

CalZeroPhase
3830
677   6.393019E+01

CalZeroTemp
3830
677   2.048111E+01

Interval
 
3830
677       0

AnCoef
 
3830
677   0.000000E+00
 1.000000E+00

Output
 
3830
677       0

SoftwareVersion3830
677       3

SoftwareBuild
3830
677       7

4. Instrument-specific service properties
None.
5. Initial service property values for this deployment
powerOnDelaySec=0

nominalCurrentMa=1000

summaryVars=Oxygen 

serviceName=Aanderaa Optode s/n 677

packetSetSize=10

sampleSchedule=R */0 */5 */0 */0 * * * * GMT *

defaultSkipInterval=-1

errCacheLimit=5

UUID=26abd9d3-a8be-11da-bb39-00306e389969

timeSynch=false

maxSampleTries=3

peakCurrentMa=5000

summaryTriggerCount=12

sampleTimeoutMsec=5000

serviceStatus=0

startDelayMsec=1500

isiID = 1524

powerPolicy = sampling

commPowerPolicy = sampling

sampleSchedule = 300

summaryTriggerCount = 12

maxSummarySamples = 12

currentLimitMa = 3000

defaultSkipInterval = -1

startDelayMsec=1500
6. Results of acceptance test procedures
     All OK.
7. Summarization

Not applicable.

8. Instrument clock synchronization
Not implemented.

9. Known-service bugs/deficiencies
None.






