ASIMET - QUICK CHECK OUT

14 January, 2002

This procedure is for a module that has been newly manufactured or repaired and calibrated.  The module should also be prepared or "preped" prior to quick check out by zeroing the flash memory card, setting the time and data to UTC, and installing a new 12vdc or 15vdc battery.  The module will have the battery disconnected if it has just been shipped to save energy.

1.  Open the housing with a large flat blade screwdriver and remove the housing.

2.  Using a computer running ProCom (or other terminal program) at (9600, N, 8, 1), power supply, and RS485 converter, plug into the module as shown in [CONNECT.DWG] into the P2 connector and with the JP1 set toward the edge of the card (or toward the P2 connector in older units).

3.  Turn on power and monitor the current or if a new battery is installed, plug the battery into P1.  

4.  Type [#WND01A] or HRH for humidity etc.  Watch the return:  WND01.  Communications have been established.

5.  Type [#WND01H] to get the help menu.  Ref the web site for the return.

6.  Type [#WND01L] to get the ID, SN, cal etc info.  Check the time.  Check the memory or "Records Used" = 0.

7.  Type [#WND01T] to go into the test mode.  Check that the output is reasonable for the module being tested. Use <esc> to exit the test mode.

8.  Type [#WND01A] to get the WND01 return.  This checks that the unit is out of the test mode and is collecting one minute data.

9.  Re-install the housing.  Check the o-rings.  Connect the bottom connector to the end cap.  Install the end cap using the large flat blade screwdriver.  Install a dummy plug.

The module is ready to install.

