snubber_1482-SBE37SM_1534

1. Manually-configured instrument settings not handled by service

SBE37-SM-PUCK V 1.0-01  SERIAL NO. 2765    01 Jan 1996  02:51:50

logging not started

sample interval = 60 seconds

samplenumber = 0, free = 183202

do not transmit real-time data

do not output salinity with each sample

do not output sound velocity with each sample

store time with each sample

number of samples to average = 4

serial sync mode disabled

wait time after serial sync sampling = 30 seconds

PUCK mode enabled

internal pump not installed

temperature = 23.14 deg C

SBE37-SM-PUCK V 1.0-01  2765

temperature:  12-dec-02

    -4.215792e-05

    2.839508e-04

    -3.000133e-06

    1.696669e-07

conductivity:  12-dec-02

    -1.051158e+00

    1.466496e-01

    -9.150801e-05

    3.157882e-05

    -9.570000e-08

    3.250000e-06

    0.000000e+00

pressure S/N 2598, range = 160 psia:  09-dec-02

    5.945600e-02

    7.603297e-03

    -7.831820e-10

    5.468956e+01

    1.021594e+00

    -2.910683e-02

    2.494063e+01

    9.250000e-04

    0.000000e+00

    0.000000e+00

rtc:  12-dec-02

    9.999726e-01

    1.708880e-06

    -3.062022e-08

2. Device metadata returned in SIAM datastream

See (1)

3. Instrument-specific service properties

None

4. Initial service property values for this deployment

powerOnDelaySec=0

instrumentName=UNKNOWN

nominalCurrentMa=1000

className=org.mbari.isi.interfaces.InstrumentServiceAttributes

serviceName=Seabird-37SM s/n 37SM30960-2765

parentID=1482

currentLimitMa=5000

sampleSchedule=R */0 */10 */0 */0 * * * * GMT *

extendedVersion=UNK

frameworkVersion=$Name: march14_2006 $

UUID=5832530d-ad76-11da-bb39-00306e389969

powerPolicy=ALWAYS

timeSynch=false

maxSampleTries=3

isiID=1534

sampleTimeoutMsec=15000

commPowerPolicy=SAMPLING

serviceStatus=0

startDelayMsec=2000

5. Results of acceptance test procedures

All Passed

6. Known-service bugs/deficiencies

unable to access XML file

instrument time/date not correct







