Surface_1446-MSP430_1447

1. Manually-configured instrument settings not handled by service

Detection thresholds (not currently used)

2. Device metadata returned in SIAM datastream

$PSID, V2.3.1-Q [Mar 12 2004], *1811

$PALEVEL, PL0.000000, PH0.000000, TL0.000000, TH0.000000, HL0.000000, HH0.000000, GFLL0.000000, GFLH0.000000, GFHL0.000000, GFHH0.000000, CL0.000000, CH0.000000, TCL0.000000, TCH0.000000, *984

$PAMASK, 00, *729

$PALARM, 00, *729

3. Instrument-specific service properties

turnsSetpoint=0

setTurns=false

4. Initial service property values for this deployment

powerOnDelaySec=0

instrumentName=UNKNOWN

nominalCurrentMa=1000

className=moos.devices.msp430.MSP430$Attributes

serviceName=MSP430 (sidearm4.25)

parentID=1446

currentLimitMa=500

sampleSchedule=R */0 */10 */0 */0 * * * * GMT *

extendedVersion=UNK

frameworkVersion=$Name: march14_2006 $

UUID=a92f763a-29ff-11da-80c8-00306e389969

powerPolicy=ALWAYS

timeSynch=false

maxSampleTries=3

isiID=1447

sampleTimeoutMsec=10000

commPowerPolicy=ALWAYS

serviceStatus=0

startDelayMsec=500

5. Results of acceptance test procedures

All passed

6. Known-service bugs/deficiencies

unable to access XML file

Turns counter does not work?

GFL= -450 uA







