Instrument name here

1. Commands issued by service at initialization 

        stopSampling();  //"STOP\r".

        if (_attributes.log) {

            // Take sample, log to instrument's FLASH

            _requestSample = "TSS\r".getBytes();

            _requestSampleEcho = "TSS\r\n".getBytes();

        }

        else {

            // Take sample, do NOT log to instrument's FLASH

            _requestSample = "TS\r".getBytes();

            _requestSampleEcho = "TS\r\n".getBytes();

        }

                getPrompt(); // "\r\n"
        // Set data format

        _toDevice.write(getFormatForSummaryCmd());//"FORMAT=1\r"
        // Set the Seabird clock to "now"

        setClock(System.currentTimeMillis());//"MMDDYY="…"HHMMSS="

2. Commands issued by service for safe mode, if applicable

       // Get to a prompt.

        getPrompt(); // \r\n
        // Set the internal clock before going into autonomous mode.

        setClock(System.currentTimeMillis());//"MMDDYY="…"HHMMSS="
        // Get to a prompt again.

        getPrompt();//\r\n
        // Set sample interval (seconds). Default is 300s.

        String interval = ("INTERVAL=" + _attributes.safeSampleIntervalSec + "\r\n");

        _toDevice.write(interval.getBytes());

        // Start autonomous sampling immediately.

        _toDevice.write(_startNow); // "Startnow\r"

3. Device metadata returned in SIAM datastream
powerOnDelaySec=0

instrumentName=UNKNOWN

nominalCurrentMa=1000

className=moos.devices.seabird.Seabird$Attributes

summaryVars=temperature conductivity 

serviceName=Seabird-37SM s/n 37SMP41725-4590

parentID=1482

maxSummarySamples=0

currentLimitMa=3000

packetSetSize=10

sampleSchedule=R */0 */10 */0 */0 * * * * GMT *

defaultSkipInterval=0

extendedVersion=UNK

frameworkVersion=$Name: june09_2006 $

errCacheLimit=5

UUID=7ba4f00c-94fb-11da-92c6-00306e389969

powerPolicy=ALWAYS

timeSynch=false

maxSampleTries=3

isiID=1508

log=true

peakCurrentMa=5000

summaryTriggerCount=0

sampleTimeoutMsec=15000

commPowerPolicy=SAMPLING

safeSampleIntervalSec=600

serviceStatus=2

startDelayMsec=2000

isiID = 1508

serviceName = Seabird-37SM s/n 37SMP41725-4590

sampleSchedule = 600

powerPolicy = always

commPowerPolicy = sampling

safeSampleIntervalSec = 600

currentLimitMa = 3000

UUID = 7ba4f00c-94fb-11da-92c6-00306e389969

SBE37-SM-PUCK V 1.0-01  SERIAL NO. 4590    11 Jun 2006  22:12:24

logging not started

sample interval = 300 seconds

samplenumber = 463, free = 182739

do not transmit real-time data

do not output salinity with each sample

do not output sound velocity with each sample

store time with each sample

number of samples to average = 4

serial sync mode disabled

wait time after serial sync sampling = 30 seconds

PUCK mode enabled

internal pump is installed

temperature = 24.72 deg C

SBE37-SM-PUCK V 1.0-01  4590

temperature:  22-jan-06

    -8.721949e-05

    2.893646e-04

    -3.293638e-06

    1.810673e-07

conductivity:  22-jan-06

    -1.012003e+00

    1.410226e-01

    -1.667405e-04

    3.524536e-05

    -9.570000e-08

    3.250000e-06

    -6.806242e-06

pressure S/N 9315, range = 508 psia:  18-jan-06

    -1.479434e-03

    2.409061e-02

    2.906055e-09

    2.686157e+01

    2.526582e-01

    -1.190681e-02

    2.502950e+01

    1.000000e-04

    0.000000e+00

    0.000000e+00

rtc:  22-jan-06 

    9.999854e-01

    1.596637e-06

    -3.227476e-08

4. Instrument-specific service properties
5. Initial service property values for this deployment
6. Results of acceptance test procedures
N/A
7. Known-service bugs/deficiencies

none






