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MVP Demo Cheat Sheet

The slit and focus stages are run on Ramansbc through PCAnywhere/VNC/NetMeeting

Shortcut to mvpdemo is on Ramansbc desktop 

(executable file is at C:\program files\National Aperture’s Mvpdemo\ )

Command File Processing mode

A set of command files have been created for regular commands.  Each file is prefaced by “slit” or “focus” to indicate which stage it operates.


slit_setup.cmd
— used on startup to initiate, home, and rezero stage


slit_home.cmd
— moves stage to home position


(for the focus stage, home is aft in the pressure housing)


slit_incrXX.cmd
— increments by XX microns (increment moves away from home)


slit_decrXX.cmd
— decrements by XX microns (decrement moves towards home)

Manual Command Entry or Terminal Emulation mode

Every command begins with a 1 or a 2 to indicate which stage it operates (1 = slit, 2 = focus)


st
query node
1 st


lr
load relative
1 lr 1000


la
load absolute
1 la 1000


m
move
1 m


ab
abort
1 ab


pos
query position
1 pos


ho
define home position
1 ho


NOTE:  When in Terminal Emulation mode, position data is shown in hex

Distance reference for slit stage:


10 μm 
= 20 counts


1 μ 
= 2 counts


½ μ 
= 1 count


16000 counts = 8 mm

NOTE:  Sometimes cursor is not visible.  Tab and arrow keys can help to navigate the window.

NOTE:  Don’t drive focus stage past 167,000.  If the probe head hits the window it will jam the motor stage.

Configuration

COM 3
19200 baud

NOTE:  If using serial port of computer, set baud to 19200, use straight cable (NOT null).
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