DORISS/PUP Mtg Minutes

2006-01-09

Present:  Peter Brewer, Rachel Dunk, Bill Kirkwood, Ed Peltzer, Alana Sherman & Peter Walz.

DORISS I relocation to Brewer lab:


1.  Re-organize lab benches in the lab:  PMW to speak with Greg Voogd.


2.  Dis-mount electronics from inside of end-bells?   YES.


3.  Bench vs wall-mount:  Bench; but build something to protect hardware.


4.  Fix slit lens  (by Bill Kirkwood).

5.  Use Mark II probhead.


Estimated completion date:  by April.

DORISS II heater tests:


1.  Temp fix (heating tape) appears to be working; use 2/6-8.


2.  Long-term fix: do right = add internal heaters and temp sensors



To be installed AFTER February dives



Alana is looking for components including additional temp sensors.



When heaters are installed, temp sensors will be calibrated.



Also, the cooling fans for the PC104 stack will be re-wired to start at




power-up.

DORISS II / PUP Mob day = 2/3.


*** Full systems comms check on Ventana.

Other DORISS issues:


1.  We need Karen S. to build us a new image for FLASH memory plus fix some spares that were corrupted during dives last year.


2.  Alana will fix the change made to VIC.


3.  When will we interface the Mark II probe-head?


4.  Maintain a latent interest for future upgrades and improvements.

Fiber-optic cable upgrade / replacement


Current cables are getting old and lose signal under pressure.


Most sub-sea cables are not flexible enough for our purposes (>1m bend radius).

Bill K suggests that we adapt Kaiser cables with pressure compensation.


Kaiser cables need to be tested at full operating lengths  (20m).


New cables should use telephone standard fibers and connectors.

Annual report due soon: send bullets to Etp3 by 1/16.

