DORISS/PUP System Test Checklist

Communications check:


Check Mudpuppy (or Ramandesktop)
C:\> ping mudpuppy



IP address should be 134.89.10.93 (or 134.89.10.92)


Check Galil (134.89.10.96)
C:\> ping galil


Check Axis Server (134.89.10.95)
C:\> ping axis


Check Ramansbc
C:\> ping ramansbc or doriss



IP address should be 134.89.10.94  (It takes 2 minutes to fully boot up.)

PUP GUI check:

Double-click on PUP.exe icon to start PUP GUI

Check all features

· Make sure Z and R axes numbers look reasonable

· Make sure tilt, temperature, humidity numbers look reasonable

· Turn camera on

· Double-click on Internet Explorer and check camera output

· Turn lasers on and off

· Turn light on and off

· Turn camera off

Check positioning

· Jog Z axis 50 mm in the positive direction, 30 mm in the negative direction

· Jog R axis 50 mm in the positive direction, 30 mm in the negative direction

· Run Homing command (double click on button with house)

· Make sure R and then Z positions decrease and set to zero 

· (homing completed when error numbers are ~0)

· Go to position 150, 150

· Go to position 0, 0

DORISS check:

Open PCAnywhere/VNC/NetMeeting and login to ramansbc (Administrator, LRS)

· Verify that temperature & humidity date are scrolling

· Open LRSClient (should take a few minutes, click Ignore for any error message)

· (If it doesn’t work close it, wait 5 minutes, then try to open it again)

Open HoloGRAMS by double-clicking on the HoloGRAMS icon on the topside desktop


(I2C error is normal)

· In Acquire tab, set exposure to 1 sec, click Focus button (may have to hit button again).

· Verify that spectrum refreshes every second.

· Stop Focus mode and exit HoloGRAMS

After checking all features, exit everything except for PCAnywhere/VNC/NetMeeting.  

On ramansbc desktop select Shutdown – turn off power to system after 35 seconds.

