To: Andrew Hamilton, MBARI

From: Owen Raybould

Date: September 13, 2010

Subject:  Wave Generation Thermocouple Experiment

Objective

The purpose of this experiment was to develop a simulation with similar conditions to that which the thermocouple will be used in the Wave Generation project at sea. It is important to test the accuracy and response time of the thermocouple in its changing environment before being deployed.

Procedure

In order to create an environment similar to that which the thermocouple will be used in a pressure cylinder was assembled. The flange at the end of the cylinder has four outlets: an exhaust with a ½” release valve to release the pressure quickly, a pressure transducer to record the pressure change, a ‘T’ type thermocouple to record temperature change and an intake with shutoff valve to pressurize the cylinder. Once everything was connected securely, the tank was pressurized to roughly 250psi with Nitrogen, allowed to equilibrate to the surrounding air’s temperature and then released. Measurements from the pressure transducer and thermocouple were taken from the Tektronix TSP 2015 Oscilloscope and then transferred into an excel spreadsheet. An amplifier was installed between the thermocouple and the oscilloscope to amplify the voltage in order for more accurate readings. From these results a time constant could be estimated to justify the selection of this type and size of thermocouple.

Results

Amplifier Gain

Before any results could be analyzed the amplifier had to be calibrated to know the approximate gain and offset. To do so a precision voltage source was connected to the amplifier with accuracy down to 0.1mV’s. The output from the amplifier was then read on the Oscilloscope and recorded onto excel and compared to the input from the precision voltage source. From there a relation could be determined between input and output of the amplifier to determine the approximate gain as seen in figure 1. Once the gain was known a correlation between temperature and input voltage to the amplifier was determined. To do so two known temperature were used, room temp (19.7°C) and ice water (4.8°C) were used to see the voltage output of the thermocouple, by doing so the offset voltage created by not having a cold junction in the amplifier was able to be accounted for. By comparing these voltages from known temperatures to published data on ‘T’ type thermocouple a polynomial relation was determined between the temperature and amplifier input voltage as seen in figure 2.

Figure 1: By using the precision voltage source a known input voltage accurate to 0.1mV was put into the amplifier where the output was read by the Oscilloscope. From there a linear relation of input to output was determined by the equation y = 1.6423x-1.0953 where y in the input and x is the amplified output voltage.

Figure 2: Relation of temperature to amplifier input voltage (thermocouple output) was used to get the polynomial equation of temperature change to voltage so that the temperature change in the experiment can be determined.

Compressed Nitrogen

Once the amplification level was calibrated and the offset was accounted for the results from the three tests could be analyzed and graphed. First the pressure drop was graphed and an exponential tread line was fit to determine alpha, a value used to determine the temperature change in an adiabatic process.

Pt= P0e-αt

From this equation the changing pressure can be used to determine the theoretical temperature change by rearranging the equation for a reversible adiabatic process shown below.

P(t)1-γT(t)γ= P01-γT0γ              (γ=1.4 for adiabatic process)

In all three tests the values of alpha were within 2% of each other which validates the values used to determine the theoretical temperatures. Finally, using the equations from the amplifier calibration tests the experimental temperatures could be graphed from the thermocouples output voltage. As seen in figures 3-5 the experimental temp does not drop as fast as the theoretical temp; this is most likely due to heat given off by the surrounding steel cylinder. As the pressure drops it cools down the cylinder as well as the gas therefore the temperature of the nitrogen never drops as much as the theoretical value predicts. After the three experiments were conducted and the results analyzed, it was demonstrated that the thermocouples response time is sufficient for its application so the chosen size and type of thermocouple is appropriate for this project.

Figure 3: First recorded measurements from the oscilloscope used to demonstrate the lag in the response time of the thermocouple. The theoretical temperature (red) drops quicker than the experimental temperature (green) proving there is lag in the thermocouple. But it must also be noted that the surrounding steel cylinder conducts heat into the Nitrogen warming it up as the pressure drops meaning the assumption of an adiabatic process is not entirely accurate. It is due to this heat transfer that the experimental temperature does not drop as low as the theoretical temp and could also contribute to the lag in temperature change.  In this test alpha equals 0.709.

Figure 4: Second recorded measurements from the oscilloscope used to demonstrate the lag in the response time of the thermocouple.  Here alpha equals 0.713 which is very close to the constant determined in the first experiment.

Figure 5: Third recorded measurements from the oscilloscope used to demonstrate the lag in the response time of the thermocouple and used to validate the above results.

Ice Bath to Ambient Air Time Constant

Table 1: Results from time constant experiment with thermocouple initially in ice bath (0.1 °C) and quickly moved to surrounding air (20.4°C). After ten tests the average time constant was 145.5ms. The column with ms/div more or less represents how accurate the measurement was.

|  | Ice bath to ambient air |  |  |  |
| --- | --- | --- | --- | --- |
|  | ∆V (mV) | 63.2% of ∆V (mV) | Time Constant, τ (ms) | ms/div |
|  | 700 | 442.4 | 117 | 25 |
|  | 650 | 410.8 | 138 | 25 |
|  | 660 | 417.12 | 132 | 50 |
|  | 720 | 455.04 | 188 | 50 |
|  | 660 | 417.12 | 156 | 50 |
|  | 680 | 429.76 | 158 | 50 |
|  | 680 | 429.76 | 128 | 50 |
|  | 680 | 429.76 | 138 | 50 |
|  | 700 | 442.4 | 146 | 50 |
|  | 780 | 492.96 | 154 | 50 |
| Averages | 691 | 436.712 | 145.5 | 45 |

Figure 6: Example of the data points used to determine the time constant of the thermocouple when removed from ice bath into surrounding air. In this case the time constant is much larger than when moved from air to ice because some of the water remains on the thermocouple junction so it takes longer for the water to change to room temperature. For this reason, the resulting time constant for this part of the experiment should not be taken into consideration when deciding if this is the appropriate thermocouple for this application since there will not be a liquid median in the real project.

Ambient Air to Ice Bath Time Constant

Table 2: Results from time constant experiment with thermocouple initially in ambient air (20.4°C) and quickly placed into the ice bath (0.1 °C). After ten tests the average time constant was 9.26ms. The column with ms/div more or less represents how accurate the measurement was; in this case the measurements are more accurate than from ice to air.

|  | Ambient air to ice bath |  |  |  |
| --- | --- | --- | --- | --- |
|  | ∆V (mV) | 63.2% of ∆V (mV) | Time Constant, τ (ms) | ms/div |
|  | 890 | 562.48 | 9.2 | 10 |
|  | 890 | 562.48 | 8.6 | 10 |
|  | 890 | 562.48 | 8.8 | 10 |
|  | 880 | 556.16 | 8.4 | 10 |
|  | 888 | 561.216 | 7.2 | 10 |
|  | 880 | 556.16 | 11.8 | 10 |
|  | 872 | 551.104 | 9.6 | 10 |
|  | 840 | 530.88 | 9.6 | 10 |
|  | 940 | 594.08 | 9.4 | 10 |
|  | 860 | 543.52 | 10 | 10 |
| Averages | 883 | 558.056 | 9.26 | 10 |

Figure 7: Example of the data points used to determine the time constant of the thermocouple when dipped into an ice bath. The time constant is found from the time it takes for 63.2% of the total change in voltage to occur.