EyeRIS (Remote Imaging System)

MBARI Project 800139

- Functional Requirements (General)

  - EyeRIS shall be configurable for a range of particle sizes, particle concentrations, flow rates, and platform profiling speeds for a specific use case Add a table of min size vs flow rate to show tradeoff

  - EyeRIS shall measure particle sizes in the range 250-10000 microns with a resolution of 100 um

  - EyeRIS shall be able to bin particle sizes with 10% accuracy (min 100 particles per bin)

  - EyeRIS shall provide images sufficient for morphological identification of objects in the range 1 mm to 20 cm

  - EyeRIS shall provide a quantifiable imaged volume

  - EyeRIS shall provide two direct color LED illumination, one for minimal animal disturbance [Red or Near IR]

- Functional Requirements (Cabled Instrument):

  - EyeRIS shall stream live video to a surface vessel

  - EyeRIS shall record data on an external high-resolution recorder on a surface vessel

  - EyeRIS shall operate to a depth of 4000 meters

  - EyeRIS shall be compatible with existing ROV power and communications

  - EyeRIS shall be deployable on a manipulator arm of an ROV

- Functional Requirements (Autonomous Instrument)

  - EyeRIS shall record image and video data locally

  - EyeRIS shall operate to a depth of 4000 meters

  - EyeRIS shall operate autonomously for 24 hours, 0.5 Hz frame rate, daily profiles?

- Functional Requirements (3DPIV):

  - EyeRIS shall be able to resolve objects in the size range from 1-20 cm (x/y dimension)

  - EyeRIS shall be able to resolve objects in the size range 1-20 cm (z dimension)

  - EyeRIS shall measure fluid velocities from 1mm/s to 1 m/s

- Functional Requirements (3D(dt)R):

  - EyeRIS shall resolve features on the order of 1% of size on objects in the size range 1-20 cm (x/y/z).

  - EyeRIS shall resolve mucus and gelatinous structures (> 50% opacity)

  - EyeRIS shall be able to resolve objects in the size range from 1-20 cm (x/y dimension)

  - EyeRIS shall size targets with accuracy better than 5% x/y/z

  - EyeRIS may use auxiliary/multiple cameras to verify target measurements.

  - EyeRIS may use structured illumination to enhance surface reconstruction

- Functional Requirements (Plankton Imaging, CeNCOOS):

  - EyeRIS shall detect and size particles with size 150 um and larger with resolution of 50 um

  - EyeRIS shall yield quantifiable particle abundance estimates for particles larger than 150 um.

  - EyeRIS shall provide size estimates from particles with resolution up to the quantified effective resolution of the main lens and micro-lens-array combination, roughly 100 um from current estimates.

  - EyeRIS shall provide data output in formats compatible with CeNCOOS workflows

  - EyeRIS shall provide raw camera data for still image capture for posterity