HELIUM LEAK TEST DATA SHEET

	Test Engr: Gene Massion
	Date: Feb 14, 2008

	Project: MARS LV/COMMS Housing Wet Side

	Maximum Acceptable Leak Rate: 5 * 10e-7 std cc/sec He

	Comments: Test dry side before ship to UW


	Equipment Used: MBAR Pfeiffer HLT260 SN 4011006016

	Cal Date: Factory Cal March 2006, Internal Cal Feb 14, 2008

	Gas: HE 5.0 Ultra High Purity


	Test Results

	Initial Leak Rate (before He):
	3.5 * 10e-7 and dropping

	Apply Helium at the Following Locations

	Circ. of inner bulkhead
	MV telemetry connector

	All inner bulkhead penetrators
	Pressure relief

	Circ. Of end cap

	All 8 ODI penetrators

	All of the above with both fill ports open and tester at bottom port: Leak rate 5e-8 and dropping

	Close upper port, add He to upper port fill plumbing

	Close lower port, open upper port wit He

	Leak rate still 3e-8 and dropping
	

	
	

	Maximum Measured Leak Rate:
	

	Comments: 

	

	

	

	Pass/Fail 
	Pass

	Signature: Gene Massion


