AquaFlash Connector
The AquaFlash uses a 7 pin interface connector, the MBH-7-MP from Impulse. An
optional 2 ft mating cable with locking sleeve is available for sale with the AquaFlash for
splicing into a camera system. Longer cables are available upon request, or directly from Impulse. The AquaFlash incorporates cast threads to mate with the mating cable retainer. Note that this threaded section is thin, and is not intended to handle great stresses.
The following table details the connections.

Pin Name Description
1 Mode Select Pull low to activate. Internal 82K resistor to 5V
2 Supply Nominal 12VDC: 9V-15V allowable range. 16V Absolute
Maximum! Peak current draw is about 3 amps.
3 Quench Pull low to stop flash, in TTL mode. Internal 82K resistor to 5V
4 Enable Pull low to enable charging and flash output. Internal 82K
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resistor to 5V.
5 Trigger Pull low to initiate flash. Internal 82K resistor to 5V. In slave
mode, flash will be terminated when Trigger goes high.
6 Ground Signal & Power common.
7 Ready Charge Status: 1K to Supply when charged, Open if not active,
and toggling if charging. 
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