## BED 3
2015-09-29
Pressure sensor offset calibration | [13 empty] | Forward calculation for External 3-wire Pressure Sensor 0-70 BAR Absolute, Ratiometric, 0.5-4.5V Output; US331-000004-070BA
| V sensor [V] (voltage measured at sensor output, 0-5V range) | [8 empty] | Note | Sensor SN | Sensor model |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
Pressure applied to sensor in addition to atmospheric pressure [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 70 | 0.5 | 4.5 | -8.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -96.9974
BED/Sensor Number (Volts) | [14 empty] | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 70 | 0.5 | 4.5 | 0 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -9.9485
BED 3 | 0.51 | 0.72 | 0.92 | 1.13 | [7 empty] | PHE163 100 bar |  |  | 65535 | 11.14% | 7302.472 | 2.5 | 0.27857 | 0.55714 | 0 | 70 | 0.5 | 4.5 | 1 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 7.4613e-06
BED 3 | 0.5105 | 0.7116 | 0.9193 | 1.1289 | 1.3378 | 1.5401 | 1.7423 | 1.9378 | 2.1316 | DVK 9/23/16, fluke 189 | 20298001 | PHE163 100 bar |  |  | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 70 | 0.5 | 4.5 | 13.125 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 120.625
BED 4 | [11 empty] | PHE163 100 bar |  |  | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 70 | 0.5 | 4.5 | 35 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 338.2473
BED 5 | [11 empty] | PHE163 100 bar |  |  | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 70 | 0.5 | 4.5 | 56.875 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 555.8697
BED 6 | [11 empty] | PHE163 100 bar |  |  | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 70 | 0.5 | 4.5 | 70 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 686.4431
[15 empty] | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 70 | 0.5 | 4.5 | 78.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 773.4921
BEDS 2 | [9 empty] | DVK 10/8/2015 |  | PHE163 100 bar
BED 7 | [9 empty] | DVK 4/4/2016 | 21110001 | PHE163 100 bar
BED 8 | [9 empty] | DVK 4/4/2016 | 21110002 | PHE163 100 bar |  |  | Forward calculation for External 3-wire Pressure Sensor 0-100 bar Sealed Gauge, Ratiometric, 0.5-4.5V Output; PHE163 100 bar
BED 9 | [9 empty] | DVK 4/4/2016 | 21110003 | PHE163 100 bar |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
BED 10 | [9 empty] | DVK 4/4/2016 | 21110004 | PHE163 100 bar |  |  | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 100 | 0.5 | 4.5 | -12.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -124.3556
BED 11 | [9 empty] | DVK 4/4/2016 | 21110005 | PHE163 100 bar |  |  | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 100 | 0.5 | 4.5 | 0 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 0
Pressure on sensor + P atmosphere [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 11.14% | 7302.4713 | 2.5 | 0.27857 | 0.55714 | 0 | 100 | 0.5 | 4.5 | 1.4286 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 14.2121
[15 empty] | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 100 | 0.5 | 4.5 | 18.75 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 186.5335
[15 empty] | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 100 | 0.5 | 4.5 | 50 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 497.4225
[15 empty] | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 100 | 0.5 | 4.5 | 81.25 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 808.3116
[13 empty] | Legend: |  | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 100 | 0.5 | 4.5 | 100 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 994.8451
volts = 2.50 * (counts/65536) | [12 empty] | constant |  | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 100 | 0.5 | 4.5 | 112.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 1119.2007
(pressFullScale/2) * volts - pressFullScale/8 | [12 empty] | auto calculated
for 100 bar sensor: | [12 empty] | input |  | Forward calculation - Equations, plus simplified equations for PHE163 100 bar sensor
50.0 * volts - 12.5 | [14 empty] | Constant = 2^16-1 | Input | Counts FS * Counts%FS | Constant | Vref * Counts / Counts FS | Vadc * 2 | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | (Pmax-Pmin) / (Vsens max-Vsens min) * (Vsens- Vmin) + Pmin | Constant (sensor parameter) | Constant (atmospheric pressure) | Constant (gravity) | Constant (water density) | 1/ g [m/s2] / R [kg/m3] | k [m/Pa] * 1E5 [Pa/bar] | k [m/bar] * (Pcalc [bar] + Pref [bar] - Patm [bar])
[15 empty] | 65535 | input | 65535 * input | 2.5 | 2.5 * counts / 65536 | 5 * counts /65536 | 0 | 100 | 0.5 | 4.5 | Pmax / 4 * Vsens - Pmax /8 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 9.948* Pcalc [bar]
Calibration Vsens vs Pressure
Linear 4 point calibration |  | Linear 9 point calibration |  | polynomial 4 point |  | polynomial 9 point |  | Nominal
slope | 24.2696 |  | 24.5456 | a | 0 |  | 0.5773 | slope | 25
intercept | -12.401 |  | -12.6182 | b | 24.297 |  | 23.0199 | Intercept | -12.5
|  |  |  | c | -12.401 |  | -11.7697
R2 | 0.99991 |  | 0.99984 |  | 0.99999 |  | 0.99996
Water depth differences according to different calibrations, BED# 3
[13 empty] | Nominal Depth, m |  | Water depth difference meters, Lineal- polynomial | Water depth difference, meters Lineal -nominal | Water depth difference ,meters Polynomial - Nominal
| Pressure,Bars Linear calibration | Pressure, Bars Polynomial 4 point calibration | Pressure, Bars Polynomial 9 point calibration | Pressure, Bars Nominal calibration |  | Delta pressure, bars, Linear - Polinomial 4 point | Delta pressure, bars, Linear - Nominal | Delta pressure, bars, Polynomial 4 point - Nominal | Delta pressure, bars, Polynomial 4 point - Polynomial 9 point | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar]
Volts | [14 empty] | 0 |  | -0.13667 | -2.6489 | -2.5122
0.5 | -0.26626 | -0.25252 | -0.11548 | 0 |  | -0.013738 | -0.26626 | -0.25252 | -0.13704 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 62.1778 |  | -0.20483 | -4.4656 | -4.2607
0.75 | 5.8011 | 5.8217 | 5.8199 | 6.25 |  | -0.020589 | -0.44887 | -0.42828 | 0.0018262 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 124.3556 |  | -0.27298 | -6.2822 | -6.0093
1 | 11.8685 | 11.896 | 11.8274 | 12.5 |  | -0.02744 | -0.63148 | -0.60404 | 0.06853 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 186.5335 |  | -0.34114 | -8.0989 | -7.7578
1.25 | 17.9359 | 17.9702 | 17.9071 | 18.75 |  | -0.03429 | -0.81409 | -0.7798 | 0.063071 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 248.7113 |  | -0.40929 | -9.9156 | -9.5063
1.5 | 24.0033 | 24.0444 | 24.059 | 25 |  | -0.041141 | -0.9967 | -0.95556 | -0.01455 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 310.8891 |  | -0.47745 | -11.7323 | -11.2549
1.75 | 30.0707 | 30.1187 | 30.283 | 31.25 |  | -0.047992 | -1.1793 | -1.1313 | -0.16433 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 373.0669 |  | -0.5456 | -13.549 | -13.0034
2 | 36.1381 | 36.1929 | 36.5792 | 37.5 |  | -0.054843 | -1.3619 | -1.3071 | -0.38628 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 404.1558 |  | -0.57968 | -14.4574 | -13.8777
2.125 | 39.1718 | 39.23 | 39.7544 | 40.625 |  | -0.058268 | -1.4532 | -1.395 | -0.52431 | 9.8066 | 1025 | 9.9485e-05 | 9.9485

## BED 4
2015-09-29
Pressure sensor offset calibration | [13 empty] | Forward calculation for External 3-wire Pressure Sensor 0-70 BAR Absolute, Ratiometric, 0.5-4.5V Output; US331-000004-070BA
| V sensor [V] (voltage measured at sensor output, 0-5V range) | [8 empty] | Note | Sensor SN | Sensor model |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
Pressure applied to sensor in addition to atmospheric pressure [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 70 | 0.5 | 4.5 | -8.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -96.9974
BED/Sensor Number (Volts) | [14 empty] | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 70 | 0.5 | 4.5 | 0 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -9.9485
BED 3 | [11 empty] | PHE163 100 bar |  |  | 65535 | 11.14% | 7302.472 | 2.5 | 0.27857 | 0.55714 | 0 | 70 | 0.5 | 4.5 | 1 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 7.4613e-06
BED 4 | 0.4935 | 0.7 | 0.9 | 1.1 | [7 empty] | PHE163 100 bar |  |  | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 70 | 0.5 | 4.5 | 13.125 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 120.625
BED 4 | 0.495 | 0.694 | 0.901 | 1.109 | 1.3175 | 1.517 | 1.718 | 1.912 | 2.106 | DVK 9/23/16, fluke 75 & 189 | 20298002 | PHE163 100 bar |  |  | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 70 | 0.5 | 4.5 | 35 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 338.2473
BED 5 | [11 empty] | PHE163 100 bar |  |  | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 70 | 0.5 | 4.5 | 56.875 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 555.8697
BED 6 | [11 empty] | PHE163 100 bar |  |  | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 70 | 0.5 | 4.5 | 70 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 686.4431
[15 empty] | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 70 | 0.5 | 4.5 | 78.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 773.4921
BEDS 2 | [9 empty] | DVK 10/8/2015 |  | PHE163 100 bar
BED 7 | [9 empty] | DVK 4/4/2016 | 21110001 | PHE163 100 bar
BED 8 | [9 empty] | DVK 4/4/2016 | 21110002 | PHE163 100 bar |  |  | Forward calculation for External 3-wire Pressure Sensor 0-100 bar Sealed Gauge, Ratiometric, 0.5-4.5V Output; PHE163 100 bar
BED 9 | [9 empty] | DVK 4/4/2016 | 21110003 | PHE163 100 bar |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
BED 10 | [9 empty] | DVK 4/4/2016 | 21110004 | PHE163 100 bar |  |  | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 100 | 0.5 | 4.5 | -12.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -124.3556
BED 11 | [9 empty] | DVK 4/4/2016 | 21110005 | PHE163 100 bar |  |  | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 100 | 0.5 | 4.5 | 0 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 0
Pressure on sensor + P atmosphere [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 11.14% | 7302.4713 | 2.5 | 0.27857 | 0.55714 | 0 | 100 | 0.5 | 4.5 | 1.4286 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 14.2121
[15 empty] | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 100 | 0.5 | 4.5 | 18.75 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 186.5335
[15 empty] | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 100 | 0.5 | 4.5 | 50 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 497.4225
[15 empty] | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 100 | 0.5 | 4.5 | 81.25 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 808.3116
[13 empty] | Legend: |  | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 100 | 0.5 | 4.5 | 100 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 994.8451
volts = 2.50 * (counts/65536) | [12 empty] | constant |  | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 100 | 0.5 | 4.5 | 112.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 1119.2007
(pressFullScale/2) * volts - pressFullScale/8 | [12 empty] | auto calculated
for 100 bar sensor: | [12 empty] | input |  | Forward calculation - Equations, plus simplified equations for PHE163 100 bar sensor
50.0 * volts - 12.5 | [14 empty] | Constant = 2^16-1 | Input | Counts FS * Counts%FS | Constant | Vref * Counts / Counts FS | Vadc * 2 | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | (Pmax-Pmin) / (Vsens max-Vsens min) * (Vsens- Vmin) + Pmin | Constant (sensor parameter) | Constant (atmospheric pressure) | Constant (gravity) | Constant (water density) | 1/ g [m/s2] / R [kg/m3] | k [m/Pa] * 1E5 [Pa/bar] | k [m/bar] * (Pcalc [bar] + Pref [bar] - Patm [bar])
[15 empty] | 65535 | input | 65535 * input | 2.5 | 2.5 * counts / 65536 | 5 * counts /65536 | 0 | 100 | 0.5 | 4.5 | Pmax / 4 * Vsens - Pmax /8 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 9.948* Pcalc [bar]
Calibration Vsens vs Pressure
Linear 4 point |  | Linear 9 point |  | polynomial 4 point |  | polynomial 9 point |  | Nominal
slope | 24.7571 |  | 24.7079 | a | 0.98642 | a | 0.57435 | slope | 25
intercept | -12.2654 |  | -12.3111 | b | 23.1859 | b | 23.214 | Intercept | -12.5
|  |  |  | c | -11.69 | c | -11.4965
R2 | 0.99994 |  | 0.99985 |  | 0.99999 |  | 0.99999
Water depth differences according to different calibrations, BED# 4
[15 empty] | Nominal Depth, m |  | Water depth difference meters, 9 points Lineal- 9 point polynomial | Water depth difference, meters 9 points Lineal -nominal | Water depth difference, meters 9 point Polynomial - Nominal | Water depth difference, 4 point Polinomial - 9 point polynomial
| Pressure,Bar, 9 point Linear calibration | Pressure, Bars, 4 point Polynomial calibration | Pressure, Bars, 9 point Polynomial calibration | Pressure, Bars Nominal calibration |  | Delta pressure, bars, 9 points Linear - 9 point Polinomial | Delta pressure, bars, 9 point Linear - Nominal | Delta pressure, bars, 9 point Polynomial - Nominal | Delta pressure, bars, 4 point Polinomial - 9 point polynomial | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar]
Volts | [14 empty] | 0 |  | -2.1014 | 0.42671 | 2.5281 | -1.0404
0.5 | 0.042892 | 0.14954 | 0.25412 | 0 |  | -0.21123 | 0.042892 | 0.25412 | -0.10458 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 62.1778 |  | -0.17155 | -0.29982 | -0.12827 | 0.17068
0.75 | 6.2199 | 6.2543 | 6.2371 | 6.25 |  | -0.017244 | -0.030137 | -0.012893 | 0.017157 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 124.3556 |  | 1.044 | -1.0263 | -2.0704 | 1.8942
1 | 12.3968 | 12.4823 | 12.2919 | 12.5 |  | 0.10494 | -0.10317 | -0.20811 | 0.1904 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 186.5335 |  | 1.5454 | -1.7529 | -3.2982 | 4.1301
1.25 | 18.5738 | 18.8336 | 18.4185 | 18.75 |  | 0.15534 | -0.17619 | -0.33153 | 0.41515 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 248.7113 |  | 1.3325 | -2.4794 | -3.8119 | 6.8785
1.5 | 24.7508 | 25.3083 | 24.6168 | 25 |  | 0.13394 | -0.24922 | -0.38316 | 0.69141 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 310.8891 |  | 0.40537 | -3.2059 | -3.6113 | 10.1393
1.75 | 30.9277 | 31.9062 | 30.887 | 31.25 |  | 0.040747 | -0.32225 | -0.363 | 1.0192 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 373.0669 |  | -1.236 | -3.9324 | -2.6964 | 13.9125
2 | 37.1047 | 38.6274 | 37.229 | 37.5 |  | -0.12424 | -0.39528 | -0.27104 | 1.3985 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 404.1558 |  | -2.3245 | -4.2957 | -1.9712 | 15.9913
2.125 | 40.1932 | 42.0343 | 40.4269 | 40.625 |  | -0.23366 | -0.43179 | -0.19814 | 1.6074 | 9.8066 | 1025 | 9.9485e-05 | 9.9485

## BED 5
2015-09-29
Pressure sensor offset calibration | [13 empty] | Forward calculation for External 3-wire Pressure Sensor 0-70 BAR Absolute, Ratiometric, 0.5-4.5V Output; US331-000004-070BA
| V sensor [V] (voltage measured at sensor output, 0-5V range) | [8 empty] | Note | Sensor SN | Sensor model |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
Pressure applied to sensor in addition to atmospheric pressure [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 70 | 0.5 | 4.5 | -8.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -96.9974
BED/Sensor Number (Volts) | [14 empty] | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 70 | 0.5 | 4.5 | 0 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -9.9485
BED 3 | [11 empty] | PHE163 100 bar |  |  | 65535 | 11.14% | 7302.472 | 2.5 | 0.27857 | 0.55714 | 0 | 70 | 0.5 | 4.5 | 1 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 7.4613e-06
BED 4 | [11 empty] | PHE163 100 bar |  |  | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 70 | 0.5 | 4.5 | 13.125 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 120.625
BED 5 | 0.528 | 0.734 | 0.94 | 1.15 | [7 empty] | PHE163 100 bar |  |  | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 70 | 0.5 | 4.5 | 35 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 338.2473
BED 6 | [11 empty] | PHE163 100 bar |  |  | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 70 | 0.5 | 4.5 | 56.875 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 555.8697
[15 empty] | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 70 | 0.5 | 4.5 | 70 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 686.4431
BEDS 2 | [9 empty] | DVK 10/8/2015 |  | PHE163 100 bar |  |  | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 70 | 0.5 | 4.5 | 78.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 773.4921
BED 7 | [9 empty] | DVK 4/4/2016 | 21110001 | PHE163 100 bar
BED 8 | [9 empty] | DVK 4/4/2016 | 21110002 | PHE163 100 bar
BED 9 | [9 empty] | DVK 4/4/2016 | 21110003 | PHE163 100 bar |  |  | Forward calculation for External 3-wire Pressure Sensor 0-100 bar Sealed Gauge, Ratiometric, 0.5-4.5V Output; PHE163 100 bar
BED 10 | [9 empty] | DVK 4/4/2016 | 21110004 | PHE163 100 bar |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
BED 11 | [9 empty] | DVK 4/4/2016 | 21110005 | PHE163 100 bar |  |  | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 100 | 0.5 | 4.5 | -12.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -124.3556
Pressure on sensor + P atmosphere [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 100 | 0.5 | 4.5 | 0 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 0
[15 empty] | 65535 | 11.14% | 7302.4713 | 2.5 | 0.27857 | 0.55714 | 0 | 100 | 0.5 | 4.5 | 1.4286 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 14.2121
[15 empty] | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 100 | 0.5 | 4.5 | 18.75 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 186.5335
[15 empty] | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 100 | 0.5 | 4.5 | 50 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 497.4225
[15 empty] | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 100 | 0.5 | 4.5 | 81.25 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 808.3116
volts = 2.50 * (counts/65536) | [12 empty] | Legend: |  | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 100 | 0.5 | 4.5 | 100 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 994.8451
(pressFullScale/2) * volts - pressFullScale/8 | [12 empty] | constant |  | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 100 | 0.5 | 4.5 | 112.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 1119.2007
for 100 bar sensor: | [12 empty] | auto calculated
50.0 * volts - 12.5 | [12 empty] | input |  | Forward calculation - Equations, plus simplified equations for PHE163 100 bar sensor
[15 empty] | Constant = 2^16-1 | Input | Counts FS * Counts%FS | Constant | Vref * Counts / Counts FS | Vadc * 2 | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | (Pmax-Pmin) / (Vsens max-Vsens min) * (Vsens- Vmin) + Pmin | Constant (sensor parameter) | Constant (atmospheric pressure) | Constant (gravity) | Constant (water density) | 1/ g [m/s2] / R [kg/m3] | k [m/Pa] * 1E5 [Pa/bar] | k [m/bar] * (Pcalc [bar] + Pref [bar] - Patm [bar])
[15 empty] | 65535 | input | 65535 * input | 2.5 | 2.5 * counts / 65536 | 5 * counts /65536 | 0 | 100 | 0.5 | 4.5 | Pmax / 4 * Vsens - Pmax /8 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 9.948* Pcalc [bar]
Calibration Vsens vs Pressure
Linear |  | polynomial |  | Nominal
slope | 24.1307 | a | -0.5621 | slope | 25
intercept | -12.7216 | b | 25.074 | Intercept | -12.5
|  | c | -13.087
R2 | 0.99998 |  | 0.99994
Water depth differences according to different calibrations, BED# 5
| Pressure,Bars Linear calibration | Pressure, Bars Polynomial calibration | Pressure, Bars Nominal calibration |  | Delta pressure, bars, Linear - Polinomial | Delta pressure, bars, Linear - Nominal | Delta pressure, bars, Polynomial - Nominal | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] |  | Nominal Depth, m |  | Water depth difference meters, Lineal- polynomial | Water depth difference, meters Lineal -nominal | Water depth difference ,meters Polynomial - Nominal
Volts
0.5 | -0.65619 | -0.69052 | 0 |  | 0.034337 | -0.65619 | -0.69052 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 0 |  | 0.3416 | -6.5281 | -6.8697
0.75 | 5.3765 | 5.4023 | 6.25 |  | -0.025823 | -0.8735 | -0.84768 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 62.1778 |  | -0.2569 | -8.69 | -8.4331
1 | 11.4092 | 11.4249 | 12.5 |  | -0.015721 | -1.0908 | -1.0751 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 124.3556 |  | -0.1564 | -10.852 | -10.6956
1.25 | 17.4419 | 17.3772 | 18.75 |  | 0.064644 | -1.3081 | -1.3728 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 186.5335 |  | 0.64311 | -13.0139 | -13.657
1.5 | 23.4745 | 23.2593 | 25 |  | 0.21527 | -1.5255 | -1.7407 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 248.7113 |  | 2.1416 | -15.1759 | -17.3175
1.75 | 29.5072 | 29.0711 | 31.25 |  | 0.43616 | -1.7428 | -2.1789 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 310.8891 |  | 4.3391 | -17.3379 | -21.677
2 | 35.5399 | 34.8126 | 37.5 |  | 0.72731 | -1.9601 | -2.6874 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 373.0669 |  | 7.2356 | -19.4998 | -26.7355
2.125 | 38.5563 | 37.657 | 40.625 |  | 0.89924 | -2.0687 | -2.968 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 404.1558 |  | 8.946 | -20.5808 | -29.5268

## BED 6
2015-09-29
Pressure sensor offset calibration | [13 empty] | Forward calculation for External 3-wire Pressure Sensor 0-70 BAR Absolute, Ratiometric, 0.5-4.5V Output; US331-000004-070BA
| V sensor [V] (voltage measured at sensor output, 0-5V range) | [8 empty] | Note | Sensor SN | Sensor model |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
Pressure applied to sensor in addition to atmospheric pressure [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 70 | 0.5 | 4.5 | -8.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -96.9974
BED/Sensor Number (Volts) | [14 empty] | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 70 | 0.5 | 4.5 | 0 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -9.9485
BED 3 | [11 empty] | PHE163 100 bar |  |  | 65535 | 11.14% | 7302.472 | 2.5 | 0.27857 | 0.55714 | 0 | 70 | 0.5 | 4.5 | 1 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 7.4613e-06
BED 4 | [11 empty] | PHE163 100 bar |  |  | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 70 | 0.5 | 4.5 | 13.125 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 120.625
BED 5 | [11 empty] | PHE163 100 bar |  |  | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 70 | 0.5 | 4.5 | 35 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 338.2473
BED 6 | 0.505 | 0.71 | 0.918 | 1.113 | 1.33 | 1.53 | 1.73 | 1.93 | 2.12 |  |  | PHE163 100 bar |  |  | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 70 | 0.5 | 4.5 | 56.875 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 555.8697
[15 empty] | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 70 | 0.5 | 4.5 | 70 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 686.4431
BEDS 2 | [9 empty] | DVK 10/8/2015 |  | PHE163 100 bar |  |  | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 70 | 0.5 | 4.5 | 78.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 773.4921
BED 7 | [9 empty] | DVK 4/4/2016 | 21110001 | PHE163 100 bar
BED 8 | [9 empty] | DVK 4/4/2016 | 21110002 | PHE163 100 bar
BED 9 | [9 empty] | DVK 4/4/2016 | 21110003 | PHE163 100 bar |  |  | Forward calculation for External 3-wire Pressure Sensor 0-100 bar Sealed Gauge, Ratiometric, 0.5-4.5V Output; PHE163 100 bar
BED 10 | [9 empty] | DVK 4/4/2016 | 21110004 | PHE163 100 bar |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
BED 11 | [9 empty] | DVK 4/4/2016 | 21110005 | PHE163 100 bar |  |  | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 100 | 0.5 | 4.5 | -12.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -124.3556
Pressure on sensor + P atmosphere [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 100 | 0.5 | 4.5 | 0 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 0
[15 empty] | 65535 | 11.14% | 7302.4713 | 2.5 | 0.27857 | 0.55714 | 0 | 100 | 0.5 | 4.5 | 1.4286 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 14.2121
[15 empty] | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 100 | 0.5 | 4.5 | 18.75 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 186.5335
[15 empty] | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 100 | 0.5 | 4.5 | 50 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 497.4225
[15 empty] | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 100 | 0.5 | 4.5 | 81.25 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 808.3116
volts = 2.50 * (counts/65536) | [12 empty] | Legend: |  | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 100 | 0.5 | 4.5 | 100 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 994.8451
(pressFullScale/2) * volts - pressFullScale/8 | [12 empty] | constant |  | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 100 | 0.5 | 4.5 | 112.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 1119.2007
for 100 bar sensor: | [12 empty] | auto calculated
50.0 * volts - 12.5 | [12 empty] | input |  | Forward calculation - Equations, plus simplified equations for PHE163 100 bar sensor
[15 empty] | Constant = 2^16-1 | Input | Counts FS * Counts%FS | Constant | Vref * Counts / Counts FS | Vadc * 2 | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | (Pmax-Pmin) / (Vsens max-Vsens min) * (Vsens- Vmin) + Pmin | Constant (sensor parameter) | Constant (atmospheric pressure) | Constant (gravity) | Constant (water density) | 1/ g [m/s2] / R [kg/m3] | k [m/Pa] * 1E5 [Pa/bar] | k [m/bar] * (Pcalc [bar] + Pref [bar] - Patm [bar])
[15 empty] | 65535 | input | 65535 * input | 2.5 | 2.5 * counts / 65536 | 5 * counts /65536 | 0 | 100 | 0.5 | 4.5 | Pmax / 4 * Vsens - Pmax /8 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 9.948* Pcalc [bar]
Calibration Vsens vs Pressure
Linear |  | polynomial |  | Nominal
slope | 24.6656 | a | 0.4777 | slope | 25
intercept | -12.575 | b | 23.14 | Intercept | -12.5
|  | c | -11.88
R2 | 0.99986 |  | 0.99994
Water depth differences according to different calibrations
| Pressure,Bars Linear calibration | Pressure, Bars Polynomial calibration | Pressure, Bars Nominal calibration |  | Delta pressure, bars, Linear - Polinomial | Delta pressure, bars, Linear - Nominal | Delta pressure, bars, Polynomial - Nominal | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] |  | Nominal Depth, m |  | Water depth difference meters, Lineal- polynomial | Water depth difference, meters Lineal -nominal | Water depth difference ,meters Polynomial - Nominal
Volts
0.5 | -0.24224 | -0.19058 | 0 |  | -0.051662 | -0.24224 | -0.19058 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 0 |  | -0.51396 | -2.4099 | -1.8959
0.75 | 5.9242 | 5.7437 | 6.25 |  | 0.18046 | -0.32584 | -0.50629 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 62.1778 |  | 1.7953 | -3.2416 | -5.0368
1 | 12.0906 | 11.7377 | 12.5 |  | 0.35286 | -0.40944 | -0.7623 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 124.3556 |  | 3.5104 | -4.0733 | -7.5837
1.25 | 18.257 | 17.7914 | 18.75 |  | 0.46556 | -0.49304 | -0.95859 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 186.5335 |  | 4.6316 | -4.9049 | -9.5365
1.5 | 24.4234 | 23.9048 | 25 |  | 0.51854 | -0.57664 | -1.0952 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 248.7113 |  | 5.1587 | -5.7366 | -10.8953
1.75 | 30.5898 | 30.078 | 31.25 |  | 0.51181 | -0.66024 | -1.172 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 310.8891 |  | 5.0917 | -6.5683 | -11.66
2 | 36.7562 | 36.3108 | 37.5 |  | 0.44537 | -0.74383 | -1.1892 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 373.0669 |  | 4.4307 | -7.4 | -11.8307
2.125 | 39.8394 | 39.4496 | 40.625 |  | 0.38975 | -0.78563 | -1.1754 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 404.1558 |  | 3.8774 | -7.8158 | -11.6933

## BED 2
2015-09-29
Pressure sensor offset calibration | [13 empty] | Forward calculation for External 3-wire Pressure Sensor 0-70 BAR Absolute, Ratiometric, 0.5-4.5V Output; US331-000004-070BA
| V sensor [V] (voltage measured at sensor output, 0-5V range) | [8 empty] | Note | Sensor SN | Sensor model |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
Pressure applied to sensor in addition to atmospheric pressure [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 70 | 0.5 | 4.5 | -8.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -96.9974
BED/Sensor Number (Volts) | [14 empty] | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 70 | 0.5 | 4.5 | 0 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -9.9485
BED 3 | [11 empty] | PHE163 100 bar |  |  | 65535 | 11.14% | 7302.472 | 2.5 | 0.27857 | 0.55714 | 0 | 70 | 0.5 | 4.5 | 1 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 7.4613e-06
BED 4 | [11 empty] | PHE163 100 bar |  |  | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 70 | 0.5 | 4.5 | 13.125 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 120.625
BED 5 | [11 empty] | PHE163 100 bar |  |  | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 70 | 0.5 | 4.5 | 35 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 338.2473
BED 6 | [11 empty] | PHE163 100 bar |  |  | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 70 | 0.5 | 4.5 | 56.875 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 555.8697
[15 empty] | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 70 | 0.5 | 4.5 | 70 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 686.4431
BEDS 2 | 0.484 | 0.684 | 0.892 | 1.095 | 1.302 | 1.502 | 1.703 | 1.895 | 2.09 | DVK 10/8/2015 |  | PHE163 100 bar |  |  | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 70 | 0.5 | 4.5 | 78.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 773.4921
BED 7 | [9 empty] | DVK 4/4/2016 | 21110001 | PHE163 100 bar
BED 8 | [9 empty] | DVK 4/4/2016 | 21110002 | PHE163 100 bar
BED 9 | [9 empty] | DVK 4/4/2016 | 21110003 | PHE163 100 bar |  |  | Forward calculation for External 3-wire Pressure Sensor 0-100 bar Sealed Gauge, Ratiometric, 0.5-4.5V Output; PHE163 100 bar
BED 10 | [9 empty] | DVK 4/4/2016 | 21110004 | PHE163 100 bar |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
BED 11 | [9 empty] | DVK 4/4/2016 | 21110005 | PHE163 100 bar |  |  | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 100 | 0.5 | 4.5 | -12.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -124.3556
Pressure on sensor + P atmosphere [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 100 | 0.5 | 4.5 | 0 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 0
[15 empty] | 65535 | 11.14% | 7302.4713 | 2.5 | 0.27857 | 0.55714 | 0 | 100 | 0.5 | 4.5 | 1.4286 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 14.2121
[15 empty] | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 100 | 0.5 | 4.5 | 18.75 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 186.5335
[15 empty] | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 100 | 0.5 | 4.5 | 50 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 497.4225
[15 empty] | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 100 | 0.5 | 4.5 | 81.25 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 808.3116
volts = 2.50 * (counts/65536) | [12 empty] | Legend: |  | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 100 | 0.5 | 4.5 | 100 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 994.8451
(pressFullScale/2) * volts - pressFullScale/8 | [12 empty] | constant |  | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 100 | 0.5 | 4.5 | 112.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 1119.2007
for 100 bar sensor: | [12 empty] | auto calculated
50.0 * volts - 12.5 | [12 empty] | input |  | Forward calculation - Equations, plus simplified equations for PHE163 100 bar sensor
[15 empty] | Constant = 2^16-1 | Input | Counts FS * Counts%FS | Constant | Vref * Counts / Counts FS | Vadc * 2 | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | (Pmax-Pmin) / (Vsens max-Vsens min) * (Vsens- Vmin) + Pmin | Constant (sensor parameter) | Constant (atmospheric pressure) | Constant (gravity) | Constant (water density) | 1/ g [m/s2] / R [kg/m3] | k [m/Pa] * 1E5 [Pa/bar] | k [m/bar] * (Pcalc [bar] + Pref [bar] - Patm [bar])
[15 empty] | 65535 | input | 65535 * input | 2.5 | 2.5 * counts / 65536 | 5 * counts /65536 | 0 | 100 | 0.5 | 4.5 | Pmax / 4 * Vsens - Pmax /8 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 9.948* Pcalc [bar]
Calibration Vsens vs Pressure
Linear |  | polynomial |  | Nominal
slope | 24.819 | a | 0.5469 | slope | 25
intercept | -12.1185 | b | 23.411 | Intercept | -12.5
|  | c | -11.36
R2 | 0.99987 |  | 0.99997
Water depth differences according to different calibrations, BED #2
| Pressure,Bars Linear calibration | Pressure, Bars Polynomial calibration | Pressure, Bars Nominal calibration |  | Delta pressure, bars, Linear - Polinomial | Delta pressure, bars, Linear - Nominal | Delta pressure, bars, Polynomial - Nominal | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] |  | Nominal Depth, m |  | Water depth difference meters, Lineal- polynomial | Water depth difference, meters Lineal -nominal | Water depth difference ,meters Polynomial - Nominal
Volts
0.5 | 0.29098 | 0.48223 | 0 |  | -0.19125 | 0.29098 | 0.48223 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 0 |  | -1.9026 | 2.8948 | 4.7974
0.75 | 6.4957 | 6.5059 | 6.25 |  | -0.01016 | 0.24572 | 0.25588 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 62.1778 |  | -0.10107 | 2.4445 | 2.5456
1 | 12.7005 | 12.5979 | 12.5 |  | 0.10256 | 0.20046 | 0.0979 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 124.3556 |  | 1.0204 | 1.9943 | 0.97395
1.25 | 18.9052 | 18.7583 | 18.75 |  | 0.14693 | 0.15521 | 0.0082813 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 186.5335 |  | 1.4617 | 1.5441 | 0.082386
1.5 | 25.11 | 24.987 | 25 |  | 0.12293 | 0.10995 | -0.012975 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 248.7113 |  | 1.2229 | 1.0938 | -0.12908
1.75 | 31.3147 | 31.2841 | 31.25 |  | 0.030562 | 0.064694 | 0.034131 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 310.8891 |  | 0.30405 | 0.6436 | 0.33955
2 | 37.5194 | 37.6496 | 37.5 |  | -0.13016 | 0.019437 | 0.1496 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 373.0669 |  | -1.2949 | 0.19336 | 1.4883
2.125 | 40.6218 | 40.858 | 40.625 |  | -0.23616 | -0.0031918 | 0.23297 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 404.1558 |  | -2.3494 | -0.031754 | 2.3177

## BED 7
2015-09-29
Pressure sensor offset calibration | [13 empty] | Forward calculation for External 3-wire Pressure Sensor 0-70 BAR Absolute, Ratiometric, 0.5-4.5V Output; US331-000004-070BA
| V sensor [V] (voltage measured at sensor output, 0-5V range) | [8 empty] | Note | Sensor SN | Sensor model |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
Pressure applied to sensor in addition to atmospheric pressure [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 70 | 0.5 | 4.5 | -8.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -96.9974
BED/Sensor Number (Volts) | [14 empty] | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 70 | 0.5 | 4.5 | 0 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -9.9485
BED 3 | [11 empty] | PHE163 100 bar |  |  | 65535 | 11.14% | 7302.472 | 2.5 | 0.27857 | 0.55714 | 0 | 70 | 0.5 | 4.5 | 1 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 7.4613e-06
BED 4 | [11 empty] | PHE163 100 bar |  |  | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 70 | 0.5 | 4.5 | 13.125 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 120.625
BED 5 | [11 empty] | PHE163 100 bar |  |  | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 70 | 0.5 | 4.5 | 35 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 338.2473
BED 6 | [11 empty] | PHE163 100 bar |  |  | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 70 | 0.5 | 4.5 | 56.875 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 555.8697
[15 empty] | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 70 | 0.5 | 4.5 | 70 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 686.4431
BEDS 2 | [9 empty] | DVK 10/8/2015 |  | PHE163 100 bar |  |  | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 70 | 0.5 | 4.5 | 78.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 773.4921
BED 7 | 0.497 | 0.694 | 0.899 | 1.106 | 1.312 | 1.512 | 1.711 | 1.907 | 2.1 | DVK 4/4/2016 | 21110001 | PHE163 100 bar
BED 8 | [9 empty] | DVK 4/4/2016 | 21110002 | PHE163 100 bar
BED 9 | [9 empty] | DVK 4/4/2016 | 21110003 | PHE163 100 bar |  |  | Forward calculation for External 3-wire Pressure Sensor 0-100 bar Sealed Gauge, Ratiometric, 0.5-4.5V Output; PHE163 100 bar
BED 10 | [9 empty] | DVK 4/4/2016 | 21110004 | PHE163 100 bar |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
BED 11 | [9 empty] | DVK 4/4/2016 | 21110005 | PHE163 100 bar |  |  | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 100 | 0.5 | 4.5 | -12.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -124.3556
Pressure on sensor + P atmosphere [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 100 | 0.5 | 4.5 | 0 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 0
[15 empty] | 65535 | 11.14% | 7302.4713 | 2.5 | 0.27857 | 0.55714 | 0 | 100 | 0.5 | 4.5 | 1.4286 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 14.2121
[15 empty] | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 100 | 0.5 | 4.5 | 18.75 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 186.5335
[15 empty] | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 100 | 0.5 | 4.5 | 50 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 497.4225
[15 empty] | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 100 | 0.5 | 4.5 | 81.25 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 808.3116
volts = 2.50 * (counts/65536) | [12 empty] | Legend: |  | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 100 | 0.5 | 4.5 | 100 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 994.8451
(pressFullScale/2) * volts - pressFullScale/8 | [12 empty] | constant |  | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 100 | 0.5 | 4.5 | 112.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 1119.2007
for 100 bar sensor: | [12 empty] | auto calculated
50.0 * volts - 12.5 | [12 empty] | input |  | Forward calculation - Equations, plus simplified equations for PHE163 100 bar sensor
[15 empty] | Constant = 2^16-1 | Input | Counts FS * Counts%FS | Constant | Vref * Counts / Counts FS | Vadc * 2 | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | (Pmax-Pmin) / (Vsens max-Vsens min) * (Vsens- Vmin) + Pmin | Constant (sensor parameter) | Constant (atmospheric pressure) | Constant (gravity) | Constant (water density) | 1/ g [m/s2] / R [kg/m3] | k [m/Pa] * 1E5 [Pa/bar] | k [m/bar] * (Pcalc [bar] + Pref [bar] - Patm [bar])
[15 empty] | 65535 | input | 65535 * input | 2.5 | 2.5 * counts / 65536 | 5 * counts /65536 | 0 | 100 | 0.5 | 4.5 | Pmax / 4 * Vsens - Pmax /8 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 9.948* Pcalc [bar]
Calibration Vsens vs Pressure
Linear |  | polynomial |  | Nominal
slope | 24.8296 | a | 0.4693 | slope | 25
intercept | -12.3833 | b | 23.611 | Intercept | -12.5
|  | c | -11.719
R2 | 0.99989 |  | 0.99996
Water depth differences according to different calibrations, BED #7
| Pressure,Bars Linear calibration | Pressure, Bars Polynomial calibration | Pressure, Bars Nominal calibration |  | Delta pressure, bars, Linear - Polinomial | Delta pressure, bars, Linear - Nominal | Delta pressure, bars, Polynomial - Nominal | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] |  | Nominal Depth, m |  | Water depth difference meters, Lineal- polynomial | Water depth difference, meters Lineal -nominal | Water depth difference ,meters Polynomial - Nominal
Volts
0.5 | 0.031465 | 0.20383 | 0 |  | -0.17236 | 0.031465 | 0.20383 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 0 |  | -1.7147 | 0.31303 | 2.0277
0.75 | 6.2389 | 6.2532 | 6.25 |  | -0.01436 | -0.011129 | 0.0032313 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 62.1778 |  | -0.14286 | -0.11072 | 0.032146
1 | 12.4463 | 12.3613 | 12.5 |  | 0.084977 | -0.053723 | -0.1387 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 124.3556 |  | 0.84539 | -0.53446 | -1.3799
1.25 | 18.6537 | 18.528 | 18.75 |  | 0.12565 | -0.096317 | -0.22197 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 186.5335 |  | 1.25 | -0.95821 | -2.2082
1.5 | 24.8611 | 24.7534 | 25 |  | 0.10766 | -0.13891 | -0.24657 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 248.7113 |  | 1.0711 | -1.382 | -2.453
1.75 | 31.0685 | 31.0375 | 31.25 |  | 0.031013 | -0.18151 | -0.21252 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 310.8891 |  | 0.30853 | -1.8057 | -2.1142
2 | 37.2759 | 37.3802 | 37.5 |  | -0.1043 | -0.2241 | -0.1198 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 373.0669 |  | -1.0376 | -2.2294 | -1.1918
2.125 | 40.3796 | 40.5736 | 40.625 |  | -0.19395 | -0.2454 | -0.051442 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 404.1558 |  | -1.9295 | -2.4413 | -0.51177

## BED 8
2015-09-29
Pressure sensor offset calibration | [13 empty] | Forward calculation for External 3-wire Pressure Sensor 0-70 BAR Absolute, Ratiometric, 0.5-4.5V Output; US331-000004-070BA
| V sensor [V] (voltage measured at sensor output, 0-5V range) | [8 empty] | Note | Sensor SN | Sensor model |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
Pressure applied to sensor in addition to atmospheric pressure [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 70 | 0.5 | 4.5 | -8.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -96.9974
BED/Sensor Number (Volts) | [14 empty] | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 70 | 0.5 | 4.5 | 0 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -9.9485
BED 3 | [11 empty] | PHE163 100 bar |  |  | 65535 | 11.14% | 7302.472 | 2.5 | 0.27857 | 0.55714 | 0 | 70 | 0.5 | 4.5 | 1 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 7.4613e-06
BED 4 | [11 empty] | PHE163 100 bar |  |  | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 70 | 0.5 | 4.5 | 13.125 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 120.625
BED 5 | [11 empty] | PHE163 100 bar |  |  | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 70 | 0.5 | 4.5 | 35 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 338.2473
BED 6 | [11 empty] | PHE163 100 bar |  |  | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 70 | 0.5 | 4.5 | 56.875 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 555.8697
[15 empty] | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 70 | 0.5 | 4.5 | 70 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 686.4431
BEDS 2 | [9 empty] | DVK 10/8/2015 |  | PHE163 100 bar |  |  | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 70 | 0.5 | 4.5 | 78.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 773.4921
BED 7 | [9 empty] | DVK 4/4/2016 | 21110001 | PHE163 100 bar
BED 8 | 0.492 | 0.691 | 0.899 | 1.108 | 1.316 | 1.517 | 1.718 | 1.913 | 2.108 | DVK 4/4/2016 | 21110002 | PHE163 100 bar
BED 9 | [9 empty] | DVK 4/4/2016 | 21110003 | PHE163 100 bar |  |  | Forward calculation for External 3-wire Pressure Sensor 0-100 bar Sealed Gauge, Ratiometric, 0.5-4.5V Output; PHE163 100 bar
BED 10 | [9 empty] | DVK 4/4/2016 | 21110004 | PHE163 100 bar |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
BED 11 | [9 empty] | DVK 4/4/2016 | 21110005 | PHE163 100 bar |  |  | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 100 | 0.5 | 4.5 | -12.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -124.3556
Pressure on sensor + P atmosphere [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 100 | 0.5 | 4.5 | 0 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 0
[15 empty] | 65535 | 11.14% | 7302.4713 | 2.5 | 0.27857 | 0.55714 | 0 | 100 | 0.5 | 4.5 | 1.4286 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 14.2121
[15 empty] | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 100 | 0.5 | 4.5 | 18.75 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 186.5335
[15 empty] | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 100 | 0.5 | 4.5 | 50 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 497.4225
[15 empty] | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 100 | 0.5 | 4.5 | 81.25 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 808.3116
volts = 2.50 * (counts/65536) | [12 empty] | Legend: |  | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 100 | 0.5 | 4.5 | 100 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 994.8451
(pressFullScale/2) * volts - pressFullScale/8 | [12 empty] | constant |  | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 100 | 0.5 | 4.5 | 112.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 1119.2007
for 100 bar sensor: | [12 empty] | auto calculated
50.0 * volts - 12.5 | [12 empty] | input |  | Forward calculation - Equations, plus simplified equations for PHE163 100 bar sensor
[15 empty] | Constant = 2^16-1 | Input | Counts FS * Counts%FS | Constant | Vref * Counts / Counts FS | Vadc * 2 | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | (Pmax-Pmin) / (Vsens max-Vsens min) * (Vsens- Vmin) + Pmin | Constant (sensor parameter) | Constant (atmospheric pressure) | Constant (gravity) | Constant (water density) | 1/ g [m/s2] / R [kg/m3] | k [m/Pa] * 1E5 [Pa/bar] | k [m/bar] * (Pcalc [bar] + Pref [bar] - Patm [bar])
[15 empty] | 65535 | input | 65535 * input | 2.5 | 2.5 * counts / 65536 | 5 * counts /65536 | 0 | 100 | 0.5 | 4.5 | Pmax / 4 * Vsens - Pmax /8 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 9.948* Pcalc [bar]
Calibration Vsens vs Pressure
Linear |  | polynomial |  | Nominal
slope | 24.633 | a | 0.5558 | slope | 25
intercept | -12.1926 | b | 23.188 | Intercept | -12.5
|  | c | -11.406
R2 | 0.99985 |  | 0.99996
Water depth differences according to different calibrations, BED #8
| Pressure,Bars Linear calibration | Pressure, Bars Polynomial calibration | Pressure, Bars Nominal calibration |  | Delta pressure, bars, Linear - Polinomial | Delta pressure, bars, Linear - Nominal | Delta pressure, bars, Polynomial - Nominal | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] |  | Nominal Depth, m |  | Water depth difference meters, Lineal- polynomial | Water depth difference, meters Lineal -nominal | Water depth difference, meters Polynomial - Nominal
Volts
0.5 | 0.12391 | 0.32695 | 0 |  | -0.20304 | 0.12391 | 0.32695 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 0 |  | -2.0199 | 1.2327 | 3.2526
0.75 | 6.2822 | 6.2976 | 6.25 |  | -0.015478 | 0.032159 | 0.047637 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 62.1778 |  | -0.15399 | 0.31993 | 0.47392
1 | 12.4404 | 12.3378 | 12.5 |  | 0.10261 | -0.059592 | -0.1622 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 124.3556 |  | 1.0208 | -0.59285 | -1.6136
1.25 | 18.5987 | 18.4474 | 18.75 |  | 0.15122 | -0.15134 | -0.30256 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 186.5335 |  | 1.5044 | -1.5056 | -3.01
1.5 | 24.7569 | 24.6265 | 25 |  | 0.13035 | -0.2431 | -0.37345 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 248.7113 |  | 1.2968 | -2.4184 | -3.7152
1.75 | 30.9152 | 30.8751 | 31.25 |  | 0.040016 | -0.33485 | -0.37486 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 310.8891 |  | 0.3981 | -3.3312 | -3.7293
2 | 37.0734 | 37.1932 | 37.5 |  | -0.1198 | -0.4266 | -0.3068 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 373.0669 |  | -1.1918 | -4.244 | -3.0522
2.125 | 40.1525 | 40.3783 | 40.625 |  | -0.22576 | -0.47247 | -0.24672 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 404.1558 |  | -2.2459 | -4.7004 | -2.4544

## BED 9
2015-09-29
Pressure sensor offset calibration | [13 empty] | Forward calculation for External 3-wire Pressure Sensor 0-70 BAR Absolute, Ratiometric, 0.5-4.5V Output; US331-000004-070BA
| V sensor [V] (voltage measured at sensor output, 0-5V range) | [8 empty] | Note | Sensor SN | Sensor model |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
Pressure applied to sensor in addition to atmospheric pressure [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 70 | 0.5 | 4.5 | -8.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -96.9974
BED/Sensor Number (Volts) | [14 empty] | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 70 | 0.5 | 4.5 | 0 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -9.9485
BED 3 | [11 empty] | PHE163 100 bar |  |  | 65535 | 11.14% | 7302.472 | 2.5 | 0.27857 | 0.55714 | 0 | 70 | 0.5 | 4.5 | 1 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 7.4613e-06
BED 4 | [11 empty] | PHE163 100 bar |  |  | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 70 | 0.5 | 4.5 | 13.125 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 120.625
BED 5 | [11 empty] | PHE163 100 bar |  |  | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 70 | 0.5 | 4.5 | 35 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 338.2473
BED 6 | [11 empty] | PHE163 100 bar |  |  | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 70 | 0.5 | 4.5 | 56.875 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 555.8697
[15 empty] | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 70 | 0.5 | 4.5 | 70 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 686.4431
BEDS 2 | [9 empty] | DVK 10/8/2015 |  | PHE163 100 bar |  |  | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 70 | 0.5 | 4.5 | 78.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 773.4921
BED 7 | [9 empty] | DVK 4/4/2016 | 21110001 | PHE163 100 bar
BED 8 | [9 empty] | DVK 4/4/2016 | 21110002 | PHE163 100 bar
BED 9 | 0.508 | 0.712 | 0.917 | 1.124 | 1.333 | 1.534 | 1.733 | 1.933 | 2.124 | DVK 4/4/2016 | 21110003 | PHE163 100 bar |  |  | Forward calculation for External 3-wire Pressure Sensor 0-100 bar Sealed Gauge, Ratiometric, 0.5-4.5V Output; PHE163 100 bar
BED 10 | [9 empty] | DVK 4/4/2016 | 21110004 | PHE163 100 bar |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
BED 11 | [9 empty] | DVK 4/4/2016 | 21110005 | PHE163 100 bar |  |  | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 100 | 0.5 | 4.5 | -12.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -124.3556
Pressure on sensor + P atmosphere [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 100 | 0.5 | 4.5 | 0 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 0
[15 empty] | 65535 | 11.14% | 7302.4713 | 2.5 | 0.27857 | 0.55714 | 0 | 100 | 0.5 | 4.5 | 1.4286 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 14.2121
[15 empty] | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 100 | 0.5 | 4.5 | 18.75 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 186.5335
[15 empty] | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 100 | 0.5 | 4.5 | 50 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 497.4225
[15 empty] | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 100 | 0.5 | 4.5 | 81.25 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 808.3116
volts = 2.50 * (counts/65536) | [12 empty] | Legend: |  | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 100 | 0.5 | 4.5 | 100 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 994.8451
(pressFullScale/2) * volts - pressFullScale/8 | [12 empty] | constant |  | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 100 | 0.5 | 4.5 | 112.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 1119.2007
for 100 bar sensor: | [12 empty] | auto calculated
50.0 * volts - 12.5 | [12 empty] | input |  | Forward calculation - Equations, plus simplified equations for PHE163 100 bar sensor
[15 empty] | Constant = 2^16-1 | Input | Counts FS * Counts%FS | Constant | Vref * Counts / Counts FS | Vadc * 2 | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | (Pmax-Pmin) / (Vsens max-Vsens min) * (Vsens- Vmin) + Pmin | Constant (sensor parameter) | Constant (atmospheric pressure) | Constant (gravity) | Constant (water density) | 1/ g [m/s2] / R [kg/m3] | k [m/Pa] * 1E5 [Pa/bar] | k [m/bar] * (Pcalc [bar] + Pref [bar] - Patm [bar])
[15 empty] | 65535 | input | 65535 * input | 2.5 | 2.5 * counts / 65536 | 5 * counts /65536 | 0 | 100 | 0.5 | 4.5 | Pmax / 4 * Vsens - Pmax /8 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 9.948* Pcalc [bar]
Calibration Vsens vs Pressure
Linear |  | polynomial |  | Nominal
slope | 24.6497 | a | 0.5421 | slope | 25
intercept | -12.6416 | b | 23.221 | Intercept | -12.5
|  | c | -11.85
R2 | 0.99987 |  | 0.99997
Water depth differences according to different calibrations, BED #9
| Pressure,Bars Linear calibration | Pressure, Bars Polynomial calibration | Pressure, Bars Nominal calibration |  | Delta pressure, bars, Linear - Polinomial | Delta pressure, bars, Linear - Nominal | Delta pressure, bars, Polynomial - Nominal | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] |  | Nominal Depth, m |  | Water depth difference meters, Lineal- polynomial | Water depth difference, meters Lineal -nominal | Water depth difference, meters Polynomial - Nominal
Volts
0.5 | -0.31679 | -0.10398 | 0 |  | -0.21282 | -0.31679 | -0.10398 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 0 |  | -2.1172 | -3.1516 | -1.0344
0.75 | 5.8456 | 5.8707 | 6.25 |  | -0.025061 | -0.40438 | -0.37932 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 62.1778 |  | -0.24932 | -4.023 | -3.7736
1 | 12.008 | 11.9131 | 12.5 |  | 0.094934 | -0.49197 | -0.5869 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 124.3556 |  | 0.94445 | -4.8943 | -5.8387
1.25 | 18.1704 | 18.0233 | 18.75 |  | 0.14717 | -0.57955 | -0.72672 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 186.5335 |  | 1.4641 | -5.7656 | -7.2297
1.5 | 24.3329 | 24.2012 | 25 |  | 0.13164 | -0.66714 | -0.79878 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 248.7113 |  | 1.3096 | -6.637 | -7.9466
1.75 | 30.4953 | 30.4469 | 31.25 |  | 0.048344 | -0.75472 | -0.80307 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 310.8891 |  | 0.48095 | -7.5083 | -7.9893
2 | 36.6577 | 36.7604 | 37.5 |  | -0.10271 | -0.84231 | -0.7396 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 373.0669 |  | -1.0218 | -8.3797 | -7.3579
2.125 | 39.7389 | 39.9425 | 40.625 |  | -0.20365 | -0.8861 | -0.68245 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 404.1558 |  | -2.026 | -8.8154 | -6.7894

## BED 10
2015-09-29
Pressure sensor offset calibration | [13 empty] | Forward calculation for External 3-wire Pressure Sensor 0-70 BAR Absolute, Ratiometric, 0.5-4.5V Output; US331-000004-070BA
| V sensor [V] (voltage measured at sensor output, 0-5V range) | [8 empty] | Note | Sensor SN | Sensor model |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
Pressure applied to sensor in addition to atmospheric pressure [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
BED/Sensor Number (Volts) | [14 empty] | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 70 | 0.5 | 4.5 | -8.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -96.9974
BED 3 | [11 empty] | PHE163 100 bar |  |  | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 70 | 0.5 | 4.5 | 0 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -9.9485
BED 4 | [11 empty] | PHE163 100 bar |  |  | 65535 | 11.14% | 7302.472 | 2.5 | 0.27857 | 0.55714 | 0 | 70 | 0.5 | 4.5 | 1 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 7.4613e-06
BED 5 | [11 empty] | PHE163 100 bar |  |  | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 70 | 0.5 | 4.5 | 13.125 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 120.625
BED 6 | [11 empty] | PHE163 100 bar |  |  | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 70 | 0.5 | 4.5 | 35 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 338.2473
[15 empty] | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 70 | 0.5 | 4.5 | 56.875 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 555.8697
BEDS 2 | [9 empty] | DVK 10/8/2015 |  | PHE163 100 bar |  |  | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 70 | 0.5 | 4.5 | 70 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 686.4431
BED 7 | [9 empty] | DVK 4/4/2016 | 21110001 | PHE163 100 bar |  |  | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 70 | 0.5 | 4.5 | 78.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 773.4921
BED 8 | [9 empty] | DVK 4/4/2016 | 21110002 | PHE163 100 bar
BED 9 | [9 empty] | DVK 4/4/2016 | 21110003 | PHE163 100 bar
BED 10 | 0.509 | 0.708 | 0.912 | 1.119 | 1.325 | 1.527 | 1.729 | 1.926 | 2.117 | DVK 4/4/2016 | 21110004 | PHE163 100 bar |  |  | Forward calculation for External 3-wire Pressure Sensor 0-100 bar Sealed Gauge, Ratiometric, 0.5-4.5V Output; PHE163 100 bar
BED 11 | [9 empty] | DVK 4/4/2016 | 21110005 | PHE163 100 bar |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
Pressure on sensor + P atmosphere [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 100 | 0.5 | 4.5 | -12.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -124.3556
[15 empty] | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 100 | 0.5 | 4.5 | 0 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 0
[15 empty] | 65535 | 11.14% | 7302.4713 | 2.5 | 0.27857 | 0.55714 | 0 | 100 | 0.5 | 4.5 | 1.4286 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 14.2121
[15 empty] | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 100 | 0.5 | 4.5 | 18.75 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 186.5335
[15 empty] | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 100 | 0.5 | 4.5 | 50 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 497.4225
volts = 2.50 * (counts/65536) | [12 empty] | Legend: |  | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 100 | 0.5 | 4.5 | 81.25 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 808.3116
(pressFullScale/2) * volts - pressFullScale/8 | [12 empty] | constant |  | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 100 | 0.5 | 4.5 | 100 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 994.8451
for 100 bar sensor: | [12 empty] | auto calculated |  | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 100 | 0.5 | 4.5 | 112.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 1119.2007
50.0 * volts - 12.5 | [12 empty] | input
Forward calculation - Equations, plus simplified equations for PHE163 100 bar sensor
[15 empty] | Constant = 2^16-1 | Input | Counts FS * Counts%FS | Constant | Vref * Counts / Counts FS | Vadc * 2 | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | (Pmax-Pmin) / (Vsens max-Vsens min) * (Vsens- Vmin) + Pmin | Constant (sensor parameter) | Constant (atmospheric pressure) | Constant (gravity) | Constant (water density) | 1/ g [m/s2] / R [kg/m3] | k [m/Pa] * 1E5 [Pa/bar] | k [m/bar] * (Pcalc [bar] + Pref [bar] - Patm [bar])
[15 empty] | 65535 | input | 65535 * input | 2.5 | 2.5 * counts / 65536 | 5 * counts /65536 | 0 | 100 | 0.5 | 4.5 | Pmax / 4 * Vsens - Pmax /8 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 9.948* Pcalc [bar]
Calibration Vsens vs Pressure
Linear |  | polynomial |  | Nominal
slope | 24.7336 | a | 0.4229 | slope | 25
intercept | -12.6264 | b | 23.622 | Intercept | -12.5
|  | c | -12.012
R2 | 0.9999 |  | 0.99996
Water depth differences according to different calibrations, BED #10
| Pressure,Bars Linear calibration | Pressure, Bars Polynomial calibration | Pressure, Bars Nominal calibration |  | Delta pressure, bars, Linear - Polinomial | Delta pressure, bars, Linear - Nominal | Delta pressure, bars, Polynomial - Nominal | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] |  | Nominal Depth, m |  | Water depth difference meters, Lineal- polynomial | Water depth difference, meters Lineal -nominal | Water depth difference, meters Polynomial - Nominal
Volts
0.5 | -0.25959 | -0.095275 | 0 |  | -0.16431 | -0.25959 | -0.095275 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 0 |  | -1.6347 | -2.5825 | -0.94784
0.75 | 5.9238 | 5.9424 | 6.25 |  | -0.018563 | -0.32618 | -0.30762 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 62.1778 |  | -0.18468 | -3.245 | -3.0603
1 | 12.1072 | 12.0329 | 12.5 |  | 0.074325 | -0.39278 | -0.4671 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 124.3556 |  | 0.73942 | -3.9075 | -4.6469
1.25 | 18.2906 | 18.1763 | 18.75 |  | 0.11435 | -0.45937 | -0.57372 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 186.5335 |  | 1.1376 | -4.57 | -5.7076
1.5 | 24.474 | 24.3725 | 25 |  | 0.10151 | -0.52596 | -0.62748 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 248.7113 |  | 1.0099 | -5.2325 | -6.2424
1.75 | 30.6574 | 30.6216 | 31.25 |  | 0.035814 | -0.59255 | -0.62837 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 310.8891 |  | 0.35629 | -5.895 | -6.2513
2 | 36.8409 | 36.9236 | 37.5 |  | -0.082748 | -0.65915 | -0.5764 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 373.0669 |  | -0.82321 | -6.5575 | -5.7343
2.125 | 39.9326 | 40.0944 | 40.625 |  | -0.16185 | -0.69244 | -0.53059 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 404.1558 |  | -1.6102 | -6.8888 | -5.2786

## BED 11
2015-09-29
Pressure sensor offset calibration | [13 empty] | Forward calculation for External 3-wire Pressure Sensor 0-70 BAR Absolute, Ratiometric, 0.5-4.5V Output; US331-000004-070BA
| V sensor [V] (voltage measured at sensor output, 0-5V range) | [8 empty] | Note | Sensor SN | Sensor model |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
Pressure applied to sensor in addition to atmospheric pressure [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 70 | 0.5 | 4.5 | -8.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -96.9974
BED/Sensor Number (Volts) | [14 empty] | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 70 | 0.5 | 4.5 | 0 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -9.9485
BED 3 | [11 empty] | PHE163 100 bar |  |  | 65535 | 11.14% | 7302.472 | 2.5 | 0.27857 | 0.55714 | 0 | 70 | 0.5 | 4.5 | 1 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 7.4613e-06
BED 4 | [11 empty] | PHE163 100 bar |  |  | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 70 | 0.5 | 4.5 | 13.125 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 120.625
BED 5 | [11 empty] | PHE163 100 bar |  |  | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 70 | 0.5 | 4.5 | 35 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 338.2473
BED 6 | [11 empty] | PHE163 100 bar |  |  | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 70 | 0.5 | 4.5 | 56.875 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 555.8697
[15 empty] | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 70 | 0.5 | 4.5 | 70 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 686.4431
BEDS 2 | [9 empty] | DVK 10/8/2015 |  | PHE163 100 bar |  |  | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 70 | 0.5 | 4.5 | 78.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 773.4921
BED 7 | [9 empty] | DVK 4/4/2016 | 21110001 | PHE163 100 bar
BED 8 | [9 empty] | DVK 4/4/2016 | 21110002 | PHE163 100 bar
BED 9 | [9 empty] | DVK 4/4/2016 | 21110003 | PHE163 100 bar |  |  | Forward calculation for External 3-wire Pressure Sensor 0-100 bar Sealed Gauge, Ratiometric, 0.5-4.5V Output; PHE163 100 bar
BED 10 | [9 empty] | DVK 4/4/2016 | 21110004 | PHE163 100 bar |  |  | Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
BED 11 | 0.482 | 0.68 | 0.887 | 1.096 | 1.304 | 1.508 | 1.708 | 1.908 | 2.103 | DVK 4/4/2016 | 21110005 | PHE163 100 bar |  |  | 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 100 | 0.5 | 4.5 | -12.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -124.3556
Pressure on sensor + P atmosphere [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | [5 empty] | 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 100 | 0.5 | 4.5 | 0 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 0
[15 empty] | 65535 | 11.14% | 7302.4713 | 2.5 | 0.27857 | 0.55714 | 0 | 100 | 0.5 | 4.5 | 1.4286 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 14.2121
[15 empty] | 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 100 | 0.5 | 4.5 | 18.75 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 186.5335
[15 empty] | 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 100 | 0.5 | 4.5 | 50 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 497.4225
[15 empty] | 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 100 | 0.5 | 4.5 | 81.25 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 808.3116
volts = 2.50 * (counts/65536) | [12 empty] | Legend: |  | 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 100 | 0.5 | 4.5 | 100 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 994.8451
(pressFullScale/2) * volts - pressFullScale/8 | [12 empty] | constant |  | 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 100 | 0.5 | 4.5 | 112.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 1119.2007
for 100 bar sensor: | [12 empty] | auto calculated
50.0 * volts - 12.5 | [12 empty] | input |  | Forward calculation - Equations, plus simplified equations for PHE163 100 bar sensor
[15 empty] | Constant = 2^16-1 | Input | Counts FS * Counts%FS | Constant | Vref * Counts / Counts FS | Vadc * 2 | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | (Pmax-Pmin) / (Vsens max-Vsens min) * (Vsens- Vmin) + Pmin | Constant (sensor parameter) | Constant (atmospheric pressure) | Constant (gravity) | Constant (water density) | 1/ g [m/s2] / R [kg/m3] | k [m/Pa] * 1E5 [Pa/bar] | k [m/bar] * (Pcalc [bar] + Pref [bar] - Patm [bar])
[15 empty] | 65535 | input | 65535 * input | 2.5 | 2.5 * counts / 65536 | 5 * counts /65536 | 0 | 100 | 0.5 | 4.5 | Pmax / 4 * Vsens - Pmax /8 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 9.948* Pcalc [bar]
Calibration Vsens vs Pressure
Linear |  | polynomial |  | Nominal
slope | 24.5435 | a | 0.3953 | slope | 25
intercept | -11.8411 | b | 23.522 | Intercept | -12.5
|  | c | -11.29
R2 | 0.9999 |  | 0.99996
Water depth differences according to different calibrations, BED #11
| Pressure,Bars Linear calibration | Pressure, Bars Polynomial calibration | Pressure, Bars Nominal calibration |  | Delta pressure, bars, Linear - Polinomial | Delta pressure, bars, Linear - Nominal | Delta pressure, bars, Polynomial - Nominal | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] |  | Nominal Depth, m |  | Water depth difference meters, Lineal- polynomial | Water depth difference, meters Lineal -nominal | Water depth difference, meters Polynomial - Nominal
Volts
0.5 | 0.43063 | 0.56982 | 0 |  | -0.1392 | 0.43063 | 0.56982 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 0 |  | -1.3848 | 4.2841 | 5.6689
0.75 | 6.5665 | 6.5739 | 6.25 |  | -0.0073477 | 0.31651 | 0.32386 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 62.1778 |  | -0.073099 | 3.1488 | 3.2219
1 | 12.7024 | 12.6273 | 12.5 |  | 0.075091 | 0.20239 | 0.1273 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 124.3556 |  | 0.74704 | 2.0135 | 1.2664
1.25 | 18.8383 | 18.7302 | 18.75 |  | 0.10812 | 0.088273 | -0.019844 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 186.5335 |  | 1.0756 | 0.87818 | -0.19741
1.5 | 24.9742 | 24.8824 | 25 |  | 0.09173 | -0.025845 | -0.11757 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 248.7113 |  | 0.91257 | -0.25712 | -1.1697
1.75 | 31.11 | 31.0841 | 31.25 |  | 0.025931 | -0.13996 | -0.16589 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 310.8891 |  | 0.25798 | -1.3924 | -1.6504
2 | 37.2459 | 37.3352 | 37.5 |  | -0.08928 | -0.25408 | -0.1648 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 373.0669 |  | -0.8882 | -2.5277 | -1.6395
2.125 | 40.3139 | 40.4793 | 40.625 |  | -0.16542 | -0.31114 | -0.14572 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 |  | 404.1558 |  | -1.6456 | -3.0954 | -1.4497

## data
2015-09-29
Pressure sensor offset calibration
| V sensor [V] (voltage measured at sensor output, 0-5V range) | [8 empty] | Note | Sensor SN | Sensor model
Pressure applied to sensor in addition to atmospheric pressure [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40
BED/Sensor Number (Volts)
BED 3 | 0.51 | 0.72 | 0.92 | 1.13 | [6 empty] | 20298001 | PHE163 100 bar
BED 3 | 0.5105 | 0.7116 | 0.9193 | 1.1289 | 1.3378 | 1.5401 | 1.7423 | 1.9378 | 2.1316 | DVK 9/23/16, fluke 189 | 20298001 | PHE163 100 bar
BED 4 | 0.4935 | 0.7 | 0.9 | 1.1 | [6 empty] | 20298002 | PHE163 100 bar
BED 4 | 0.495 | 0.694 | 0.901 | 1.109 | 1.3175 | 1.517 | 1.718 | 1.912 | 2.106 | DVK 9/23/16, fluke 75 & 189 | 20298002 | PHE163 100 bar
BED 5 | 0.528 | 0.734 | 0.94 | 1.15 | [7 empty] | PHE163 100 bar
BED 6 | 0.505 | 0.71 | 0.918 | 1.113 | 1.33 | 1.53 | 1.73 | 1.93 | 2.12 |  |  | PHE163 100 bar
BEDS 2 | 0.484 | 0.684 | 0.892 | 1.095 | 1.302 | 1.502 | 1.703 | 1.895 | 2.09 | DVK 10/8/2015 |  | PHE163 100 bar
BED 7 | 0.497 | 0.694 | 0.899 | 1.106 | 1.312 | 1.512 | 1.711 | 1.907 | 2.1 | DVK 4/4/2016 | 21110001 | PHE163 100 bar
BED 8 | 0.492 | 0.691 | 0.899 | 1.108 | 1.316 | 1.517 | 1.718 | 1.913 | 2.108 | DVK 4/4/2016 | 21110002 | PHE163 100 bar
BED 9 | 0.508 | 0.712 | 0.917 | 1.124 | 1.333 | 1.534 | 1.733 | 1.933 | 2.124 | DVK 4/4/2016 | 21110003 | PHE163 100 bar
BED 10 | 0.509 | 0.708 | 0.912 | 1.119 | 1.325 | 1.527 | 1.729 | 1.926 | 2.117 | DVK 4/4/2016 | 21110004 | PHE163 100 bar
BED 11 | 0.482 | 0.68 | 0.887 | 1.096 | 1.304 | 1.508 | 1.708 | 1.908 | 2.103 | DVK 4/4/2016 | 21110005 | PHE163 100 bar
Pressure on sensor + P atmosphere [bar] | 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40
volts = 2.50 * (counts/65536)
(pressFullScale/2) * volts - pressFullScale/8
for 100 bar sensor:
50.0 * volts - 12.5
Forward calculation for External 3-wire Pressure Sensor 0-70 BAR Absolute, Ratiometric, 0.5-4.5V Output; US331-000004-070BA
| Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
| 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 70 | 0.5 | 4.5 | -8.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -96.9974
| 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 70 | 0.5 | 4.5 | 0 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -9.9485
| 65535 | 11.14% | 7302.472 | 2.5 | 0.27857 | 0.55714 | 0 | 70 | 0.5 | 4.5 | 1 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 7.4613e-06
| 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 70 | 0.5 | 4.5 | 13.125 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 120.625
| 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 70 | 0.5 | 4.5 | 35 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 338.2473
| 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 70 | 0.5 | 4.5 | 56.875 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 555.8697
| 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 70 | 0.5 | 4.5 | 70 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 686.4431
| 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 70 | 0.5 | 4.5 | 78.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 773.4921
Forward calculation for External 3-wire Pressure Sensor 0-100 bar Sealed Gauge, Ratiometric, 0.5-4.5V Output; PHE163 100 bar
| Counts FS | Count %FS | Counts | Vref [V] | Vadc [V] | Vsens [V] | Pmin [bar] | Pmax [bar] | Vsens min [V] | Vsens max [V] | Pcalc [bar] | Pref [bar] | Patm [bar] | g [m/s2] | R [kg/m3] | k[m/Pa] | k [m/bar] | Dwater [m]
| 65535 | 0% | 0 | 2.5 | 0 | 0 | 0 | 100 | 0.5 | 4.5 | -12.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | -124.3556
| 65535 | 10% | 6553.5 | 2.5 | 0.25 | 0.5 | 0 | 100 | 0.5 | 4.5 | 0 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 0
| 65535 | 11.14% | 7302.4713 | 2.5 | 0.27857 | 0.55714 | 0 | 100 | 0.5 | 4.5 | 1.4286 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 14.2121
| 65535 | 25% | 16383.75 | 2.5 | 0.625 | 1.25 | 0 | 100 | 0.5 | 4.5 | 18.75 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 186.5335
| 65535 | 50% | 32767.5 | 2.5 | 1.25 | 2.5 | 0 | 100 | 0.5 | 4.5 | 50 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 497.4225
| 65535 | 75% | 49151.25 | 2.5 | 1.875 | 3.75 | 0 | 100 | 0.5 | 4.5 | 81.25 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 808.3116
| 65535 | 90% | 58981.5 | 2.5 | 2.25 | 4.5 | 0 | 100 | 0.5 | 4.5 | 100 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 994.8451
| 65535 | 100% | 65535 | 2.5 | 2.5 | 5 | 0 | 100 | 0.5 | 4.5 | 112.5 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 1119.2007
Forward calculation - Equations, plus simplified equations for PHE163 100 bar sensor
| Constant = 2^16-1 | Input | Counts FS * Counts%FS | Constant | Vref * Counts / Counts FS | Vadc * 2 | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | (Pmax-Pmin) / (Vsens max-Vsens min) * (Vsens- Vmin) + Pmin | Constant (sensor parameter) | Constant (atmospheric pressure) | Constant (gravity) | Constant (water density) | 1/ g [m/s2] / R [kg/m3] | k [m/Pa] * 1E5 [Pa/bar] | k [m/bar] * (Pcalc [bar] + Pref [bar] - Patm [bar])
| 65535 | input | 65535 * input | 2.5 | 2.5 * counts / 65536 | 5 * counts /65536 | 0 | 100 | 0.5 | 4.5 | Pmax / 4 * Vsens - Pmax /8 | 1 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | 9.948* Pcalc [bar]
Legend:
constant
auto calculated
input

## BED 1
Constant = 2^16-1 | Input | Counts FS * Counts%FS | Constant | Vref * Counts / Counts FS | Vadc * 2 | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | Constant (sensor parameter) | (Pmax-Pmin) / (Vsens max-Vsens min) * (Vsens- Vmin) + Pmin | Constant (sensor parameter) | Constant (atmospheric pressure) | Constant (gravity) | Constant (water density) | 1/ g [m/s2] / R [kg/m3] | k [m/Pa] * 1E5 [Pa/bar] | k [m/bar] * (Pcalc [bar] + Pref [bar] - Patm [bar])
65535 | input | 65535 * input | 2.5 | 2.5 * counts / 65536 | 5 * counts /65536 | 0 | 70 | 0.5 | 4.5 | 17.5 * Vsens - 8.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | K * (Pcalc [bar] -1)
65535 | input | 65535 * input | 2.5 | 2.5 * counts / 65536 | 5 * counts /65536 | 0 | 70 | 0.5 | 4.5 | 35 * Vadc - 8.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | K * (Pcalc [bar] -1)
65535 | input | 65535 * input | 2.5 | 2.5 * counts / 65536 | 5 * counts /65536 | 0 | 70 | 0.5 | 4.5 | 87.5*counts/65536 - 8.75 | 0 | 1 | 9.8066 | 1025 | 9.9485e-05 | 9.9485 | K * (Pcalc [bar] -1)