## Slide 1: BEDS Preliminary Design Review
May 15, 2012
Project Team:
PI: Charlie Paull
Lead Scientist: Bill Ussler
EE: Alana Sherman (2011), Denis Klimov (2012)
SE: Bob Herlien
ME: Dale Graves, Larry Bird
Project manager: Brian Kieft

## Slide 2: Outline?

## Slide 3: What is a Benthic Event Detector?
Characterizes motion and frequency of submarine canyon sediment debris flows
Allows for data recovery without instrument recovery
May not necessarily be recovered
Doesn’t require MBARI ships for operation

## Slide 4: BEDs High Risk Elements
Can data be retrieved from device?
Can motion be resolved in a way useful to science?

## Slide 5: BEDs Data Retrieval
Acoustic modem
  Burial test .7m @1200 bps – 2560 bps
  Vryhof Anchors: 13m in 300m water @ 300 bps
  Frequency hopping scheme for modem if needed

## Slide 6: BEDs Motion
Bound events
  Duration
  Frequency
  Orientation changes
  Accelerations
Linear Displacement
  Performed off-board initially and validated through survey

## Slide 7: BEDs Motion
Depth
Magnitude

## Slide 8: BEDs Architecture

## Slide 9: Why Invensense?
Cost $15
Low power
Sensor Fusion Algorithms onboard
Gyro drift and accel bias compensated
Event detection logic
Open source community and COTs devel boards
  Paparazzi code base includes peripherals like magnetometer, pressure
  Integrated 9 DOF solutions available for MSP430

## Slide 10: BEDs Flow Chart
Motion thresholds
Context data
Watch Cycle
Event recording
Command mode
Read Sensors
Motion detection
Thresholds
Time period (~1H)
Context data
Duty cycle ~5Min/1Hour
Activate modem
Listen for acoustic communication
Read Sensors
No Motion
thresholds
Calculate and scale data
Log data
~ 1Hz loop
~ 100Hz loop
Listen
Execute
Set upload range
Set thresholds
Read status
Specified Data upload
Decimated data
upload
Time-out
Exit command
Acoustic loop

## Slide 11: BEDs Electrical
Data and Power Budget:
Operating modes: Watch Cycle; recording Cycle; Command mode
Power of subsystems and duty cycles defined in each mode
Data rates and number of bytes per point defined
Parameter | Watch cycle | Recording cycle
Bytes/data point | 6 | 24
Data rate | 1 / Hour | 100 / sec
Event data, kB | Depends on length of deployment | Depends on length of event(s)
Parameter | Watch cycle | Recording cycle | Command mode
Acoustic modem power, W | 0.60 | 0.60 | 21.00
Acoustic modem duty cycle, % | 10% | 10% | 100%
Inertial sensor & microcontroller power, W | 0.16 | 0.40 | 0.40
Peripherals power, W | 0.04 | 0.20 | 0.20
Peripherals and sensor duty cycle, % | 20% | 100% | 100%
Cycle average power, W | 0.10 | 0.66 | 21.60
Duration in mode, hrs | Depends on length of deployment | Depends on length of event(s) | Depends on data to be transmitted
Cycle energy, WHr |  |  | 

## Slide 12: BEDs Electrical
Data and Power Budget:
Energy is limited and defines everything else
Lithium DD cells 72 W*H x 18 for deployable article
Alkaline D-cells for prototype
Available Energy
Watch Cycle
Recording cycle
Data transfer
Trade-off
Deployment length
Event recording time

## Slide 13: BEDs Electrical
Chart: Deployment duration as function of Event recording time
| 0 | 0.5 | 1 | 1.5 | 2 | 2.5
Deployment length, days | 538.75 | 433.1438 | 327.5375 | 221.9312 | 116.325 | 10.7188
Chart: Energy concumed in each operating mode
| 0 | 0.5 | 1 | 1.5 | 2 | 2.5
Watch cycle energy, WHr | 1293 | 1039.545 | 786.09 | 532.635 | 279.18 | 25.725
Recording cycle energy, WHr | 0 | 0.33 | 0.66 | 0.99 | 1.32 | 1.65
Data transfer energy, WHr | 3 | 256.125 | 509.25 | 762.375 | 1015.5 | 1268.625
Deployment length
Event recording time
Trade-off

## Slide 14: BEDs Software

## Slide 15: BEDs Mechanical

## Slide 16: Functional Requirements (Mechanical)
Operational Depth = 500 meters
Deployment duration = up to 12 months
Temperature range = -3 to 50 degrees C.
“Ruggedized” to withstand acceleration shocks associated with transport
Package density = 1.5 to 1.8 G/CC

## Slide 17: Functional Requirements (continued)
The system shall be packaged such that its density could be increased to 3 times the default package density so that self-burial can occur (i.e., it can be made more dense).
The system shall be designed as small as possible to approximate the natural motion of sediment within the canyon floor
Be deployable and recoverable from an MBARI ROV
The system shall communicate at any package orientation

## Slide 18: Main Components
Instrument housing
  7”ID X 10.5” cylinder w/modem & homer transducers
Trim/ballast/system
  15” Flotation sphere - collar w/ballast weights
Transducer guards
  UHMW “windows”

## Slide 19: Instrument Housing (internal components)
Homer transducer
Homer PCB
Homer battery pack
triaxial inertial sensor
Pressure transducer
Benthos G4 modem PCB assembly
Benthos modem transducer
Modem battery pack
Instrument battery pack

## Slide 20: Instrument Housing (materials)
Options
  Ti
    Pros – last for ever
    Cons – high cost
  Composite tube/Ti end caps
    Pros – Very long life
    Cons – higher costs than mild steel
  Aluminum
    Cons – very limited life due to corrosion and not predictable
  Mild steel w/coating and or anodes
    Pros – no crevice corrosion problems
      Corrosion is predictable and design can accommodate for this
      Inexpensive
    Cons – limited life time (2-3 years)
      Heavy compared to the Ti and Composite/Ti options

## Slide 21: Instrument Housing (internal packaging)

## Slide 22: Trim & Ballast System
Collar provides a hydrodynamic shape to the instrument housing that can be modified as needed.
  In addition, the collar provides flotation, ballast & trim adjustments

## Slide 23: Weights & Balance
Still working on this section

## Slide 24: Transducer Guards
Protects transducers
Provides hydrodynamic profile
UHMW properties provide good acoustic path.

## Slide 25: BEDs Integration and Test
Test jig
Prototype

## Slide 26: BEDs Prototype
250 m and 300 m test sites
4 km from ML harbor

## Slide 27: BEDs Future Improvements
Onboard linear displacement calculation?
Adaptive TX power reduction?
Clock drift reporting?
Adaptive Decimation (zip?)
Adaptive baud rate?
Drop weight system?
Atomic clock

## Slide 28: BEDs Resources and Schedule