# Why a BED? What’s Desired?

- Instrument that can characterize the motion and frequency of submarine canyon sediment debris flows

- Ability to recover data without recovering instrument (which may be impossible)

# Recent History

- Two main project level risks/questions:

  - Can data be brought back reliably

  - Can motion be resolved and interpreted

- Data Transfer

  - Acoustic modem (ConDR)

  - Tank tests

  - Beach face tests

  - Buried transducer test

  - Results

    - 1200bps – 2560 bps

    - Vryhof Anchors (3m)

    - Benthos frequency hopping signal scheme for added robustness (to 10m)

- Motion detection/capture

  - Linear Displacement (offboard)

    - Least squares positioning algorithm

    - Work continues on shore

  - Orientation changes

    - Preventing aliasing. 49.99Hz @ 200Hz sample rate

    - Upper range for gyros: 2000dps (5.55 rotations/sec)

  - Event detection

    - Time to wake?

  - Depth

  - Magnitude (Denis)

# System Requirements Summary

- Deployment duration

- Types of motion

- Burial

- Cost

- Clock

- Physical packaging

# Design

- Architecture

  - System Block Diagram

    - IMU (Invensense 6050) (I2C – direct connection to Atmel)

    - Atmel CPU (RS-232 serial debug port)

    - Acoustic modem (RS-232)

    - Mass storage (SPI)

    - RTC (SPI)

      - Drift up to 63 seconds per year

      - Hand calibration?

        - Systematic frequency offset (~.2 sec/day delta)

        - Temp correction for constant temp

    - Power Management (Digital I/O, Analog V and I measurements)

    - pressure sensor / A/D (SPI)

    - health sensing (Analog – onboard)

    - homer (designed to outlast main system)

    - Batteries

  - Data Flow

    - Collection

    - Visualization

- Electrical

  - Data budget

    - Currently 22 bytes/sample

    - Could be reduced? To 16 bytes/sample

    - Accel/Gyro sensitivity settings

  - Power budget

    - Number of events vs. time

    - Total time of events vs power

  - A/D circuit

- Mechanical

  - Housing design – 15” sphere initial package

  - Material? Aluminum?

  - Cost?

  - Lead time – 4-6 wks

- Software

  - Onboard

  - Post processing

  - Visualization

    - VRML

    - Other

# Testing

- Lab tests

- Sensor comparison

# Deployment Plan / Schedule

- Lab

- Sea

  - Prototype test  250 m depth (36.7908, -121.8346) ~4km from ML

  - Prototype test  300 m depth (36.7945, -121.8481) ~5km from ML

# Future Improvements

- Onboard linear displacement calculation

- Adaptive TX power reduction

- Clock drift reporting

- Adaptive baud rate

- Drop weight system?