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Hi Denis,

The oddness of the SD cards is also attributable to hot swapping the serial port.  Please don't touch the hardware when it is powered on. When the serial buffers get corrupted, the system does not respond well.

The SD card doesn't show all of its files because the software believes a key has been hit (due to buffer corruption when hot-swapping serial port). We use a key hit to stop the directory listing for cases where it is very long directory output.

I will modify the output to indicate the reason for terminating the command. 
There is nothing wrong with the SD cards.

For J8... I don't understand as it works here and your board worked here. I can get your hardware to check or we can walk through it on Zoom again now that I understand where the fragile parts are. Note that you cannot send 9600 baud characters to to J8 or the system gets very unhappy. (I have a code fix for that.)

I think the steps for J8 are:
    1. Start with the BEDS unit power off.
    2. Plug serial port into J8 and your non WinXP computer.
    3. Set up your terminal to 115200 8-N-1 (no flow control).
    4. Power on the unit.
    5. Hit return a few times.

I'm free to talk at 3pm.
Time: Dec 17, 2021 03:00 PM Pacific Time (US and Canada)
https://zoom.us/j/96628102132?pwd=d0NRb2tBZzBHYWx3NVFvVVJhc0QwQT09
Meeting ID: 966 2810 2132
Passcode: 4eFjqQ
I also have code (2.0.11) for you to update your unit with to prevent the issues with 9600 baud garbage characters going to the J8 serial port running 115200 baud. And to output the reason for early termination of the directory command.

Thanks,
Elecia
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Downloadable binaries are here:
https://www.dropbox.com/s/upvmzzaz2klrfer/20211217%20BEDS%20Release%20v2.0.11%20Binaries%20Only.zip?dl=0

Note that these mitigate the issues caused by hotswapping serial ports:
- Print message when terminating dir command due to keyboard hit
- Don't try to run commands that are junk sent at the wrong baud rate
Thanks,
Elecia
