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Inspection report
Program Number: 901006
BEDS Benthic Even Detector
Author: Denis Klimov; 831-775-1914; klimov@mbari.org
Date: 5/9/2014

Hardware affected: 

BEDS#3 (AS1 SN003 12/2013)

Cell Electrochem model CSC93DD part# 3B0036-XB, batch 2334026, SN 7331
Background. 

BEDS#3 (AS1 SN003 12/2013) was built 12/6/2013 and deployed in December 2013 in Monterey Canyon at approximate depth of 290 meters. Initial environmental conditions inside of a housing were recorded by on-board sensor set at humidity 9%, temperature 8-9C, internal pressure 14.5 PSI
4/28/2014 an elevated humidity level was detected (85%; 9C; 14.8PSI) and decision was made to recover instrument as soon as possible. 

5/1/2014 at about 11:00 PST local time another read of environmental sensors was performed (89%; 9C, 14.8 PSI). 

5/1/2014 at about 16:00 PST BEDS instrument was recovered by ROV Tiburon and R/V Western Flyer, brought to Moss Landing and stored on site until next Monday. 
5/5/2014 ~ 13:00 PST, the instrument was opened for inspection. 

Initial Inspection

Housing does not have any signs of damage. Pressure slightly elevated, large clamps were used to hold housing’s end plated until pressure released as a safety measure (no pressure read was performed right before opening). Subjectively, internal pressure was not more than 110% of atmospheric pressure. 
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Once housing was open, presence of light-green liquid and strong chlorine smell became evident
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About 25-35mL of the light-green liquid was collected from the housing into polyethylene container. 
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Housing does not have any signs of mechanical damage. Inside surface of the housing contaminated with light-green fluid. 
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About 25-35mL of the light-green liquid was collected from the housing into polyethylene container and then transferred to a vile. About half was spilled on a plastic cart while working outside of the building when wind overturn plastic container. 
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Left Black Delrin endplate holding battery sticks found missing about 1x3 cm section adjacent to Battery Stick marked L2 (means: stick “2”, oriented to the left end plate, assembly AS3-LEFT). 
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Edges of the missing section appeared “rounded” as if melted or chemically etched. 
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Internals of Electropack Assembly –AS1 look undisturbed aside of presence of light-green fluid on some surfaces, and some discoloration of metals: pressure sensor housing, solder on PCB boards and so on:
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Battery pack board (10020895-BPB1) mounted on right end plate (10020895-AS3-RIGHT) has signs of corrosion near connectors J1-J2, but later tested to be fully functional
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Battery pack board (10020895-BPB1) mounted on LEFT end plate (10020895-AS3-LEFT) is corroded near connectors J1-J2 with black-green deposits covering connectors pins; and later tested NOT to be functional. Damage to the board include dissolved traces between pins J1.1 and J2.2, and J1.1 and copper pour net name “+ BATT” on Top and BOTTOM layers. 
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Black Delrin under the LEFT battery pack board is corroded with approximately 15x5x1mm mass missing, with discoloration and pitting in surrounding areas as if the electrolyte was trapped inside the channel under the PCB board enabling galvanic corrosion between energized connector pins.
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Fused Power Harnesses delivering power from each Battery Pack Board to Mother Board are intact and later tested to be functional. 
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Connector pins in Homer power harness are damaged but since central pin is redundant positive line, Homer is most likely functional
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Other side of damaged Homer power harness connector

[image: image14.png]



Coin cell on acoustic board stack has some surface corrosion/ discoloration 
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Inspection of battery sticks and cells. Table below indicates serial numbers of the cells installed in each slot as well as voltage of each cell measure at the time of the assembly and a the time of inspection after battery failure. Drawing below illustrates slots allocation in respective AS3- assembly. (AS2 SN003 11/26/2013)
	Slot#
	Model
	Part#
	Batch#
	SN#
	Voltage 11/26/2013
	Voltage 5/9/2014
	Note

	Slot L1
	CSC93DD
	3B0036-XB
	2334026
	3027
	3.886
	3.859
	positive wire missing; signs of corrosion under paper disk

	Slot L2
	same
	same
	same
	7331
	3.886
	3.86
	failed cell; 2 pitted areas on the edge on positive side; electrolyte leaked out

	Slot L3
	same
	same
	same
	7305
	3.886
	3.344
	 

	Slot L4
	same
	same
	same
	7306
	3.885
	3.85
	severe connector corrosion

	Slot L5
	same
	same
	same
	7339
	3.885
	3.061
	 

	Slot L6
	same
	same
	same
	7340
	3.885
	3.739
	 

	Slot L7
	same
	same
	same
	7334
	3.885
	3.852
	 

	 
	 
	 
	 
	 
	 
	 
	 

	Slot R7
	same
	same
	same
	3026
	3.885
	3.856
	 

	Slot R6
	same
	same
	same
	7332
	3.885
	3.856
	other side of failed cell in slot L2

	Slot R5
	same
	same
	same
	7325
	3.885
	3.856
	 

	Slot R4
	same
	same
	same
	7336
	3.885
	3.856
	 

	Slot R3
	same
	same
	same
	7335
	3.886
	3.856
	 

	Slot R2
	same
	same
	same
	3029
	3.886
	3.857
	 

	Slot R1
	same
	same
	same
	7327
	3.886
	3.856
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Affected “battery stick” from Slot L2, SN 7331
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Affected “battery stick” from Slot L2, SN 7331, heat shrinking material partially removed
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Plastic cap with protection elements of cell CSC93DD part# 3B0036-XB, batch 2334026, SN 3027
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Conclusions of system-level inspection: 
1. Damage to the electrical system of the BEDS assembly is caused by a presence of corrosive liquid inside the housing. 

2. Source of the liquid:

a. Battery cell SN 7331 located in a Slot L2 of the LEFT half-pack

b. Possibly, presence of seawater inside of the housing

3. LEFT half-pack failed due to corrosion of battery pack board traces. Channel milled under battery pack board negatively affected reliability allowing liquid to be trapped under the board and electrochemical corrosion to occur. Due to damage to the battery pack board, other cells in the same LEFT half-pack were affected (partially discharged beyond unaffected cells) as indicated by their voltage at the time of the inspection. 
4. RIGHT half-pack retained full functionality and continued supplying power to Mother Board. 
Suggestions and improvements: 
Batteries: 

Slot 

Battery and electrical: 

· Pack configuration: Electrochem CSC93, 7s2p 
· DC load discharge range: 27.5 to 17.5 VDC
· Pulse load discharge range: 17.5 to 13.5V at 4A peak 
Configuration 
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