Dear Drs. Kemp and Reimers,

Thank you for the favorable reviews of our manuscript. We have followed the reviewer suggestions in all cases. Below we outline the specific changes made, with suggestions in bold and our replies in regular font.

Sincerely,

Christine Huffard

Reviewer: 1

Comments to the Author

The following detailed comments relate to the line number in the pdf copy of the manuscript.

73        50 mab is pretty close to the seafloor so that resuspension of bottom sediments could influence the signal, as indicated later in discussing significant short-term events (line 489). What do the CTD profiles in the area show in terms of bottom mixed-layer thicknesses in this region (temperature, salinity, turbidity)? Address in the Results section. This has no impact on the validity of the SES instrumentation, just the interpretation of the results.

We have cited previous studies at Station M that showed that resuspension of bottom sediments was not a major factor at 50 m above bottom (lines 328-335)

198-199        This is very interesting that you can sense luminescence or phosphorescence. I don’t see that anything else is mentioned about this in the paper.

We have added to lines 433-436: “No significant activity was evident on the 590 nm record for either deployment, nor was any luminescence or phosphorescence seen when the fluorometer excitation source was switched off.”

226        Please quantify the variability of the signal when no reference material is on the slide and connect this to the caution about measurements made when less than 20% of the slide is covered.

We have now included the standard error of readings of slides with no reference material on lines 228-229.

229        Define S.E. PM counts. The caption in Fig. 4 contains “standard error”, but this needs clarification the first time it is used.

We have provided a clearer definition of this concept and more appropriately called the term photomultiplier counts per second (PMC s-1) (Lines 193-195).

233       …reference targets OF VARYING SIZES […] and a 2.3 cm2 leaf cutting …

Fixed

255      The software section is pretty heavy reading and far more detailed than in most science papers. Given that this is a Methods paper this detail may be acceptable (or even necessary). I leave that up the editor.

Fixed

395-399        The figure referenced should be Figure 7, not 6.

Fixed

429        Please explain the “integrated transmissometer readings.” Searching the pdf copy, I find no other mention of a transmissometer. You refer later to light transmission two times.  The connection needs to be made. Is there a transmissometer on the SES looking at the sample slide or is it on the CTD measuring transmission in the water? If the CTD, how many measurements are made and averaged? Note that %light transmission is not linearly related to particle concentration, but if you convert %transmission to beam attenuation (/m), there is a linear correlation.

Throughout the manuscript we have clarified the text and eliminated any wording that would suggest there is a transmissometer on the SES.

432     a) should be b)

Fixed

425 – Figure 8       Note in the figure caption that the time (x) axes are different. To show this in the figure, one suggestion is to add dashed line from the beginning of the two graphs, a dashed line between the two graphs where they end, and between the major even in September in both graphs. A much better solution is to match the time scales of light attenuation and fluorescence so the eye can match the synchronous signals easily. The added white space that would result above the “3 h collect time” would be well worth the clarity of matching the two signals during the “12 h collect time.”

Fixed

442        Because sediment traps collect particles through a process of fluid exchange, not simple settling (as I’m sure you know well), it is more appropriate to use the cross-sectional area of the top of the trap than the area of the baffle opening. It is probably only a few percent different.

We have clarified wording here to give the size of the collection opening (line 394)

489        Were currents ever measured during these deployments? The flux variability could be from increased currents changing the bottom mixed layer depth and introducing resuspended sediment to the SES trap.

Although currents were not measured at the SES, previous research suggests resuspension is not a major consideration at this depth (lines 329-335).

499-500       A better transition is needed between these two paragraphs.

Fixed

500        The Bishop, 2009 reference should be after “Carbon Flux

explorer” since that is the first time you talk about it.

Fixed

535-566. Since some of the citations contain up to 9 authors, all of the citations should be complete with all authors.

Fixed

Reviewer: 2

From attached pdf

Lines 43-49 – This long, enumerated list could be broken up for easier reading…?

We have split up this list for easier reading.

Lines 93, Fig. 2, Fig. 3 – Cleaning brushes – they present a lot of surface area and collect organic

material. Presumably they serve as a great substrate for a healthy microbial community – how

fast do they foul? A brief description of what the brushes look like after recovery, whether they

clean as well months into the deployment, and whether they are easily replaced with fresh brushes would be good.

We have added mention of the cleaning brush condition following deployment at Station M. Fouling is not an issue at that site (Lines 100-102).

Lines 144-145 – Unclear what is meant by mechanical and optical distances. Can you add a

diagram of all of the optical and imaging components? While these are described in detail in the

text it would be easier to follow with a visual. AND Lines 146 – What does depth of field “+1.6/-3.2 mm” mean ? Please clarify.

We have clarified this wording (Lines 128-134) “The measured thickness of the glass and water between the camera lens and sample plate is 113 mm. However, based on the thicknesses and refractive indices of the optical elements, a computer-aided design model of the assembly calculated the in-air optical distance as 82 mm. Once the focus was set to 82 mm, the aperture was adjusted to provide a depth of field that maintains image detail from at least 3.2 mm above the top surface of the sample plate to 1.6 mm below the surface. The camera lens, adjusted using this procedure, required no further readjustments once the SES was submerged.”

Lines 172, 177 – The notation “590 x 10 nm” is unfamiliar to me. I assume it means a band

centered at 590 nm, but I don’t know whether 10 nm is the bandwidth or the half bandwidth, nor whether this is the bandwidth at half-maximum sensistivity or ….? Please be more specific.

We have clarified to read  “590×10 nm (590 nm center wavelength with 10 nm bandwidth, full width at half maximum)” (Lines 162-163)

Lines 177 – What is the spectral shape of the blue LED illumination? 480 +/- 30 nm is quite

wide so the initial spectrum is important as well.

We have clarified to read “Fluorescence excitation is provided by a blue 470×25 nm LED (Lumileds model LXHL-NB98) followed by a 480×30 nm bandpass filter to minimize crosstalk between optical channels. “ (Lines 167-169)

Lines 209-230 – Fluorometer calibration vs. area of 6mm-thick fluorescent-dyed PMA is

interesting and a good first pass. However, this procedure did not test the possibility of selfshading by optically thick accumulated particles (they could be optically thick if geometrically thick, or if there is just a high density of pigment). By “self-shading” I mean a particle can be thick or strongly pigmented enough that the interior may not absorb as much of the excitation light as the exterior. It would be useful to vary the pigmentation or the thickness of the calibration disks as well as the area, as these variations would make a better parallel to actual particles.

We have added the following text to clarify (Lines 215-217):

“We did not vary thickness of the reference objects to test for self-shading, which may occur when samples overlap on the slide and interfere with the sensor’s ability to excite and absorb a signal from underlying particles.”

Lines 263-339, Figures 5 and 6. I think the description of the Mission Plan (Lines 295-306) and

general discussion of the system’s flexibility for different missions is useful to the scientific

reader. However, the rest of the information in this section can be removed without

compromising the paper’s ultimate goals. This is not to say that details such as the software

architecture and mission text file format are unimportant – but they are not necessary to a

scientific description of the instrument. Rather, they would make more sense to include in the

SES technical manual.

This entire section has been trimmed significantly to address this reviewer’s suggestion.

Line 401 – I think you mean that seawater filters out the longer wavelengths? (or shorter

frequencies?)

Fixed to read: “Note the blue tint of the background light in the front-lit image, but not the back-lit image. Both the front and back lights use the same type of white LED, but only the backlight passes through the sample plate material, absorbing the shorter wavelengths.” (lines 364-366).

Lines 413-416 – Please explain more precisely how you computed the reported attenuation

factors. Is it correct that the reported value is a function both of the stage area covered and the

thickness of the accumulated material?

Lines 368-386 have been clarified to provide better details about how to replicate this brightness value.

Figure 8a: Left-hand y-axis maximum tick label should be 100%, not 1000%.

Fixed

Figure 8b: Could you replace the red dots with horizontal bars that show the 10-day collection

time period for each trap sample? Also, add a text annotation in red pointing out the “clog” in

the sediment trap. Could you format panels 8 and b so that their x-axes align, and so the reader

can more easily see when 690 nm fluorescence aligns with light attenuation and when it does

not?

We have made the requested changes to the figure

Lines 439-441 – Please provide more detail on trap sample handling. Specific questions that are

relevant here: In what depth of water were the station M traps deployed? (Were they at the

same depth as the SES?) Were the sample bottles filled with poisoned brine? How long were

samples held prior to retrieval? How were samples handled between return to the lab and freezedrying? Perhaps there is another reference that has already detailed the Station M trap sampling protocols, that you could cite here—that would also be acceptable.

We have provided a few more details, left critical information such as trap depth, and added a citation to where full sample processing have been explained in a previous publication (lines 389-394).

Lines 451-453 – This non-linearity may be true at high loadings but there is no reason to assume

that attenuance by particles is not a linear function of mass at low loadings. If the accumulated

material is not completely opaque then Beer’s Law should apply. See Estapa et al. (2013) and

Bishop et al. (2016) for related discussions.

We have added reference to these citations. These references also do not appear to work under an assumption of a linear relationship between mass accumulation and light attenuation (lines 401-404).

Lines 475-476 – Was there a qualitative difference visible in the images of the collected material

between these two periods? If so, it would be illustrative to show some of the image data.

We clearly saw higher counts of salp tunics and fecal pellets in the SES images in the earlier pulse, as noted. Due to the high coverage over the slide unfortunately it wasn’t possible to make other comparisons.

Paragraph starting line 500: This “compare and contrast” section feels abrupt and out of place in

the discussion. It might be better to move some of the material into the introduction/background

section, perhaps returning briefly to table 2 in the discussion if you feel the point-by-point

comparison is necessary. At the very least, a “segue” sentence leading into this paragraph is

necessary. Also, note that Bishop et al. (2016) contains the most recent CFE data and

description. I believe that paper states that the CFE does NOT transmit image data via satellite,

and there may be other small inaccuracies. You may want to double check the details in this text

and in your table 2.

Response: A better segue was added and the table has been updated.

Paragraph starting line 515 (“Future work…”): This instrument would also be useful

(particularly the fluorescence capability) for upper-mesopelagic work if a high-aspect ratio

cylindrical tube could be substituted for the conical trap, and the mechanism mounted on a

surface tethered, drifting array. See Buesseler et al. (2007) for a discussion of hydrodynamic

issues with conical traps in shallow water, however.

There are many possible adaptations of the SES. We are happy the reviewer has identified one of many alternate applications. At this early stage we think it is preliminary to offer a discussion of configuration changes for use in other environments. This discussion would be incomplete and speculative.

References:

Bishop, J. K. B., M. B. Fong, and T. J. Wood. 2016. Robotic observations of high wintertime

carbon export in California coastal waters. Biogeosciences Discuss. 2016: 1–32. doi:10.5194/bg-

2016-62

Estapa, M. L., K. Buesseler, E. Boss, and G. Gerbi. 2013. Autonomous, high-resolution

observations of particle flux in the oligotrophic ocean. Biogeosciences 10: 5517–5531.

doi:10.5194/bg-10-5517-2013