SeaBat T51-R

Rack-mounted High-Resolution Multibeam Sonar System

Operator's Manual

Part Number 1016039

Version 7B

October 2024

| Copyright Teledyne RESON A/S 2021-2024 |  |  |
| --- | --- | --- |
| Copyright Statement | All rights reserved. Reproduction, adaptation, or translation without prior written permission is prohibited, as per the copyright laws of Denmark. SeaBat T51-R Operator's Manual written by: Teledyne RESON A/S Fabriksvangen 35 3550 Slangerup Telephone: +45 47 38 00 22 Facsimile: +45 47 38 00 66 |  |
|  | Web: http://www.teledynemarine.com/reson | E-mail: reson-support@teledyne.com |
| Notice | The information contained in this document is subject to change without notice. Teledyne RESON makes no warranty of any kind with regard to this material, including (but not limited to) the implied warranties of merchant liability and fitness for a particular purpose. Teledyne RESON shall not be liable for errors contained herein or for incidental or consequential damages in connection with the furnishing, performance, or use of this material. |  |
| Copyright Release | No part of this manual may be copied or distributed, transmitted, transcribed, stored in a retrieval system, or translated into any human or machine readable language, in any form or by any means for disclosure to third parties without the express written consent of Teledyne RESON A/S. |  |
| Version | October 2024; Version 7B. Teledyne RESON strives to maintain up-to-date information and may, as necessary, review and revise this manual. If the information contained in this manual is unclear, please contact your nearest Teledyne RESON office for clarification. |
| Software Release | 3.0 or higher. This manual is consistent with the indicated release versions. |
| Electronic File Version | This manual is available from Teledyne RESON in the Adobe® Acrobat® Portable Document Format (PDF). The selected PDF process preserves the interactive index and cross-reference features of the original Microsoft Word document. |  |

Preface – About This Manual

This manual is designed to give detailed procedures for the correct installation, operation, and maintenance of the SeaBat T51-R sonar system and is equally useful for both the first time user and the experienced technician.

Read this manual thoroughly and follow the steps provided to maintain optimal safety standards and to obtain the best system performance.

If you require additional information or need clarification of any part of this document, please contact Teledyne RESON Customer Support for assistance at reson-support@teledyne.com.

Document Conventions

The following information will make it easier for the reader to interpret the document content.

| Convention | Definition |
| --- | --- |
| Hyperlink | Indicates a cross-reference to Ctrl+Click to reach the specified item. |
| Numbered list | Indicates stepwise instructions to be followed in a particular order. |
| Bulleted list | Indicates items of a list without any particular order. |

Notes, Cautions, and Warnings

Notes, Cautions, and Warnings are included to provide valuable information and to help the operator prevent damage to the sonar or to personnel. These are the formats:

NOTE

Note boxes provide explanatory information that may be useful to the operator, but is not necessarily vital to the operation of your Teledyne RESON system.

CAUTION

Caution boxes provide important information regarding your Teledyne RESON system. Disregarding Cautions may result in accidental misuse or damage to the system.

WARNING!

Warning boxes provide important information regarding your Teledyne RESON system. Disregarding Warnings may prove harmful to personnel working with or near the system.

Conversion of Units of Measurement

| Metric Conversion | Metric Unit; Imperial Unit. Millimeter (mm); 0.0394in. Centimeter (cm); 0.3937in. Meter (m); 1.0936yd. Kilometer (km); 0.6214mi. Kilogram (kg); 2.2046lb. |
| --- | --- |
| All measurements in this manual are referenced in metric terms. Use this reference table to convert between metric (e.g. cm, m, and km) and imperial (e.g. in, ft, and yd) measurements. |  |

| Temperature Conversion | Formula: F =; 1.8 x C + 32. Example:. If Temperature =; 30C, then:. F =; 1.8 x 30 + 32. F =; 86. |
| --- | --- |
| All temperatures in this manual are referenced in degrees Celsius. This is the formula for converting Celsius (C) to Fahrenheit (F). |  |

Glossary of Terms

| Entry | Definition |
| --- | --- |
| 7k | SeaBat 7000 series of multibeam sonars |
| AUV | Autonomous Underwater Vehicle |
| BITE | Built-In Test Environment |
| CE | Conformité Européenne |
| CFE | customer-furnished equipment |
| CW | Continuous wave |
| Cylindrical loss | The loss of sound by spreading in the form of a cylinder |
| DFD | Data Format Definition |
| DWG | Drawing |
| EA | Equiangular |
| ED | Equidistant |
| FM | Frequency modulation |
| INS | Inertial Navigation System |
| LED | light-emitting diode |
| N/A | not available or not applicable |
| OM | Operator’s Manual |
| PN | Part Number |
| RH | relative humidity |
| RMA | Return Material Authorization |
| ROV | Remotely Operated Vehicle |
| RSP | Rack-mounted Sonar Processor |
| Rx | receiver |
| SNR | signal-to-noise ratio |
| Spherical loss | The loss of sound by spreading in the form of a sphere |
| SVP | Sound Velocity Probe |
| TBD | To Be Defined |
| TPU | Total Propagated Uncertainty |
| TVG | Time-Varied Gain is a control intended to vary the gain of an amplifier over a predetermined time period (usually from a low to a high gain state) |
| Tx | transmitter |
| UI | user interface |
| UUV | Unmanned Underwater Vehicle |
| VDM | Vertical Detection Mode |
| WEEE | Waste Electrical and Electronic Equipment |
| ZDA | Time and date identifier |

Table of Contents

1	Introduction	1

1.1	System Overview	1

1.2	System Architecture	2

1.3	System Operation	3

1.4	System Features	3

1.5	Typical Applications	6

1.6	Available SeaBat T51-R Configurations	6

2	Safety Precautions	7

2.1	Operator Safety	7

2.2	Equipment Safety	7

2.2.1	Safe Handling	7

2.2.2	Cleaning and Maintenance	8

2.3	Corrosion Avoidance	10

2.4	Electrical Isolation	11

2.5	Exposure to Sunlight	11

2.6	Safe Disposal of Waste (WEEE Directive)	11

2.7	CE Marking	12

3	System Startup and Shutdown	13

3.1	Chapter Overview	13

3.2	System Startup Instructions	14

3.3	Starting the Sonar User Interface	14

3.4	System Shutdown Instructions	16

3.5	System Wake Up	17

4	Active Sonar Usage	18

4.1	Chapter Overview	18

4.2	Enhancing Sonar Performance	19

4.2.1	Correct Range Setting	19

4.2.2	Transmit Power	20

4.2.3	Ping Rate	21

4.2.4	Receiver Gain	21

4.2.5	Frequency Selection	22

4.2.6	Pulse Length	23

4.2.7	Pulse Type	23

4.2.8	Vari-Swath	23

4.2.9	Use of Beam Mode	27

4.2.10	Use of Gates	29

4.2.11	Maximizing Quality of Backscatter for Snippets Data	32

4.3	Enhancing System Efficiency	32

4.3.1	Multi-Detect Operation	32

4.3.2	Tracker Autopilot	33

4.3.3	Roll Stabilization	33

4.3.4	Head Tilt	34

4.3.5	Vertical Detection Mode (VDM)	35

4.4	Compensation for Environmental Conditions	35

4.4.1	Loss of Sound by Absorption	36

4.4.2	Loss of Sound by Spreading	37

4.4.3	Sound Velocity	38

4.5	Limitations	39

4.5.1	Cavitation	39

4.5.2	Machinery Noise	40

4.5.3	Interference from Other Sonar Systems	40

4.5.4	Speed	40

4.5.5	Air Bubbles	41

4.5.6	Environment	41

5	System Performance and Data Types	42

5.1	System Components	44

5.1.1	Rack-mounted Sonar Processor (RSP)	44

5.1.2	EM7222 Receiver Unit	45

5.1.3	TC2186 Projector Unit	45

5.2	Selectable Beam Density	45

5.2.1	Beam Modes	46

5.2.2	Equi-Distant, Equi-Angle, and Intermediate Modes	46

5.3	Data Types	48

5.3.1	Bathymetry Data	48

5.3.2	Side-Scan Data	48

5.3.3	Snippets Backscatter	49

5.4	Real-Time Pipe Detection and Tracking	51

5.5	Multi-Detect	51

5.6	Vertical Detection Mode	52

5.7	X-Range Performance	52

5.8	Quality	53

5.8.1	Brightness Test	53

5.8.2	Colinearity Test	54

5.8.3	Quality Filter	54

5.9	Uncertainty Estimation	54

5.10	Hydrodynamic Drag	55

5.10.1	Drag Simulation Setup	55

5.10.2	Drag Simulation Conclusions	55

5.10.3	Drag Simulation Results	56

5.11	Ping Rate Table	58

Appendix A	Troubleshooting	59

A.1	System Error and Warning Messages	59

A.2	Error Messages	59

A.2.1	General System Startup Errors	59

A.2.2	Data Reading Errors	60

A.2.3	System Resource Errors	60

A.2.4	Input Signal (Timing and Triggers) Errors	60

A.2.5	Normalization Errors	61

A.3	Warning Messages	61

A.3.1	Configuration Warnings	61

A.3.2	Data Processing Warnings	61

A.3.3	BITE Data Warnings	61

A.3.4	Input Signal Warnings	62

A.3.5	System Stability Warnings	62

A.3.6	Data Subscription Warnings	62

A.3.7	Network Components Warnings	62

A.3.8	External Trigger Warnings	62

A.4	Error Messages – Master List in Alphabetical Order	63

A.5	Warning Messages – Master List in Alphabetical Order	66

Appendix B	Installation Instructions	68

B.1	Components List	68

B.2	Installation Checklist	69

B.3	System Interfacing	70

B.3.1	Interfacing Diagram for Standard Configurations	70

B.3.2	Interfacing Diagram for Configurations with Integrated INS	72

B.4	Dry-End Installation	75

B.4.1	Rack-mounted Sonar Processor (RSP)	75

B.4.2	Sonar Processor Mounting	75

B.4.3	Sonar Processor Cable Connections	76

B.4.4	Rack-mounted Sonar Processor Power Button and BITE LED	78

B.5	Auxiliary Sensors	78

B.5.1	Standard SeaBat T51-R Configurations	78

B.5.2	SeaBat T51-R Configurations with Integrated Applanix AP+ INS	79

B.5.3	SeaBat T51-R Configurations with Integrated INS	79

B.5.4	Sound Velocity Input	80

B.5.5	Time Input	80

B.5.6	Time Output (for optional INS configurations)	80

B.5.7	Motion Input	81

B.5.8	Motion Output (for optional INS configurations)	81

B.5.9	Standard Serial Ports	81

B.6	SeaBat T51-R Cables	82

B.7	Wet-End Installation	83

B.7.1	EM7222 Receiver Unit (PN 1010473-DK)	83

B.7.2	TC2186 Projector Unit (PN 1010498-DK)	83

B.7.3	Wet-End Cable Connections	84

B.7.4	Sonar Head Assembly Orientation	84

B.7.5	Sonar Reference Point	85

B.7.6	Wet-End Mounting	85

B.7.7	Hardware	86

B.8	Remote Operation	87

B.9	System Check	88

B.9.1	Configuration	88

B.9.2	Receiver and Projector Test	89

B.9.3	Receiver Test (performed in air or water)	89

B.9.4	Projector Test (always performed in water)	89

Appendix C	Integrated INS – Applanix AP+	90

C.1	Access to the AP+	90

C.2	GNSS Antenna Types	90

C.3	Set Up the GAMS Baseline	91

C.4	INS Remote Settings	92

C.4.1	Graphic Prerequisites	93

C.4.2	Reference	93

C.4.3	IMU Remote	93

C.4.4	GNSS Antenna	94

C.4.5	Sensor 1 / Sensor 2	94

C.4.6	Heave Definition	95

C.5	I/O Configuration	95

C.5.1	Time (NMEA ZDA Ethernet UDP)	96

C.5.2	Motion for Roll Stab (NMEA PASHR Serial COM2)	96

C.5.3	Trimble GSOF 65 (Sensor 1) Output for Teledyne PDS	97

C.5.4	True Heave Output for Teledyne PDS	97

C.6	GAMS Calibration	97

C.7	Logging POSPac Data	99

Appendix D	PPS In and Trigger In	101

D.1	Electrical Specifications for PPS In and Trigger In	101

D.2	Typical PPS In and Trigger Performance	101

D.3	Hints and Advice	102

Appendix E	Data Rates	103

Appendix F	Network Configuration	104

F.1	System Network Configuration	104

F.2	Change IP Configuration	104

Appendix G	Configuration of Integrated INS (optional)	106

G.1	POS MV – How to Perform a GAMS Calibration	106

G.1.1	MV-POSView – Standby Mode	106

G.1.2	GAMS Parameter Setup	106

G.1.3	Positioning Quality	107

G.1.4	GAMS Calibration – Procedure	108

G.2	POS MV – How to Set Up for a Multibeam Survey	112

G.2.1	Conventions	112

G.2.2	Coordinate Frames	112

G.2.3	Offsets – Example	113

G.2.4	Alignments – Example	114

G.2.5	MV-POSView Software	114

G.2.6	Installation Parameters – Lever Arms & Mounting Angles	115

G.2.7	Installation Parameters – Sensor Mounting	116

G.2.8	Installation Parameters – Tags, Multipath &  AutoStart	117

G.2.9	Installation Parameters – Heave Filter	117

G.2.10	POS MV Output– Serial Data	117

G.2.11	POS MV Output– Ethernet Realtime	120

G.2.12	Acquisition Software	120

G.2.13	Patch Test Results	120

G.2.14	POS Statistics	121

Appendix H	Interfacing Sensors To Teledyne PDS	122

H.1	Establishing a Connection to the RSP	122

H.2	Interfacing to Teledyne PDS	124

H.2.1	Configuring the Sonar UI to Output Data	124

H.2.2	Configuring Teledyne PDS	125

Appendix I	Reference Documentation	128

I.1	SeaBat T51-R Documentation	128

I.2	SeaBat T51-R Design Documents	129

I.2.1	Receiver, EM7222 Outline	130

I.2.2	Projector, TC2186 Outline	131

I.2.3	T51 Bracket Installation: Receiver EM7222 and Projector TC2186	132

I.2.4	T51 Bracket Installation with Fairing (optional)	133

I.2.5	T51 Bracket Installation: Sonar Reference Point	134

I.2.6	T51 Bracket Installation with IMU Reference Point	135

I.2.7	Cable Assembly, Rack-mounted Sonar Processor to Receiver	136

I.2.8	Cable Assembly, Rack-mounted Sonar Processor to Projector	137

I.2.9	Rack-mounted Sonar Processor	138

I.2.10	Rack-mounted Sonar Processor with INS (optional)	139

I.2.11	Rack-mounted Sonar Processor with Applanix AP+ INS (optional)	140

I.2.12	IMU Outline (optional)	141

I.2.13	IMU Motion Reference Points on T-Series Brackets (optional)	142

I.2.14	IMU Cable	143

I.2.15	Antenna Outline	144

I.2.16	Antenna Cable	145

I.2.17	T51 Dual Head Bracket Installation (optional)	146

I.2.18	T51 Dual Head Bracket Installation with Fairing (optional)	147

I.2.19	T51 Dual Head Bracket Installation: Sonar Reference Point (optional)	148

I.2.20	T51 Dual Head Bracket Installation with IMU Reference Point (optional)	149

I.2.21	Cable Assembly, Sonar Processor to SVP 70 (optional)	150

I.3	7kCenter Licenses – Copyright Information	151

Appendix J	SeaBat T51-R System Options	156

J.1	Extended Warranty Contract	156

J.2	Fairings	156

J.3	Integrated Inertial Navigation System	156

J.4	SeaBat Cables	156

J.5	Service Level Agreement	157

J.6	Sound Velocity Probe	157

J.7	System Integration and Training	157

J.8	Teledyne PDS Data Acquisition Software	158

Appendix K	Standard Warranty Information	159

K.1	Limited Warranty	159

K.2	Warranty Exclusions	159

K.3	Warranty Disclaimer	159

K.4	Servicing During Warranty Period	159

Appendix L	Support and Service	160

L.1	Teledyne RESON Support	160

L.2	Equipment Service Procedure	160

L.2.1	Policy for Servicing Equipment	160

L.2.2	Returning Equipment for Service	161

List of Figures

Figure 1: SeaBat T51-R System Configuration	2

Figure 2: SeaBat T51-R System Configuration with Integrated INS (optional)	2

Figure 3: Processor Filter Location	10

Figure 4: Range Setting Diagrams	19

Figure 5: Range Setting – Too Long	19

Figure 6: Range Setting – Too Short	20

Figure 7: Range Setting – The Exception	20

Figure 8: Example of Reduced Swath Width	25

Figure 9: Horizontal Swath Steering	26

Figure 10: Outer Beams for the Different Beam Modes with a Tilted Sonar Head	27

Figure 11: Graphic Illustration of FlexMode	29

Figure 12: Depth Gates, With No Roll	30

Figure 13: Depth Gates, With Roll	30

Figure 14: Depth Gates, With Roll and Motion Sensor Input	31

Figure 15: Depth Gate Tilt	31

Figure 16: Submerged Oil Detected by Multi-Detect	32

Figure 17: Swath under Zero Roll Conditions	33

Figure 18: Laterally Displaced Swaths under Roll Conditions	34

Figure 19: Swath under Roll Stabilization	34

Figure 20: Head Tilt Functionality	35

Figure 21: Head Tilt for Harbor Inspection	35

Figure 22: Acoustic Absorption Reference	37

Figure 23: Typical Sound Velocity Values	38

Figure 24: Rack-mounted Sonar Processor	44

Figure 25: EM7222 Receiver Unit	45

Figure 26: TC2186 Projector Unit	45

Figure 27: Beam Density	46

Figure 28: Beam Spacing vs. Beam Angle	47

Figure 29: Beam Spacing	47

Figure 30: Across-Track Sounding Location on a Flat Bottom	48

Figure 31: Graphic Illustration of Snippets Data	50

Figure 32: SeaBat Real-Time Pipe Detection and Tracking	51

Figure 33: SeaBat FlexMode	51

Figure 34: Standard Detection	52

Figure 35: Multi-Detect with enhanced detail on Quay Wall	52

Figure 36: Multi-Detect	52

Figure 37: Vertical Detection of Quay Wall	52

Figure 38: Brightness Test	53

Figure 39: Colinearity Flags	54

Figure 40: T50 Sonar Head Assembly in Bracket with Optional Front Fairing	56

Figure 41: Flow at 10 knots around Single T50 with Fairing	56

Figure 42: Flow around Single T50 with Fairing, Cross Section of Symmetry Plane	56

Figure 43: Flow at 10 knots around Single T50 without Fairing	57

Figure 44: Flow around Single T50 without Fairing, Cross Section of Symmetry Plane	57

Figure 45: T51-R, interfacing diagram	71

Figure 46: T51-R with integrated Applanix AP+ INS, interfacing diagram	73

Figure 47: T51-R with integrated INS, interfacing diagram	74

Figure 48: RSP Standard Configuration – Cable Connections	76

Figure 49: RSP Applanix AP+ INS Configuration (optional) – Cable Connections	76

Figure 50: RSP INS Configuration (optional) – Cable Connections	77

Figure 51: SVP, Motion, and Time Standard Connections	78

Figure 52: SVP and Motion Connections with Integrated AP+ INS (optional)	79

Figure 53: SVP, Motion, and Time Connections with Integrated INS (optional)	79

Figure 54: SeaBat T51-R System Cables	82

Figure 55: Configuration Check via Sonar UI Service Menu	88

Figure 56: Required PPS Pulse Width	101

Figure 57: PPS In Delay	102

Figure 58: PPS/Trigger Signal	102

Figure 59: Tate/Bryant Co-ordinate/Rotation System	112

Figure 60: Vessel Outline – Example	113

Figure 61: IMU misalignment – Example	116

Figure 62: System Documentation Link (Example Only)	128

Figure 63: Receiver, EM7222 Outline	130

Figure 64: Projector, TC2186 Outline	131

Figure 65: T51 Bracket Installation: Receiver EM7222 and Projector TC2186	132

Figure 66: T51 Bracket Installation with Fairing (optional)	133

Figure 67: T51 Bracket Installation: Sonar Reference Point	134

Figure 68: T51 Bracket Installation with IMU Reference Point	135

Figure 69: Cable Assembly, Rack-mounted Sonar Processor to Receiver	136

Figure 70: Cable Assembly, Rack-mounted Sonar Processor to Projector	137

Figure 71: Rack-mounted Sonar Processor	138

Figure 72: Rack-mounted Sonar Processor with INS (optional)	139

Figure 73: Rack-mounted Sonar Processor with Applanix AP+ INS (optional)	140

Figure 74: IMU Outline (optional)	141

Figure 75: IMU Motion Reference Points on T-Series Brackets (optional)	142

Figure 76: IMU Cable	143

Figure 77: Antenna Outline	144

Figure 78: Antenna Cable	145

Figure 79: T51 Dual Head Bracket Installation (optional)	146

Figure 80: T51 Dual Head Bracket Installation with Fairing (optional)	147

Figure 81: T51 Dual Head Bracket Installation: Sonar Reference Point (optional)	148

Figure 82: T51 Dual Head Bracket Installation with IMU Reference Point (optional)	149

Figure 83: Cable Assembly, Sonar Processor to SVP 70 (optional)	150

List of Tables

Table 1: SeaBat T51-R Feature Pack	3

Table 2: SeaBat T51-R Available Configurations	6

Table 3: Incorrect Range Settings – Examples and Solutions	19

Table 4: TVG Formula	22

Table 5: SeaBat T51-R System Performance	43

Table 6: Beam Modes with No. of Beams and Swath Coverage	46

Table 7: Ping Rate Table	58

Table 8: SeaBat T51-R Components	68

Table 9: RSP Technical Specifications	75

Table 10: RSP Port “SVP Com 3”	80

Table 11: RSP Port “Time In Com 1”	80

Table 12: RSP Port “Time Out Com 1” for optional INS	80

Table 13: RSP Port “Motion In Com 2”	81

Table 14: RSP Port “Motion Out Com 2” for optional INS	81

Table 15: RSP Ports “Com 4”, “Com 3-5” for optional INS	81

Table 16: EM7222 Receiver Unit Technical Specifications	83

Table 17: TC2186 Projector Unit Technical Specifications	83

Table 18: Applanix AP+ and Teledyne Axis Conventions	91

Table 19: Offsets from IMU Reference Mark to IMU Sensing Center	94

Table 20: PPS In and Trigger In Signals	101

Table 21: Theoretical Max. Data Rates – SeaBat T51-R Single Head	103

Table 22: IP Configuration for RSP	104

Table 23: Sensor Offsets	113

Table 24: Sensor Alignments	114

Table 25: POS MV Output Serial Data – COM Ports 1-2	118

Table 26: SeaBat T51-R Design Documents	129

Table 27: LAPACK – Copyright Information	151

Table 28: WinPCap – Copyright Information	152

Table 29: SVP Technical Specifications	157

# Introduction

## System Overview

The SeaBat T51-R sonar system is a part of the leading SeaBat T-series product range, engineered from the ground up to evolve with your business. SeaBat T51-R is built on the shoulders of the renowned SeaBat T50 – but with four times the resolution – and designed for very fast mobilization on any type of survey vessels, securing minimal interfacing and low space requirements. The series is available in various configurations.

| The SeaBat T51-R: Combined with a very compact Rack-mounted Sonar Processor (RSP), the SeaBat T51-R produces unprecedented clean data, providing faster operational surveys and reduced processing time. - The SeaBat T51-R comes with the revolutionary high-frequency band of 700-800kHz which yields the highest level of detail while maintaining seven times water depth survey efficiency. - In addition, SeaBat T51-R comes with a flexible 350-430kHz lower frequency range for the surveys requiring extended range performance. |  |
| --- | --- |
| Integrated Inertial Navigation System: For SeaBat T51-R, the rack-mounted sonar processor is available in an extended version for dual-head sonar head configurations and an optional industry-leading fully integrated Inertial Navigation System for accurate sensor time tagging and motion stabilization. |  |

For available configurations, see section 1.6 below.

For available options, see Appendix J SeaBat T51-R System Options.

## System Architecture

The SeaBat T51-R multibeam sonar systems consist of a rack-mounted sonar processor, a receiver unit, and a projector unit. User displays and system control originate within the sonar processor.

For technical specifications, see Appendix B Installation Instructions.

Figure 1: SeaBat T51-R System Configuration

The SeaBat T51-R multibeam sonar systems are also available with an optional state-of-the-art built-in and fully integrated Inertial Navigation System (INS).

Figure 2: SeaBat T51-R System Configuration with Integrated INS (optional)

For additional options, see Appendix J SeaBat T51-R System Options.

## System Operation

The projector unit of the sonar head assembly transmits a pulse of acoustic energy, which travels through the water and is reflected by the seafloor or any objects in its path. The reflected signal is received by the receiver unit of the sonar head assembly, digitized by internal electronics, and sent to the topside rack-mounted sonar processor for beamforming and processing. The sonar processor formats a digital output to be used by a peripheral data processing system while functioning as the interface between the operator PC and the sonar system. Available data packages are documented in the Data Format Definition Document (see Appendix I Reference Documentation).

The CFE laptop/PC running the RESON Sonar User Interface receives beam data, water column data, bottom detection data etc. from the RSP and generates the wedge, water column, snippets, and side-scan displays as well as other displays for sensor monitoring and general system Built-In Test Environment overview.

## System Features

Table 1: SeaBat T51-R Feature Pack

| Feature | Base | Optional |
| --- | --- | --- |
| Channel normalization |  |  |
| Constant seafloor spacing |  |  |
| FlexMode |  |  |
| Head tilt |  |  |
| Manual gates |  |  |
| Max. 1024 beams |  |  |
| Multi-detect |  |  |
| On the fly frequency selection |  |  |
| On the fly beam mode selection |  |  |
| Pipe detection and tracking |  |  |
| Quality filter |  |  |
| Sonar User Interface |  |  |
| Roll stabilization |  |  |
| Selectable beam density |  |  |
| Tracker |  |  |
| Uncertainty output |  |  |
| Variable frequency |  |  |
| Vari-swath |  |  |
| Vertical Detection Mode |  |  |
| Water column visualization |  |  |
| X-Range (FM) |  |  |
| Integrated INS (incl. external antennas and IMU) |  |  |
| Integrated Teledyne PDS |  |  |

NOTE

The Base feature set can be upgraded after installation.

The main features of the SeaBat T51-R sonar systems are summarized below.

1PPS pulse handling and time stamping: Data time stamping is referenced to an incoming 1PPS and associated time message.

Adaptive gating: Three different gate settings are available for application to the bottom detection process: No Gates, Absolute Gates, and Adaptive Gates.

Beam data recording: The SeaBat T51-R sonar systems are capable of producing a stream of operator configurable beam data that may be written to a shared network drive on either the CFE laptop/PC’s local hard drive or an optional NAS (Network Attached Storage) RAID hard disk array. Data logging software controls recording, playback, and data storage functions.

Beam distribution: The three standard beam spacing choices are equiangular, intermediate, and equidistant spacing. In addition, the optional FlexMode provides settings for alternative beam distribution. All options are dynamically focused.

Beamforming: The sonar processor performs initial signal processing, a combination of time delay and phase rotation beamforming.

Bottom detection: The special bottom detection algorithm uses a combination of center-of-energy and phase-zero crossing algorithms to detect the bottom in each individual beam with the highest possible accuracy.

Built-In Test Environment (BITE): BITE is an integral part of the SeaBat T51-R sonar systems, monitoring the status of internal subsystems and the wet-end units.

Compressed water column data: With this user configurable feature, it is possible to make continuous recordings of water column backscatter data at significantly reduced data volumes.

Constant seafloor spacing: This feature yields equidistant spacing between beams, even on a sloping seafloor.

Data output: The SeaBat T51-R sonar systems export the standard hydrographic records of bathymetry, side-scan, and snippets as well as water column data and other data for more scientific applications.

Depth rating: The SeaBat T51-R sonar system wet end and cables are depth-rated to 50 meters.

External trigger: The system may be triggered using a standard 7k Center software remote command or by a TTL incoming trigger pulse.

Fast update: The entire 140° (or 170°) swath is illuminated with a single transmission at an update rate of up to 50Hz (range selection dependent), maintaining 100% bottom coverage, even at high speed.

Hardware sync pulse out: A TTL pulse (synchronous with the transmit pulse) is available from the sonar processor.

Head tilt: This feature makes it possible to operate with a physically rotated sonar head, which allows for very flexible and specialized applications of the SeaBat T51-R sonar systems.

Integrated INS: An optional Inertial Navigation System (motion + position) uses GPS for position and time, and IMU for measuring roll, heave, and acceleration.

Multi-detect: Providing up to five detections within each beam, including the full water column, this feature allows the surveyor to capture enhanced details from a single survey line over a complex feature, even in noisy environments.

On the fly beam mode selection: Choice of beam mode (spacing and number of beams) can be done without interrupting the survey.

On the fly frequency change: The SeaBat T51-R multi-frequency systems may be operated at the full range of frequencies without interrupting the survey.

Operator selectable operating mode: With five different options, the operator is able to select a beam mode appropriate to the application thereby optimizing sonar operation. Equi-Angle and Intermediate modes yield a swath of 170º, while the other three modes yield a swath of 140º or 150º in Equi-Distant mode.

Real-time pipe detection and tracking: This SeaBat feature automatically detects pipes within the multibeam swath, displays them to the user in real time during survey, and sends the pipe positions to the data acquisition software. The feature builds on FlexMode for generating a high-density profile of the pipe. The majority of the beams are steered to the center of the detected pipe to track the pipe regardless of its location on the wedge display.

Real-time QC display: The SeaBat T51-R sonar systems generate a real-time display with an intensity image on the CFE monitor. Overlaid on this intensity image is the bottom detection display, giving the operator an excellent QC (quality control) tool.

Resolution: Dynamic, focused beamforming maintains ultra-high resolution, even at short ranges. A receiver sample rate of 34/66kHz (34kHz for dual-head configurations) provides unprecedented range resolution and complements the narrow beam, thus providing the user with ultimate data quality and resolution.

Selectable beam density: This feature allows the surveyor to generate optimized data volumes tailored to specific survey specifications by providing a choice of how many beams to form ranging from the maximum capability of the system down to just 10. The configuration can be changed on the fly.

Steerable swath: This feature allows swath to be steered electronically.

Time stamping of sensor data: Time stamping of sensor data interfaced to the Rack-mounted Sonar Processor (position, motion, and heading) is referenced to an incoming 1PPS and associated time message (either serial or network), ensuring that each data is time stamped to UTC to better than 1 millisecond, prior to broadcast over standard Ethernet to the survey laptop/PC.

- With the optional built-in, fully integrated INS, all time tagging will be handled by the INS.

Tracker: This feature is a complete automated mode of operation where the system sets sonar control parameters, including range, gain, power, and swath width based on the quality of the data being collected.

Uncertainty output: Uncertainty is computed for each sounding and presented graphically on various UI (user interface) displays.

Variable swath coverage: This feature gives the operator control of the swath coverage. It includes swath width settings ranging from 45º to 150º in Equi-Distant mode or 170º in Equi-Angle mode in increments of 5º as well as a choice of beam density mode.

Vertical detection mode: Vertical structures pose an acoustically challenging environment for normal bottom detection methods. The SeaBat Vertical Detection Mode employs a unique method to produce improved detections of vertical structures such as quay walls. Any damage, anodes, or other smaller objects are also better detected with this feature.

Water column visualization: Water column backscatter may be viewed in both the across-track view (the sonar wedge) and in an along-track view, which allows the surveyor to see a history of water column backscatter. Combined with the rolling water column recording buffer, water column features of interest may be viewed and recorded after they have been surveyed.

X-Range: X-Range is a software and hardware feature using frequency modulated (FM) transmit pulses which provides extended range performance (across all frequencies) and significantly improves system immunity to external noise.

## Typical Applications

The SeaBat T51-R is based around the sonar processor. The powerful scalable feature set provides unmatched utility across a wide range of applications including:

| As-built surveys Breakwater mapping Coastal mapping Environmental research Habitat mapping Harbor mapping Hydrographic surveys to IHO Special Order and to U.S. Army Corps of Engineers Special Order Oil detection | Pipeline inspection surveys Pipeline route surveys Pre- and post-dredging and condition surveys Revetment surveys River surveys Search and recovery Site clearance surveys Underwater inspection and object location |
| --- | --- |

## Available SeaBat T51-R Configurations

Table 2: SeaBat T51-R Available Configurations

| PN | System | Frequency (kHz) | Depth rating (m) |
| --- | --- | --- | --- |
| 88017090 | SYS, SEABAT T51-R, Rackmount, single head, no INS SYS | 350-430 and 700-800 | 50 |
|  | OPTION: change to RSP w/INS support |  |  |

# Safety Precautions

WARNING!

If the equipment is used in a manner not specified by the manufacturer, the protection provided by the equipment may be impaired. Teledyne RESON assumes no liability if this product is operated in an unsafe manner.

CAUTION

Use of the equipment in a manner not specified by the manufacturer may affect warranty situation (see Appendix K Standard Warranty Information).

## Operator Safety

The SeaBat T51-R system should be handled with attention to operator safety as well as protection of the hardware components. General precautions include:

DO NOT connect or disconnect cables with the power on.

DO NOT attempt to open and service the projector or receiver units.

DO NOT attempt to open and service the sonar processor without specific instructions from Teledyne RESON.

DO NOT operate the system while divers are in the water and closer than two meters from the projector and receiver units. High-powered ultrasonic energy can be hazardous to human physiology.

NOTE

Observe that the mains plug is the off switch for the rack-mounted sonar processor.

For safety reasons the SeaBat™ must be connected to protective earth. The protective earth pin in the mains appliance inlet is the protective conductor terminal of SeaBat™.

| WARNING! |  |
| --- | --- |
|  | This icon on any of the SeaBat components implies hazardous voltage on output connection. |

## Equipment Safety

Each SeaBat T51-R component is sufficiently robust for shipboard storage while in its own transit case or shipping box, which may include handles and tie-down rings. Use appropriate discretion in lifting, as the case may be awkward and heavy.

### Safe Handling

To ensure safe handling of the equipment:

Inspect each transit case or shipping box for physical damage prior to opening, and each component for physical damage before installation.

Use original shipping boxes to provide adequate packaging and shock absorption when shipping or storing the equipment.

Do not drop the equipment.

Do not connect or disconnect cables at the rear of the processor or from the sonar head assembly while the processor is running. This can damage the internal workings.

Do not bend cables beyond the recommended limits.

Do not place the sonar head in an area where the face may be scratched or damaged. If an array must be placed on the ground, it should be placed on a flat surface and covered with a protective padded sheet.

Do not place liquids on or near the equipment.

Ensure that the equipment is properly secured before putting out to sea.

Unused connectors should be sealed with dummy plugs.

Do not exceed operating and storage temperature limits.

Allow clearance around the sonar processor to provide adequate air circulation.

CAUTION

The projector must always be tested in water, whereas the receiver may be tested in air.

### Cleaning and Maintenance

#### Projector and Receiver

The acoustic window molded over the front of the projector and receiver arrays is made of high-density polyurethane.

Follow these recommendations to protect the arrays while in use or in storage.

For cleaning, use a mild, nonabrasive soap, fresh lukewarm water, and a soft brush.

Rinse the sonar with fresh water after every use.

Do not clean the array or sonar by using hot or cold pressure washers. The high-velocity water spray will cut through the polyurethane or damage the interface between the polyurethane and the sonar metal housing.

To scrape away marine growth from the arrays, use only a wooden tool.

Do not apply any antifouling paint on the acoustic window unless the type is recommended by Teledyne RESON. The antifouling paint will affect the acoustic performance. Teledyne RESON recommends:

For high-frequency transducers (>100kHz) use ClearSignal®.

For low-frequency transducers (<100kHz) use SigmaGlide.

NOTE

As these products are subject to technical modifications, it is recommended that the most up to date Product Data Sheet and Material Safety Data Sheet be acquired prior to commencing the coating work.

| CAUTION |  |
| --- | --- |
| A long range of chemicals will destroy the polyurethane, which include, but are not limited to: |  |
| Acids Automotive brake fluid Base Grease Halogenated hydrocarbons (carbon tetrachloride, trichloroethylene) Hydraulic fluids | Ketones (MEK, acetone) Nitro hydrocarbons (nitrobenzene, aniline) Oil Ozone Petrol Phosphate ester |

#### Wet-End Connectors

Use the wet connector lubrication kit (PN 1007252) to protect the wet-end connectors. Follow these guidelines:

Lubricate female connectors and dummy plugs with a silicone oil or equivalent (silicone oil based). Observe the instructions provided with the lubrication kit.

Lubrication intervals depend on mission type and frequency.

Lubricate every 5 matings.

Lubricate at regular intervals, even if the connectors are not mated/unmated, in particular for UUVs where the dive tends to rub and wash out the lubricant.

Lubricate twice a year when equipment is in storage.

DO NOT grease mating surface.

Avoid nicks and cuts around contacts, as these are the sealing surfaces.

Elastomers can be seriously degraded if exposed to direct sunlight or high ozone levels for extended periods of time.

Do not overtighten bulkhead nuts.

Grip main body of connector during mating or unmating. Do not pull on cable to disconnect.

Avoid sharp bends at cable entry to connector.

#### Sonar Processor Dust Filter

The processor is equipped with filters to prevent dust and other particles to enter the unit. The dust filters need to be changed at least once a year:

Figure 3: Processor Filter Location

- Unscrew the six screws holding the front cover plate.

- Lift off the front cover plate.

- Replace the three dust filters.

- Reassemble the processor.

OR

- Remove the dust filters and clean them with a mild, nonabrasive soap in lukewarm water.

- Put the dust filters back in their place in the processor.

CAUTION

Make sure that the filters are completely dry before placing them in the processor.

- Reassemble the processor.

#### Replacement of Battery on PC Motherboard

CAUTION

Battery to be changed by authorized personnel only. Noncompliance may affect warranty situation (see Appendix K Standard Warranty Information).

Use only the correct type of onboard CMOS battery as specified by the manufacturer. Do not install the onboard battery upside down to avoid possible explosion.

## Corrosion Avoidance

All metal parts of the SeaBat™ exposed to seawater are made of corrosion-resistant alloys. This is to ensure long operational life of the system. Combining these noble and corrosion-resistant metals with less noble metals in a saline environment will cause a cathodic reaction.

We therefore recommend electrical isolation between the sonar head and mounting structure.

CAUTION

Do not use sacrificial anodes on a SeaBat™ bracket. Sacrificial anodes lead to unwanted cathodic reaction and does not offer electrical isolation.

CAUTION

Damaged cables must be replaced (and not repaired). Otherwise the result may be corrosion and/or debonding. Noncompliance may affect warranty situation (see Appendix K Standard Warranty Information).

## Electrical Isolation

Electrical isolation for titanium components may be achieved by using nonconductive bushings, washers, and isolation plates to insulate them from the mounting brackets. Materials such as FR4, G-10, and G-Etronax EP11 are ideal, as they combine high mechanical strength with a complete electrical isolation. Unreinforced plastic materials such as PVC, Delrin, and polyamide will work from an electrical point of view but offer limited mechanical strength and have a tendency to relax over time causing interfaces to become loose.

Teledyne RESON recommends using the standard bracket delivered with the system. The standard bracket is made of high-strength fiber reinforced material and will therefore isolate the sonar head completely from the mounting arrangement. The thread inserts in the fiber bracket also offer a much easier way of isolation than washers and bushings.

CAUTION

If the sonar head is not isolated, galvanic corrosion will result in the mounting structure.

NOTE

Contact Teledyne RESON for additional information on schemes for electrical isolation.

## Exposure to Sunlight

Do not allow the projector and receiver arrays to sit in direct sunlight when not used, but cover the arrays with a bag. In addition to furthering corrosive effects, continued exposure to ultraviolet (UV) rays will damage the polyurethane array face.

## Safe Disposal of Waste (WEEE Directive)

The use of the crossed-out dustbin logo indicates that this product is subject to directive 2012/19/EU, known as the Waste from Electrical and Electronic Equipment Directive. The WEEE Directive specifies that used electric/electronic equipment may not be treated as household waste, but must be sorted separately for disposal. These rules concern all your Teledyne RESON products (computer, monitor, transducer, processor, probe, etc.) and all accessories (keyboard, mouse, etc.). By ensuring this product is disposed of correctly, you will help protect the environment.

For detailed information about the recycling of this product, please contact your local authority, your household waste disposal service provider, or Teledyne RESON Customer Support (reson-support@teledyne.com).

## CE Marking

This product carries the CE mark that indicates that it conforms to the essential health, safety, and environmental requirements of the relevant European Directives.

# System Startup and Shutdown

## Chapter Overview

This chapter includes detailed instructions on the following subjects (click the emphasized hyperlink to be taken to the section):

- System Startup Instructions: Consists of startup details.

Starting the Sonar User Interface: Consists of instructions on how to start the sonar system.

- System Shutdown Instructions: Consists of details on stopping the system.

NOTE

For information on installation (including configuration of hardware), see Appendix B Installation Instructions. For information on system configuration (including IO module configuration), refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation). For information on user interface navigation, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation). For information on the practical use of the sonar, see section 4 Active Sonar Usage.

## System Startup Instructions

After the SeaBat hardware has been installed (see Appendix B Installation Instructions), the system may be started. To do this, press the power buttons on the rack-mounted sonar processor and the CFE laptop/PC.

NOTE

The rack-mounted sonar processor power button will be red, when power is connected to the unit; it will turn yellow, when the system is turned on.

## Starting the Sonar User Interface

- Double-click the Sonar UI icon on the laptop/PC to start the Sonar User Interface.

- In the top right corner of the UI, click the sonar address drop-down menu (circled in red), and select the IP address of the connected RSP/sonar system from the list. The sonar system can be identified by its serial number (prefixed with SN) next to the IP address. The serial number can be found on the front of the RSP.

NOTE

The laptop/PC and RSP should be directly connected using an Ethernet cable to the LAN 1 or LAN 2 port of the RSP. (For further information, see appendix B.4.3 Sonar Processor Cable Connections.) By factory default, the RSP is set up with fixed IP addresses: 10.11.10.1 for LAN 1 and 10.11.10.2 for LAN 2. If the RSP is configured to acquire an automatic IP address (assigned by a DHCP server), but fails, the two Ethernet ports will revert to the static IP addresses 10.11.10.1 (LAN 1) and 10.11.10.2 (LAN 2), both with subnet mask 255.255.255.0, after typically up to 1m30s of trying to acquire DHCP contact. Make sure that the laptop/PC IP address is set to an unused IP address within the 10.11.10.XXX subnet, subnet mask 255.255.255.0. Recommended 10.11.10.3 for single-head systems Recommended 10.11.10.5 for dual-head systems. For details on IP address configuration, see Appendix F Network Configuration.

The laptop/PC is now connected to a running sonar system.

- Check the state of the BITE LED on the front of the RSP.

Red, green, or yellow solid is a sign that the sonar software is running and should be accessible.

No front light or a flashing red light is a sign that the sonar software is not running correctly.

Remove the power cable, wait 10 seconds, reconnect the power cable, and restart the RSP by pressing the power button.

- Select the settings appropriate for the present survey mission. Refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

## System Shutdown Instructions

CAUTION

Ensure that the rack-mounted sonar processor is shut down correctly.

- Click the Application Settings button in the top right corner of the Sonar UI.

- Select the Shutdown option in the Application Settings dialog.

- Click OK.

This will shut down the RSP and thereby terminate the connection to the RSP/sonar system from laptop/PC.

- Click the cross in the upper right corner of the Sonar UI to terminate the session.

The Sonar UI will resynchronize to the RSP, once the system is started again.

## System Wake Up

System can wake up via a Wake-On-Lan package. The feature uses Wake-On-LAN technology.

- Click the Application Settings button in the top right corner of the Sonar UI.

- Select the Wake up option in the Application Settings dialog.

- Click OK.

This will power up the RSP.

NOTE

The wake up function is only possible under the following conditions: - Sonar system has been powered up. - Sonar UI has been connected to the system (via the sonar address drop-down menu). - RSP has remained connected to mains power. - LAN 1 port is used.

# Active Sonar Usage

## Chapter Overview

This chapter includes detailed instructions on the following subjects (click the emphasized hyperlink to be taken to the section):

- Enhancing Sonar Performance: Consists of examples on how to use the UI controls to enhance sonar performance.

- Enhancing System Efficiency: Consists of examples on how to use the auto pilot, roll stabilization, and head tilt features.

- Compensation for Environmental Conditions: Consists of a thorough introduction into the environmental parameters that may affect sonar operation.

- Limitations: Consists of examples on the various physical limitations that may affect sonar operation.

NOTE

For information on starting, stopping, and setting up the SeaBat sonar system, see section 3 System Startup and Shutdown. For information on system configuration (including IO module configuration), refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation). For information on user interface navigation, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation). For information on installation, see Appendix B Installation Instructions.

## Enhancing Sonar Performance

### Correct Range Setting

Select the appropriate range scale to keep the horizontal bottom image at or above the widest part of the sonar wedge as visualized in Figure 4:

If the Range is set too short, the outer beams will be lost. If the Range is set too long, the results may be noisy data.

Figure 4: Range Setting Diagrams

Diagram A depicts a proper range setting.

Diagram B depicts an acceptable range setting for a swath that is beginning to fade due to the effects of signal loss in the outer swath.

Diagram C depicts an incorrect range setting. Notice that the still-useable ends of the swath are being cut off by the limits of the range setting.

Note that the range setting affects ping rate: the longer the range, the lower the ping rate and thus the fewer along-track pings. Higher range settings are required for tilted heads (see also section 4.3.4 below).

NOTE

The Range adjusts itself automatically when Tracker is selected (see also section 4.3.2 below).

For additional examples on incorrect range settings and possible solutions, see Table 3.

Table 3: Incorrect Range Settings – Examples and Solutions

Figure 5: Range Setting – Too Long

With the Range set too long, the receivers are “open” for too long and the system is recording long after the acoustic returns from the seabed have been received. Spurious background noise can be amplified, especially when TVG is used, and this can give noisier data. The ping rate is also reduced thereby reducing sounding density. This should not be used to artificially reduce the ping rate, rather the Range should be set correctly, and the ping rate should be capped using Max rate. Refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

Figure 6: Range Setting – Too Short

With the Range set too short, the outer beams formed are not long enough to reach the seabed. With no seabed in these beams, the bottom detection process can only lock onto the noise in the water column. An exception to this is if the seafloor to the outer parts of the swath is of no interest and the user wishes to concentrate on a smaller seafloor coverage thereby increasing the ping rate and hence sounding density. In this case, the user should reduce the swath coverage of the sonar. Refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

Figure 7: Range Setting – The Exception

If you are operating in relatively deep water to the limits of operation and the outer parts of the swath exceed the acoustic range of the sonar, the user should reduce the swath coverage to cover the seabed that is visible only. This will increase the ping rate and reduce noise picked up for these visible beams. Refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

### Transmit Power

For instructions on how to adjust this setting, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

NOTE

The Power adjusts itself automatically when Tracker is selected (see also section 4.3.2 below).

#### Transmit Power OFF

If the Power is set to OFF, the system will not ping, but it will continue to receive data provided that the max ping rate is not set to 0.

With the transmit power OFF, the sonar system will act as a passive sonar by only receiving data. This feature is very useful for monitoring the ambient water in order to assess the level of background noise and interference from the vessel itself.

#### Transmit Power and Pulse Type

Power level depends on pulse type:

In CW mode, typically set Power at full for the most robust bottom detection at all ranges.

In FM mode, the use of power may vary with range setting. Due to longer pulse lengths in this mode, more overall energy is transmitted into the water per ping than in CW mode.

Less power may be required at shorter ranges. Start with 200dB at minimum Range, and increase with Range.

### Ping Rate

Set the Max rate at 50p/s to ensure unhindered operation of the sonar. Note that this is only a ping rate limiter.

The ping rate is automatically set by the range setting (or safe duty cycle limit) (see also section 4.2.1 above).

NOTE

If Max rate is set to 0, the system will not ping or receive data. If Max rate is set to a value higher than the range-determined ping rate (or the duty cycle-determined ping rate), the sonar will ping at the rate determined by the range setting (or calculated duty cycle). If Max rate is set to a value lower than the range-determined ping rate, the sonar will ping at the rate specified by Max rate.

In some instances it may be beneficial to force a lower ping rate than the normal rate associated with a particular range. Examples of this would be instances when too much data is being received at short ranges; when the vessel is moving at a very slow rate and the operator wants to reduce the quantity of data generated; or when the sonar is operating in a high-reverb environment where echoes of the previous ping are contaminating the current ping.

For instructions on how to adjust this setting, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

### Receiver Gain

The Gain setting adjusts the amount of receiver gain applied to the returned signal. Typically set between 20dB and 40dB.

NOTE

The Gain adjusts itself automatically when Tracker is selected (see also section 4.3.2 below).

The SeaBat system utilizes Time-Varied Gain (TVG), which counteracts the diminishing echo strength of signals that originate from farther away. The TVG data contains the applied gain at each range (sample) and is available in record 7010. Refer to the Data Format Definition document for information (see Appendix I Reference Documentation).

The generation of the TVG curve is based on the formula in Table 4.

Table 4: TVG Formula

|  |  |  | GL | for R < RL |
| --- | --- | --- | --- | --- |
| TVG = |  |  | G + SP [log10(R) -1] + 2(α/1000)R | for R ≥ RL |
|  |  |  | 83 | for R ≥ RM |
| Where: |  |  |  |  |
| α = G = GL = R = RL = RM = SP = |  | Absorption loss in dB/km Extra gain from RxGain menu item An internally calculated gain applied until the TVG begins Range in meters Lowest range (in meters) from where the TVG formula operates Range in meters where the gain has reached highest possible value (which is 83dB) Spreading loss coefficient |  |  |
| NOTE |  |  |  |
| The values of RL, GL, and RM are not accessible, but RL is in generally limited to a few meters from the sonar. It is recommended to subscribe for record 7010 in order to obtain the TVG. |  |  |  |

For details on environmental conditions such as absorption and spreading, see section 4.4 below.

For instructions on how to adjust this setting, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

### Frequency Selection

- 700-800kHz for highest resolution at short ranges.

- 350-430kHz for greater range performance.

For general hydrographic surveys, the switch from 700-800 kHz to 350-430kHz may be necessary to maintain optimal swath coverage between depths of 60m and 100m depending on conditions.

Variable frequency selection for CW pulses can be made in steps of 10kHz. The optional add-on X-Range license will enable FM chirp pulses for longer ranges than standard CW pulse type.

NOTE

Adjusting the frequency affects beam width. Reducing the frequency produces wider beams, increasing the frequency produces narrower beams.

SeaBat also includes a Ping Sequencing option, whereby the system cycles through different frequencies each time. At each ping it produces standard outputs: bathymetry, snippets backscatter, or water column backscatter.

For instructions on how to adjust this setting, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

### Pulse Length

For instructions on how to adjust this setting, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

NOTE

The Pulse Length adjusts itself automatically when Tracker is selected (see also section 4.3.2 below).

#### Pulse Length and Image Resolution

For a given power setting, a narrower pulse length will provide a higher resolution at a shorter range. A wider pulse length will provide better range but a lower resolution image.

In certain conditions, for example in turbid waters, the narrower pulse length may not operate well. In such circumstances, the wider pulse length may give better short-range performance. This may also be achieved by changing from higher to lower frequency operation.

#### Pulse Length and Pulse Type

The pulse length setting depends on pulse type (see section 4.2.7 below).

- In CW mode, typically set Pulse Length at 30µs (range is 30µs-300µs), and increase with Range. The shorter the pulse in CW mode, the higher the resolution. The longer the pulse, the greater the range performance but the lower the resolution.

- In FM mode, typically set Pulse Length at 1ms (range is 300µs-10ms), and increase with Range. Note that the length of the FM pulse does not affect resolution in the same way as it does for CW mode. At very short ranges, a long pulse length can limit the ping rate due to the duty cycle protection mode.

### Pulse Type

CW (Continuous wave) is the traditional pulse type for multibeam survey.

Shorter pulses result in high resolution, but suffer in range performance.

Longer pulses increase range performance at the expense of resolution.

FM (Frequency modulation or X-Range) is selected for greater range performance, FRDH operation, and instances where an increase in noise immunity may be required.

FM allows much longer pulses, and thus more energy, to be transmitted into the water, resulting in greater range performance whilst maintaining full resolution.

FM also provides better immunity to external noise.

NOTE

Depending on purchased features, only CW mode may be available. X-Range is an optional add-on feature (see Table 1).

For instructions on how to adjust this setting, refer to the Sonar UI User Manual (see Appendix I Reference Documentation).

### Vari-Swath

Vari-swath is a feature that allows the operator to select a swath width appropriate to the application in addition to electronically steering the swath.

Vari-swath is available with all beam modes and can be used in conjunction with roll stabilization and FlexMode. For details on the use of the various beam modes, see section 4.2.9 below.

For instructions on how to adjust this setting, refer to the Sonar UI User Manual (see Appendix I Reference Documentation).

#### Coverage Angle

The benefits of the variable swath feature are twofold:

Increased sounding density in a small sector may be used when tracking pipes or other linear features, where the operator is not interested in data outside a defined sector. The increased sounding density helps define small targets.

A reduced swath width allows an increase in ping rate for any given depth.

#### Beam Reduction

- With beam reduction enabled, only the beams necessary to fill the given swath are formed.

- With beam reduction disabled, all beams are compressed into the selected coverage angle.

Figure 8: Example of Reduced Swath Width

#### Horizontal Steering

This setting allows the operator to steer the swath independently of the sonar head physical angle. The extent of steering is determined by the angle selected. The limitation is that the outer beam cannot be steered past 5° from the horizontal (see Figure 9). When the swath is steered this way, the individual beam widths increase with the steering angle. A beam steered at 60° is twice the width of a beam steered at 0°.

With a swath of 120°, the maximum steering angle is ±25° ((85°-(swath/2)).

With a swath of 60°, the maximum steering angle is ±55° (85°-(swath/2)).

Figure 9: Horizontal Swath Steering

#### Minimum Grazing Angle

The minimum grazing angle is a limitation for the non-equiangular beam modes of the sonar. This limitation is especially noticeable for a sonar with a tilted head (see section 4.3.4 below).

With a tilted head for Equi-Angle beam mode, there is no minimum grazing angle and the outer beams can look ‘upwards’. With a tilted head for the other beam modes the sonar is limited and cannot look ‘upwards’.

For the beam mode Intermediate, the minimum grazing angle is 5°, and for Equi-Distant and FlexMode, the angle is 15°. For these three beam modes, the outer beam cannot go higher. Figure 10 provides a graphic illustration of the minimum grazing angle.

Figure 10: Outer Beams for the Different Beam Modes with a Tilted Sonar Head

NOTE

When you want to look up with a tilted sonar head, you have to select the Equi-Angle beam mode.

### Use of Beam Mode

The following beamforming modes are available: Equi-Distant, Equi-Angle, Intermediate, and FlexMode (optional).

The system allows everything from high-density beams for exceptional image clarity and detail to just a few beams to deliver results faster. By deploying fewer beams in shallow water, the system allows the user to reduce storage requirements by gathering only the data you need. The system can also hold constant seafloor spacing between beams to provide the same high-quality results with minimal data volume.

In each mode (except FlexMode), the surveyor may choose the number of beams to be formed by using a slider. In Equi-Distant mode, an additional option is available which allows the surveyor to set the desired seafloor spacing with the slider. The system automatically detects the depth and sets the spacing as required by the surveyor. Adjustments may take place during survey: there is no requirement to interrupt the survey.

There is always a minimum number of beams that theoretically ensures “100% insonification” of the swath, which roughly speaking means that there are no gaps between beams; in other words, each beam adjoins or slightly overlaps its neighbor. IHO specifications no longer refer to the term “100% insonification”, as it cannot ever be proved conclusively; it is accepted that the situation may arise where “100% insonification” is not achieved.

#### Equi-Distant, Equi-Angle, and Intermediate Modes

- Equi-Distant mode: Up to 150° coverage angle. Typically, the default mode for general purpose surveys where it is critical to maintain a certain number of hits per cell across the maximum possible swath. Provides high-density equidistant soundings across the entire swath.

- Equi-Angle mode: Up to 170° coverage angle. Typically used for tilted heads to inspect upwards, water column data recording, or for minimizing the amount of data being recorded. Provides an equiangular distribution with the fewest beams.

- Intermediate: Up to 170° coverage angle. Designed for general purpose surveys, tilted head modes, or dual-head operations. Provides high-density soundings over the entire swath.

In Equi-Distant mode, the spacing is modified to maintain an equal horizontal distance between soundings on the seafloor assuming a flat bottom; in Equi-Angle mode, the beam center-center angular spacing is constant across the swath. It is recommended that Equi-Distant mode be used for general survey applications where the requirement to achieve a specific number of sounding per cell in the across-track is critical, or where the seafloor is generally flat.

When surveying features such as wrecks and vertical structures such as harbor walls, Intermediate or Equi-Angle mode should be used in preference to Equi-Distant mode. In extreme conditions, using Equi-Distant mode over complex structures can cause magnified artifacts due to the large overlap between beams towards the edge of the sector.

In short, these are the pros and cons for the three modes, keeping in mind that in each case the negative effect is much more dramatic with a wide swath (see also section 5.2.2 Equi-Distant, Equi-Angle, and Intermediate Modes):

| PRO | CON |
| --- | --- |
| Equidistant spacing: Constant spacing even at large angles; Very large number of beams. |
| Equiangular spacing: Low number of beams; Very wide spacing at large angles. |
| Intermediate (compromise): Moderate beam number Moderate spacing at large angles. |

#### FlexMode

Designed for pipeline or umbilical profiling. Provides equidistant beam spacing for the full swath coverage together with a high-density equiangular sector, which may be steered toward the detail area of interest.

In this mode the system is started with a 140° wedge and an alternative beam distribution is used which concentrates the majority of beams into a central sector.

Figure 11: Graphic Illustration of FlexMode

FlexMode provides a combination of a high density of beams on and around the object location and a reduced density of soundings for the rest of the swath. The low-density spacing is defined by the operator (with the slider for Constant seafloor spacing) and the high-density spacing is computed by optimization based on survey parameters.

For instructions on how to adjust this setting, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

### Use of Gates

Applying range and/or depth gates to the bottom detection process aids in noise reduction and correcting for a tilted sonar head.

- No Gates: The bottom detection process searches the entirety of each beam for the bottom.

- Absolute Gates:

Depth: Forces bottom detection to occur within the Min/Max Depth limits set. If the seabed falls outside the limits, bathymetry data will fail and cannot be recovered.

Range: Forces bottom detection to occur within the Min/Max Range limits set. If the seabed falls outside the limits, bathymetry data will fail and cannot be recovered.

- Adaptive Gates: Forces bottom detection to occur around the seabed profile within an adaptive gating window typically set to 15-20% (of depth). This value should be increased over rougher terrain or over a rapidly changing sea floor where bathymetry is extreme in the vertical values.

#### Depth Gates

Use of a motion sensor is required to rotate the depth gates with the rolling motion of the vessel to ensure soundings are not clipped as the seabed rotates with respect to the sonar.

The depth gates, with no input from a motion sensor, will orient themselves parallel to the horizontal plane of the sonar head, as shown by the Min and Max Depth lines in Figure 12. In this series of illustrations, the bottom detect dots are shown as a horizontal line (Line A or B) with dashed lines representing depth gates above and below the bottom and no vessel roll present. Visualize the bottom as being in the ‘layer’ formed by the depth gates.

Figure 12: Depth Gates, With No Roll

Figure 13 shows the same bottom image with a small roll value to port. Notice that the ends of the swath have been cut off due to the action of the static depth gates (no motion sensor applied). In this figure, and all remaining figures in this discussion, the bottom is flat.

Figure 13: Depth Gates, With Roll

When controlled by a motion sensor, the depth gates will be parallel to the local horizontal plane of the earth even as the vessel rolls. Figure 14 shows the same bottom image with dynamic depth gates applied.

Figure 14: Depth Gates, With Roll and Motion Sensor Input

For instructions on how to adjust this setting, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

#### Depth Gate Tilt

Depth Gate Tilt is a feature that allows the operator to manually specify a tilt value for absolute gates. Tilted depth gates are beneficial when it is desirable to concentrate bottom detect operation efforts, e.g. on a sloping seafloor (see Figure 15).

For instructions on how to adjust this setting, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

| Manual gates with 0° tilt on a sloping seafloor | Manual gates with -15° tilt on a sloping seafloor |
| --- | --- |

Figure 15: Depth Gate Tilt

### Maximizing Quality of Backscatter for Snippets Data

A few survey practices can be followed to maximize the quality of the backscatter signal. For instructions on adjusting settings, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

Constant frequency: Since the acoustic response from the sediment is frequency dependent, it is recommended to maintain the frequency constant over the survey area of interest.

CW mode: For similar reasons, and to avoid the extra processing steps relating to FM signals, it is recommended to perform the survey using CW rather than FM. The emitted CW pulses durations should not be shorter than 100μs and source levels not below 210dB.

- Note that the information about the FM signal characteristics, including the equivalent CW pulse width necessary for footprint computation, are fully documented in record 7002 which is part of the default record set.

Intensity coloring of bottom detection points: Extra attention should be paid to the risk of losing strong or weak signals during digitalization. It is recommended to use the intensity coloring of the bottom detection points in the SUI.

Orange or red dots: Saturated detection, too high intensity for correct digitalization, which results in clipping of values.

Blue dots: Weak detection, too low intensity for correct digitalization, which results in losing signals.

Constant settings: Lastly, keeping the sonar settings relatively constant during the survey is a good practice to make sure that changes in the backscatter levels are only due to the environment.

## Enhancing System Efficiency

### Multi-Detect Operation

Use of the Multi-detect feature is of great value in various applications in terms of both generated result and safety:

Harbor surveys: safer detection

Offshore and renewables: robust cable catenary determination

Marine research: fish school detection

Dredging: suspended sediment detection

Figure 16: Submerged Oil
Detected by Multi-Detect

Energy industry: submerged oil detection

Control the sensitivity of the Multi-detect feature with two controls:

Amplitude sensitivity: Increase to cause more objects to be detected.

Max detections: Up to 5 detections generated per beam, starting with the strongest detection to the weakest detection.

NOTE

When the setting is 1, it returns the strongest detection, not the closest detection. If the surveyor wants to find weaker detections in the water column (e.g. the mast on a sunken ship), the detection setting should be >1.

When operating in harbor environments or over wrecks, where determination of least depth (or distance to object) is critical, these sonar controls are designed to increase safety of survey missions. The controls allow the surveyor to adjust detection sensitivity in order to detect weak reflecting objects in the water column closest to the sonar head assembly in addition to the main seafloor detection.

NOTE

The standard bottom detection is always generated, such that multiple detections are provided in addition to, rather than instead of, the standard detection.

For instructions on how to adjust this setting, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

### Tracker Autopilot

With “Tracker” enabled, all primary sonar control parameters such as range, power, gain, pulse length, and swath width are set by the system based on the quality of data being collected.

Furthermore, Tracker will adjust the TVG settings of Absorption and Spreading to optimize the quality of the outer beams, and open or close the swath angle accordingly thereby ensuring that maximum swath width is delivered with the highest quality result.

As there is a slight delay while tracker is being activated or deactivated, enabling/disabling the feature should not be done in the middle of the survey line.

Once Tracker has been selected, no further operator input is required; however, user controls makes it possible for the surveyor to limit the degree of automation:

Constant Swath for the Tracker to maintain. This allows constant line spacing to be planned independently of depth.

Disable Tracker control of certain settings.

For instructions, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

### Roll Stabilization

Roll stabilization is an important feature directly related to system efficiency that allows the swath width on the seafloor to be maximized, thus increasing line spacing and reducing vessel time.

| Under zero roll conditions, the swath is vertical and the center of the swath is directly below the vessel. |  |
| --- | --- |

Figure 17: Swath under Zero Roll Conditions

When the vessel rolls the swath is rotated and the projected swath on the seafloor is laterally displaced.

Figure 18: Laterally Displaced Swaths under Roll Conditions

| This has the effect of reducing the usable swath width due to the distortions on the edge. |  |
| --- | --- |

| With the input from a motion sensor (and roll stabilization enabled) to steer each receive beam dynamically for every sample, the swath stays vertical irrespective of vessel motion thereby maximizing swath. |  |
| --- | --- |

Figure 19: Swath under Roll Stabilization

It is recommended that the motion sensor data update rate is at least 50Hz and preferably 100Hz. Baud rate and protocol details can be set in Motion Stabilization, but it should be ensured that the baud rate is sufficient to allow the desired update rate.

NOTE

A warning will appear in the SUI, if the motion sensor update rate is lower than 20Hz.

For instructions on motion sensor configuration, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

For installation instructions, see appendix B.5.7 Motion Input.

### Head Tilt

In certain applications, it is desirable to physically rotate the sonar head, for instance when surveying harbor walls where the desire is to survey as close to the waterline as possible.

NOTE

EA beams are recommended when using head tilt, as ED beams are still formed assuming a horizontal seafloor.

For instructions on head tilt configuration, see appendix B.7.7 Hardware, or refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

| Horizontal seafloor and head set at 0° tilt | Horizontal seafloor with head physically set to -15° and head tilt set to 0° | Horizontal seafloor with head physically set to -15° and head tilt set to -15° |
| --- | --- | --- |

Figure 20: Head Tilt Functionality

#### Harbor Inspection (Quaysides, Breakwaters)

In this application a single head is typically tilted up to 30° (max. setting is +/-45°) from normal to register soundings as close to the waterline as possible.

Figure 21: Head Tilt for Harbor Inspection

### Vertical Detection Mode (VDM)

Multipath of the acoustical signal around vertical structures can pose a challenge to the detection capabilities of a multibeam echosounder. This is particularly true for quay walls made from corrugated sheet piling. Vertical Detection Mode (VDM) offers an improved detection under these circumstances.

For instructions on how to use this setting, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

## Compensation for Environmental Conditions

As sound propagates through the underwater environment, the energy is absorbed by the liquid caused by molecular interactions that convert acoustic energy into heat. A propagating sound field also spreads geometrically, which means that the sound pressure decreases with increasing distance from the sonar. Other mechanisms such as bubbles, plankton, tidal currents, currents from rivers etc. may influence sound propagation, but these mechanisms are difficult to measure and model, as they vary with time and only appear in a limited number of areas. However, as absorption and geometrical spreading exist everywhere in the underwater environment, they are considered the most important loss mechanisms.

In the SeaBat system it is possible to compensate for loss caused by absorption and geometrical spreading. In addition, the SeaBat system allows for changes in the speed of sound, a parameter that is important in the beamforming process.

For instructions on how to adjust this setting, refer to the Sonar UI User Manual (see Appendix I Reference Documentation).

### Loss of Sound by Absorption

As sound moves through the water, some of the sound energy is converted into to heat by absorption. The absorption loss depends on the acoustic frequency, the salinity and temperature of the water as well as the depth. Figure 22 shows the absorption loss for three different salinities and three different temperatures at an average depth of 20 meters.

If the exact absorption loss value is not known, these are the recommended approximate values:

- 400kHz operation:

  - 110dB/km for salt water

  - 70dB/km for fresh water

- 800kHz operation:

  - 250dB/km for salt water

  - 200dB/km for fresh water

The absorption loss value is part of the formula that the SeaBat system uses to calculate the amount of gain to apply to the sonar return signal (see section 4.2.4 above).

Ideally, the purpose of the TVG is to ensure the receive gain increases with a rate that compensates for the continuously decaying intensity of the received sound field - a decay which is caused by geometrical spreading and sound absorption. Meanwhile, the TVG does not need to perform an exact compensation. In practice, the purpose of the TVG is only to ensure that the received signals continuously are within the dynamic range of the receiver. Therefore, the applied settings for the absorption and spreading only need to be approximate.

For instructions on how to adjust this setting, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

Figure 22: Acoustic Absorption Reference

### Loss of Sound by Spreading

Spreading of sound through water may be the greatest contributor to the loss of sound. As sound moves away from its source in all directions, the acoustic pressures will decrease with range.

In deep water, where the acoustic propagation is not bounded by the sea surface or the seafloor, the spreading loss is spherical. In this case the pressure decrease is proportional to 1 divided by the range R from the sonar to a given object or surface. Thus, the acoustic pressure decays as 1/R. As the sound is scattered back to the sonar the backscattered field also decreases as 1/R. Consequently, for a two-travel path the acoustic pressure in a free field decays as 1/R2.

In shallow waters, on the other hand, the sound propagation is bounded by the sea surface and the seafloor, and the spreading loss is cylindrical. In cylindrical spreading the sound pressure only decreases as 1/R½, and the two-way travel path decays as 1/R.

In practical situations the transmission loss will lie between 20 and 40dB; in other words, the spreading loss will be neither perfectly spherical nor perfectly cylindrical. Finally, if the sonar operates very close to the seabed or very close to objects in general, near-field effects may yield a transmission loss less than 20dB.

The spreading loss value is part of the formula that the SeaBat system uses to calculate the amount of gain to apply to the sonar return signal (see section 4.2.4 above). If the exact spreading loss value is not known, an initial value of 30dB is recommended.

For instructions on how to adjust this setting, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

### Sound Velocity

For instructions on how to adjust this setting, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

For installation instructions, see appendix B.5.4 Sound Velocity Input.

#### Sound Velocity and Performance

Online sound velocity readings are crucial in the beamforming process to ensure that the beams are steered in the correct directions. Ensure that a real-time sound velocity probe (SVP) is set up to deliver the sound velocity at the sonar head.

Sound velocity varies with pressure, temperature, and salinity. An increase in these three levels increases the sound velocity. The sonar data presented in the SeaBat system is dependent on the sound velocity provided.

CAUTION

An online sound velocity probe must be used for providing up-to-date velocities to ensure optimum performance. A manual velocity may be entered instead but care should be taken to ensure the value entered is representative of the real velocity.

If the local sound velocity is not known, Teledyne RESON recommends a value of 1480 to 1500m/s for open sea areas and 1425m/s for fresh water.

Figure 23 provides a general reference for typical surface sound velocity at different temperatures and salinity.

Figure 23: Typical Sound Velocity Values

#### Recommended Use of SVP Filter

The speed of sound impacts the beam angles. If a sonar ping is processed with a sound velocity value that is not reflecting the speed of sound right in front of the receiver, the beam angles will be wrong and hence the targets or bottom detects will have wrong coordinates.

Some software packages allows post-compensation of wrong sound velocity, but it is typically possible to avoid this extra complication of processing by setting a proper filter on the SVP data during real-time recording. For instructions on how to adjust this setting, refer to the Sonar UI User Manual (see Appendix I Reference Documentation).

No Filtering: Although this setting does not filter data in the classical sense, it still ignores readings outside the range 1350m/s to 1800m/s.

- Use this filter if there is a requirement for immediate response to sound velocity changes. When using this filter any false reading will cause distinct false output. This filter may be preferred for certain scientific sonar work.

SVP Light: Removes occasional false readings.

- Use this filter if there is a need for fast filter response, for instance in regions with mixing waters.

- If the SVP update rate is less than 10Hz it may be preferred to use this filter as the standard filter.

SVP Normal: Removes single and double false readings and outliers, and smoothens the SVP value across 10 readings.

- The filter is recommended for most sonar work on systems using an SVP (SVP 70) with no internal filtering and an update rate of 10-20Hz.

NOTE

The SVP Normal filter is typically the best filter to use.

SVP 70 Wizard: This filter only works with the SVP 70 if the internal SVP 70 filter is set to mode 1 (low pass filtering).

- Sensors with a low pass filter cause a special artifact in the output in case of outliers and invalid readings. The SVP 70 Wizard filter removes these artifacts and makes the system work well despite of the low pass filter.

- It is recommended to change the SVP 70 filter mode to 0 when convenient and reset the SVP Filter to SVP Normal.

For access to filtered and unfiltered values, please see the DFD messages R7510 and R7610. Refer to the Data Format Definition document for information (see Appendix I Reference Documentation).

## Limitations

Physical limitations in the functionality of the system have to be taken into consideration when operating sonar systems, including SeaBat systems. Some of the limitations are absolutes; others can to some degree be overcome.

The following subsections provide examples of different kinds of limitations.

### Cavitation

Some reduction in the transmitted source level is expected to be caused by cavitation. For means of compensating for this, see subsections in 4.2 above.

#### Propeller Cavitation

Propeller cavitation can be a significant concern for two reasons. First, the frequency of propeller cavitation noise is in the operating frequency band of most sonar equipment. Second, the source levels of the propellers installed are expected to be high unless special quieting techniques have been implemented.

Propeller energy typically propagates to the forward area sonar systems via a direct or hull-grazing path. The other arrival path of propeller energy is by the bottom bounce path. This path will be more significant in shallow water depths (less than 500 meters). The potential impact can be further quantified for particular bottom depths of proposed operational test sites.

#### Appendage Cavitation

Appendage cavitation can result from rough or unfaired edges on the hull and can even be present at lower speeds. If the source is located near the sonar installation area, it can be a significant noise deficiency. It is difficult to predict the occurrence of appendage cavitation without under-hull drawings, pictures, or visual inspection of the ship when in dry-dock.

### Machinery Noise

Machinery noise should not be a major factor for sonars that operate at frequencies above 5kHz. Mechanical machinery noise is usually prevalent at 2kHz and below, though it is occasionally detected at higher frequencies when fluid flow through pipes or cavitating pumps are involved.

Typically, the most important consideration for machinery noise is the proximity of the acoustic sensor to the machinery source. There has been one observed case where diesel engine piston frequency harmonics have been the controlling noise source at 12kHz, but in this situation, the engine room and the sonar compartment shared a common bulkhead. Normally, acoustic devices are sufficiently distanced from major machinery noise sources that machinery noise does not interfere.

A concern remains that there may be some auxiliary machinery (or other unusual equipment) located near the sonar that could produce noise interference. An inspection of the installation site and ship drawings to verify the position and proximity of all shipboard equipment should be conducted to determine if there is a potential for machinery interference prior to conducting at-sea testing. For means of compensating for this, see subsections in 4.2 above.

### Interference from Other Sonar Systems

Interference from other sonar systems can be seen as radial lines, typically moving away the minimum to maximum range scale as these pings are not correlated with the ping repetition rate of the SeaBat system. The most typical source of interference of this type is navigational sonars (often 50kHz systems) and Doppler velocity logs.

Synchronization of acoustic systems is one means to mitigate this effect. The most effective approach is to remove the source of interference entirely, by either moving the equipment away from the SeaBat or turning it off.

### Speed

Hydrodynamic flow noise can be a concern when operating at higher speeds. Externally mounted acoustic sensors should have fairings that are as flush (or near-flush) mounted to the ship hull as is practically possible. This way the hydrodynamic flow-induced vibration noise should be minimal. For means of compensating for this, see subsections in 4.2 above.

NOTE

A poorly designed sonar fairing can degrade the system performance by a factor of 4 for speeds above 8-10 knots. Typically, this noise source is the most common problem on all sonar installations.

### Air Bubbles

Bubble sweep-down can be another potential noise problem. The significance of this problem is associated with the noise generated as the bubbles cascade along the hull after being ingested in the bow wave of the ship and, more importantly, with the baffling produced by the entrained air layer between the face of the sonar sensor and the water environment.

If air bubbles are present between the sonar and the water, the system will not function properly. This phenomenon is transient in nature occurring only for a few seconds per wave period. It is typically most prevalent on shallow draft ships and increases in intensity at higher sea states and ship speeds.

Modeling of this problem only indicates that air bubble noise may be present by determining the transmission vehicle of potential bubbles. Actual at-sea measurements are usually required to fully determine the presence and severity of bubble sweep-down. At-sea diver viewing has been very successful in the past to capture the bubble sweep-down characteristics of a particular ship hull on film.

Air bubbles mask some or all of the view of the sonar where the affected area will completely blank out. Common causes for bubble masking include:

The vessel props when thrusting into reverse and pours a stream of bubbles over the sonar.

The sonar has not been deployed deeply enough and there is bubble wash from the sea surface.

Bubble ingestion as the vessel pitches into the sea.

### Environment

A high-reverb environment where echoes of the previous ping are contaminating the current ping may prove to be difficult. For means of compensating for this, see section 4.2.3 above.

# System Performance and Data Types

NOTE

There are several factors that can have a significant effect on acoustic performance. For this reason, product performance estimates are indicative of standardized conditions for temperature, salinity, bottom type or target strength, sound velocity, installation type, and assume no affect from adverse environmental conditions such as wind, precipitation, sea state, ambient or traffic noise, refraction, temporal changes or other surface effects. (Read also section 4 Active Sonar Usage.) To determine actual in-situ performance predictions requires a site or vessel survey and environmental measurements with performance modeling based on the resulting measurements. Costs for this survey may be quoted as part of a contract or independently.

The SeaBat T51-R sonar systems produce four types of data, bathymetry, side-scan, snippets backscatter and water column backscatter, suitable for the generation of high-resolution hydrographic charts that exceed international standards.

A single transmission from the projector unit illuminates the entire swath on the seafloor. The seabed return signal is received by the receiver unit, digitized, and transmitted to the sonar processor, where it is formatted for display and data output. Bottom detection bathymetry soundings, side-scan, and snippets are displayed locally or exported via gigabit Ethernet using TCP I/P protocol to an external data collection system.

A Windows®-based user interface controls the system, allowing the operator to select system configuration, data output, and storage options. This also provides built-in test environment (BITE) displays to assist the operator.

For available configurations, see section 1.6 Available SeaBat T51-R Configurations.

Table 5: SeaBat T51-R System Performance

| Parameters | 350-430kHz | 700-800kHz |
| --- | --- | --- |
| Frequency range | 350-430kHz (CW) 400kHz (FM) | 700-800kHz (CW) 775kHz (FM) |
| Across-track receive beam widths* (nominal values) | 0.5° | 0.27° |
| Along-track transmit beam widths (nominal values) | 1° | 0.5° |
| Number of beams | Max. 1024, min. 10 |  |
| Max. swath coverage | 150° ED/170° EA |  |
| Depth, typical | CW: 0.5-200m FM: 0.5-225m | CW: 0.5-85m FM: 0.5-85m |
| Depth, max. | CW: 300m FM: 350m | CW: 125m FM: 150m |
| Ping rate (range dependent) | Up to 50 pings/s |  |
| Pulse lengths CW | 30-300µs |  |
| Pulse lengths FM | 300µs-10ms |  |
| Depth resolution | 6mm |  |
| Depth rating | 50m |  |
| Temperature (dry end) | Operating: -5° to +45° Storage: -30ºC to +70ºC |  |
| Temperature (wet end) | Operating: -2ºC to +36ºC Storage: -30ºC to +70ºC |  |

| * Receive center beam widths: | 0.54° (at 400) / 0.27° (at 800kHz) |
| --- | --- |

## System Components

### Rack-mounted Sonar Processor (RSP)

| The sonar processor is a high-performance industrial unit that manages data flow and signal processing using a state-of-the-art processing architecture. This highly integrated design offers reliability, maintainability, and high performance. For physical specifications, see appendix B.4.1 Rack-mounted Sonar Processor (RSP). | Figure 24: Rack-mounted Sonar Processor |
| --- | --- |

The sonar processor offers a flexible platform for a number of functions:

Receives digitized sonar data from the receiver unit.

Receives operational settings directly through the user interface or remotely from an external system.

Provides beamforming and initial processing of acoustical data.

Controls, formats, and generates data to external systems. This can include making beamformed data and preprocessed image data available to external systems over a fast Ethernet connection.

Performs built-in test routines and alerts the operator to any abnormal conditions.

Provides an interface and single point time stamping for serial sensor data:

1PPS and time recovery message (GNSS).

Position

Heading

Motion

Sound velocity

Displays the sonar image and the detected seabed as well as the sonar control user interface.

Generates the various data telegrams to be recorded by data acquisition systems.

Supports optional beam data storage on a large, external RAID array.

### EM7222 Receiver Unit

The EM7222 receiver is a dual-array unit that operates at multiple frequencies, 350-430kHz and 700-800kHz.

For physical specifications, see appendix B.7.1 EM7222 Receiver Unit (PN 1010473-DK).

| Front view | Rear view |
| --- | --- |

Figure 25: EM7222 Receiver Unit

### TC2186 Projector Unit

| Front view The TC2186 projector unit operates at multiple frequencies between 400kHz and 800kHz. The entire swath is illuminated in one transmission, with the maximum ping rate of 50Hz, providing 100% bottom coverage even at high vessel speed. For physical specifications, see appendix B.7.2 TC2186 Projector Unit. | Rear view |
| --- | --- |

Figure 26: TC2186 Projector Unit

## Selectable Beam Density

Selectable beam density is a feature that allows the surveyor to decide how many beams to form ranging from the maximum capability of the system down to just 10. This allows the surveyor to be able to capture the highest detail using the highest beam density available, but alternatively maximize efficiency by using fewer beams in very shallow water where data volumes traditionally become large quickly due to high beam numbers and high ping rates.

A special mode is the ability to define the desired equidistant spacing between beams on the seafloor and keep it constant. The system will “hold” this automatically, as depth varies, within the constraints defined by the minimum and maximum number of beams.

Selectable beam density adds to the flexibility of the system with several benefits:

Optimized data volumes tailored to survey purpose/specifications

Save of up to 75% of data volume in very shallow water

Faster processing

Reduced time and costs from survey to deliverable

### Beam Modes

Beam mode selection is made on the fly via the sonar user interface. For instructions, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation). Available selections depend on options purchased (see also Table 1).

Table 6: Beam Modes with No. of Beams and Swath Coverage

|  | No. of beams |  |  | Swath |
| --- | --- | --- | --- | --- |
|  | T51 | T51 |  | T51 |
|  | Min. | Std. | Max. | Max. swath |
| Equi-distant (with constant no. of beams) | 10 | 1024 |  | 150 |
| Equi-distant (with constant beam spacing) | 10 | 1024 |  | 150 |
| Intermediate | 10 | 1024 |  | 170 |
| Equi-angle | 10 | 1024 |  | 170 |
| FlexMode | 1024 | 1024 |  | 150 |

For examples on how to use this feature, see section 4.2.9 Use of Beam Mode.

### Equi-Distant, Equi-Angle, and Intermediate Modes

As the beam steering angle increases from nadir, the number of beams per degree of swath stays constant for equiangular spacing while it increases for equidistant spacing. Figure 27 presents the number of beams as a function of steering angle for a 1° nadir beam spacing.

Figure 27: Beam Density

| Figure 28 shows that the horizontal spacing of the Intermediate mode increases with beam angle, unlike Equi-distant mode where the spacing stays constant. However, the Intermediate mode increases the spacing much slower than the Equi-angle mode so the outer beam density is still quite good. | Figure 28: Beam Spacing vs. Beam Angle |
| --- | --- |

Assuming a flat bottom and using the beam footprint generated by the nadir beam, it may be seen that the center-to-center spacing for each beam successively reduces as the departure from nadir increases until at ±65° the beam density is approximately 2.5 times greater than at nadir.

Assuming one sounding is generated from the intersection of the beam center with the seafloor, with equidistant spacing, the appropriate number of beams is formed to maintain this nadir sounding spacing across the entire swath. See also Figure 29.

Figure 29: Beam Spacing

In order to achieve this uniform sounding spacing, it is necessary to decrease the beam center-to-center spacing, which requires a significant increase in the number of beams formed.

The beamformer is able to generate a receive beam centered on any angle within the available swath and with equidistant spacing a set of coefficients is selected that form a set of receive beams that are centered on the midpoints of the footprints.

Figure 30 shows the difference in distribution between equiangular and equidistant spacing.

Figure 30: Across-Track Sounding Location on a Flat Bottom

## Data Types

### Bathymetry Data

The SeaBat T51-R system generates between 10 and 1024 soundings per ping.

The system broadcasts from the sonar processor to the data acquisition software, where they are corrected for mechanical offsets, motion, heading, refraction, tide or depth, and position. To perform these corrections, the appropriate sensors must be interfaced to the data acquisition software. For instructions on sensor configuration, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

Depending on the beam angle, bottom type, and signal characteristics, the bottom detection method employed may be magnitude center-of-energy, phase slope zero-crossing, or a blend of the two methods. The bottom detection method can vary from ping to ping and beam to beam.

### Side-Scan Data

Side-scan swath is measured and displayed without degrading any of the traditional survey capabilities of the SeaBat.

NOTE

Side-scan cannot be used to accurately measure true depths, but it can provide a more detailed picture of the seafloor. This image can be used together with bathymetry to identify features and to help ensure that the survey does not miss any small but significant targets.

#### Side-Scan Process

The side-scan process averages adjacent pairs of beams to eliminate noise and then carries out a peak detection over beams from the left and right sides of the swath. The peak detection is done at the sample rate of the sonar.

The bathymetry and side-scan data are independent in the beamforming process and in how the output data is generated.

#### Side-Scan Data Output

The array of intensity values is a series of amplitudes, one for each sample interface for each side-scan beam.

The number of intensity values reported in a side-scan packet is a function of range. The number of values per packet can be calculated using the following formula:

Range ½SoundVelocity *SampleRate=SamplesPerBeam

Side-scan is one of the many records available for subscription from the 7k Center software. Refer to the 7k Data Format Definition for further information regarding side-scan data output (see also Appendix I Reference Documentation).

### Snippets Backscatter

#### Snippets Backscatter Data

Snippets backscatter data is available from the SeaBat T51-R. Each snippet consists of a windowed selection of amplitudes taken from the bottom detection point of each beam.

A snippets data sample contains backscatter data from the ‘footprint’ on the seabed illuminated by a single sonar ping. The number of snippets in a swath is a function of the number of sonar beams. The length of each snippet depends on the operating mode, beam number, and depth. All of these settings may be selected by the operator.

Figure 31: Graphic Illustration of Snippets Data

Each snippets data packet contains pertinent information such as time stamp, sequential ping number, sample rate, sound velocity, and operator settings such as power, gain, absorption, and range scale.

Refer to the 7k Data Format Definition for further information regarding snippet data output (see also Appendix I Reference Documentation).

## Real-Time Pipe Detection and Tracking

Real-time pipe detection and tracking is a SeaBat feature that automatically detects pipes within the multibeam swath, displays them to the user in real time during survey, and exports the position of the pipe top to the data acquisition software. The feature builds on the robustness of FlexMode for generating a high-density profile of the pipe.

The user enters the diameter of the pipe to be detected in the SeaBat UI and sets parameters to tune the pipe detection algorithm. When enabled, the real-time pipe detection algorithm computes the pipe position in the SeaBat’s coordinate system and shows a circle representing the position in the swath to the operator. The detected pipe position is also sent to the data acquisition software.

Figure 32: SeaBat Real-Time Pipe Detection and Tracking

Real-time pipe detection in the SeaBat system provides a unique advantage that improves the end product. Starting from the initial detected pipe position, the diameter of the pipe is used to assist the bottom detection near the pipe in subsequent pings to enhance bottom detection robustness, and thus the pipeline profile. This unique feedback loop from pipe position to the bottom detection algorithm results in much cleaner pipeline profiles and, therefore, much improved detected pipe positions. This is of particular advantage in instances of plastic coated (acoustically difficult) pipes where backscatter from the sides can be weak to the point that bottom detection performance is compromised, even though the sonar acoustically detects the pipe.

Figure 33: SeaBat FlexMode

Benefits of the automated pipe detection and tracking in FlexMode include:

Reduced manual pipeline processing effort

Improved end product

Reduced risk to subsea assets

Reduced risk of operator error

## Multi-Detect

| Multi-detect is a feature that provides multiple detections within each beam, including the full water column. It allows the surveyor to capture enhanced detail from a single survey line over a complex feature, or ensure that full detail is captured of any object in the water column, even in noisy environments. Up to five detections may occur anywhere within each beam. (See also section 4.3.1 Multi-Detect Operation.) NOTE. Multi-detect is not equivalent to multiple soundings across a beam footprint on the seabed.. With the multi-detect feature you achieve: Enhanced detail over complex, small, and fine objects. Improved and faster decision making. Easier and faster water column processing. More project opportunities. In the scenario in Figure 36, without multi-detect the sonar would detect either the anchor chain (blue dots) or the seabed (red dots) or flicker between the two features. With multi-detect, the sonar can detect both the anchor chain and the seabed. | Figure 34: Standard Detection Figure 35: Multi-Detect with enhanced detail on Quay Wall Figure 36: Multi-Detect |
| --- | --- |

## Vertical Detection Mode

Vertical Detection Mode (VDM) is a feature that provides optimized detection performance on vertical structures, such as quay walls.

When enabled, the detection algorithm will automatically identify vertical structures and provide the best possible detection data quality of the structure. It will also reveal features, or any anomaly present on its surface. VDM is available for all SeaBat T-Series systems and works irrespective of beam mode, single or dual head, head tilt or sonar settings.

Figure 37: Vertical Detection of Quay Wall

## X-Range Performance

X-Range is a feature that provides extended range performance at the full range of frequencies and significantly improves system immunity towards external noise. The basis of the X-Range feature is a frequency-modulated (FM) transmission combined with advanced signal processing techniques to extract the maximum possible performance from the system.

## Quality

The quality of each bathymetric bottom detection point is evaluated by the 7k Center. For instructions on how to customize some of the visual properties of the sonar image, refer to the Sonar UI User Manual (see Appendix I Reference Documentation).

### Brightness Test

The brightness test measures the signal-to-noise ratio (SNR) of the signal strength (magnitude) near the bottom detection point (as determined by the bottom detection algorithm) versus the background noise in the rest of the beam.

The signal strength is the mean of the square of the sample intensity around the bottom detection point. The number of samples used is marked ‘A’ in Figure 38 and is a small percentage of the total number of samples. The noise is based on a mean of the square of the samples in the second half of the “tail” of the detection window (marked ‘B’ in Figure 38) and the square of the samples in the “head” of the water column (before the first return – marked ‘C’ in Figure 38).

Beams, whose SNR is greater than a predetermined threshold, will be flagged as passing the brightness test.

For instructions on how to display the test results, refer to the Sonar UI User Manual (see Appendix I Reference Documentation).

Figure 38: Brightness Test

### Colinearity Test

Figure 39: Colinearity Flags

In general terms, this test will pass if a bottom detection point sample number for a given beam falls on or near a line drawn through its neighboring beams' bottom detection points. This is achieved using a moving polynomial fit of the points on either side of each point. If a given point is within the window of the line fit, it is marked colinear.

This process is done in stages that begin with the central beams and moving outward on each side of the wedge in turn.

For instructions on how to display the test results, refer to the Sonar UI User Manual (see Appendix I Reference Documentation).

### Quality Filter

A quality filter makes it possible to not display or export bottom detection points that do not pass the quality criteria. Enabling the quality filter has the following effects:

The corrected bathymetric data record will only contain the detection points that pass the quality filtering. A flag will be set in the record to indicate that the quality filter was applied.

The dedicated snippets record will only contain the detection points that pass the quality filtering. A flag will be set in the records to indicate that the quality filter was applied.

Only detections points that pass the quality filtering will be shown in the various displays.

For details on settings for the quality filter, refer to the Sonar UI User Manual (see Appendix I Reference Documentation).

## Uncertainty Estimation

The calculation of Total Propagated Uncertainty (TPU) is used to aid the surveyor determine the sounding quality of the bottom detection point. The calculation is done by statistical methods, typically in the data acquisition software. One input into the TPU calculation is the sounding measurement error as computed in the sonar system. Other error sources include position, draft, squat, load, tide geodetic calculations, lever arms calculations, timing offsets, speed, heading, attitude, and sound velocity.

In the sonar processor, the error associated with bottom detection is computed separately for amplitude, phase, and blended detection methods. This value is exported as a field in both raw and corrected bathymetry records.

The uncertainty value is normalized to the detection point by dividing the absolute range error (number of samples) by the detection point (fractional sample number). For example:

Absolute range error: 0.8 samples

Fractional sample number of detection point: 120.2

Uncertainty = 0.8/120.2 = 0.006655574

When the Uncertainty option is active, a color legend is shown on the wedge display.

For instructions on how to display the uncertainty, refer to the Sonar UI User Manual (see Appendix I Reference Documentation).

## Hydrodynamic Drag

Hydrodynamic drag has been determined for sonar head assemblies in brackets mounted on a short pole traveling through the water at a speed of 10 knots. Although the drag simulations have been done for a SeaBat T50 setup, the results are equally relevant for the T51 sonar head assembly, provided the installation is done in a similar way.

### Drag Simulation Setup

Sonar head assemblies:

T50 single head with front fairing (Figure 41 and Figure 42)
or without front fairing (Figure 43 and Figure 44).

Pole mount: 350mm long, 100mm diameter pole attached to the mounting flange of the brackets. The pole mount itself also causes drag.

Speed: 10 knots.

Finite Element simulation tool: ANSYS was used to determine the drag and evaluate water flow around the sonar heads and brackets.

NOTE

Cable routing and accessory parts like a SVP or a motion sensor added to the bracket will influence drag.

### Drag Simulation Conclusions

The front fairings help reduce drag and keep a steady and clean water flow around the receiver and projector arrays. Need of fairing depends on survey and transit speed.

The hydrodynamic drag on the system is a combination of skin friction and pressure drag. Both are within reasonable speed proportional to square of speed. With this it is possible to estimate drag figures for speeds that differ from 10 knots.

The drag figures are relevant for a pole mount over the side, a sonar head installation through a moon pool, or a sonar head mounted on an adapter bracket under the hull of the survey vessel.

In general, Teledyne RESON recommends the use of the optional fairing to maintain a steady water flow around the array and reduce load on the pole/bracket.

### Drag Simulation Results

#### T50 Single Head Drag

Figure 41 illustrates the flow at 10 knots around the T50 single head with front fairing. The drag on the pole, sonar head, and bracket is in the region of 650N. The same situation without front fairing is shown in Figure 43.

Figure 40: T50 Sonar Head Assembly in Bracket with Optional Front Fairing

Figure 41: Flow at 10 knots around Single T50 with Fairing

Figure 42 is a cross section of the symmetry plane of the T50 single head. The front fairing and bracket assure a steady water flow around the receiver and projector. The same situation without front fairing is shown in Figure 44.

Figure 42: Flow around Single T50 with Fairing, Cross Section of Symmetry Plane

Figure 43 illustrates the flow at 10 knots around the T50 single head without front fairing. The drag on the pole, sonar head, and bracket is in the region of 1000N. The same situation without front fairing is shown in Figure 41.

Figure 43: Flow at 10 knots around Single T50 without Fairing

Figure 44 is a cross section of the symmetry plane of the T50 single head. Without the front fairing the flow is still reasonable, but not as steady as with the fairing in front of the receiver. The same situation with front fairing is shown in Figure 42.

Figure 44: Flow around Single T50 without Fairing, Cross Section of Symmetry Plane

## Ping Rate Table

Table 7: Ping Rate Table

| Range setting (m) | Theoretical max ping rate (1500m/s)[1] | Ping rate [Hz] |
| --- | --- | --- |
| 5 | 150 | 50 |
| 7 | 107.1 | 50 |
| 10 | 75.0 | 50 |
| 15 | 50.0 | 43.6 |
| 20 | 37.5 | 33.8 |
| 25 | 30.0 | 27.6 |
| 30 | 25.0 | 23.3 |
| 35 | 21.4 | 20.2 |
| 40 | 18.8 | 17.8 |
| 50 | 15.0 | 14.4 |
| 70 | 10.7 | 10.4 |
| 100 | 7.5 | 7.3 |
| 140 | 5.4 | 5.3 |
| 160 | 4.7 | 4.6 |
| 180 | 4.2 | 4.1 |
| 200 | 3.8 | 3.7 |
| 250 | 3.0 | 3.0 |
| 300 | 2.5 | 2.5 |

Troubleshooting

System Error and Warning Messages

NOTE

Most errors and warnings do not result in a system failure. In the case of repeated errors or warnings, contact Teledyne RESON Customer Support (reson-support@teledyne.com).

This chapter contains a global list of error and warning messages generated by the 7k Center software. The messages appear in the Messages pane of the Sonar UI.

For instructions on how to monitor the system, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

General information (no error/warning) is marked with an  in the Messages pane.

Error Messages

NOTE

The error messages listed in this section are accurate up to the document release date. For a complete list of up-to-date error messages, contact Teledyne RESON Customer Support (reson-support@teledyne.com).

Errors are marked with a red  in the Messages pane.

General System Startup Errors

| Message | Description | Action |
| --- | --- | --- |
| Could not create UDP server | Indicates a network resource was unavailable to the software from the operating system. | Check user permissions, network drivers. |
| Could not create TCP server |  |  |
| Could not create PCI server | Software detected a failure to configure the main data acquisition and distribution module (PCI). This can be caused by errors in the configuration files, or corrupted installation. | Reboot the PC, call Teledyne RESON support. |
| Error creating PCI Reader thread |  |  |
| Error creating PCI Copier thread |  |  |
| Create TCP socket failed, code… | Operating system did not allow creating of network resources. | Check Windows installation, user settings, and network resources. |
| Error creating non-sonar data reader thread | BITE module has failed to initialize. | Check Windows installation and resources. |
| Internal 7k Center error | One or more startup conditions have failed. | Check previous alarms or warnings and verify software installation. |
| Internal error, 7kCenter NOT fully operational |  |  |
| Inbound memory initialization failed | PC resources are inadequate. | Check RAM or virtual memory resources |
| MMF Outbound memory creation … |  |  |
| Required file: "7k_0.xml" or "7k.xml" Not found. | A critical startup configuration file was missing. | Reinstall the software. |
| Required file: "7k_Bite.htm" Not found. |  |  |
| SET UP: FAILED TO LOAD BIT FILE "XXX" for Device <x>, Enumerator <x> |  |  |
| Socket initialization failed, error code <x> | Windows has failed to allocate a network resource. | Check network permissions and resources or Reinstall operating system. |

Data Reading Errors

| Message | Description | Action |
| --- | --- | --- |
| DMA Transfer Timeout. <x> of <x> bytes received | This error generally indicates that the uplink from the receiver is interrupted at some point. | Check the BITE tab for receiver BITE values. Restore factory defaults. Reboot the PC. |
| FATAL: DMA Read size limitation reached. | A general error occurred while transferring data to the software. | Reboot the system. |
| Lost sync with DMA read |  |  |

System Resource Errors

| Message | Description | Action |
| --- | --- | --- |
| Error starting <name> thread. | A module has failed to start correctly. | Reboot the system. |

Input Signal (Timing and Triggers) Errors

| Message | Description | Action |
| --- | --- | --- |
| Failed to register external trigger, ERROR = | External Trigger Firmware cannot be initialized or is not supported. | Reboot the system. Contact Teledyne RESON Customer Support if the problem persists. |
| Failed to unregister external trigger, ERROR = |  |  |
| External triggering has failed. Check/update software installation. |  |  |
| FATAL ERROR: PPS Timestamp error Check 1 PPS input | The pulse is not being detected. Time stamping will be done using system time only. 1 millisecond accuracy not possible. | Check input signal cabling. |

Normalization Errors

| Message | Description | Action |
| --- | --- | --- |
| Normalization Failed: Too much noise | Normalization has failed. | Check BITE page and hardware and cable installation to the receiver. Check environmental noise. Contact Teledyne RESON Customer Support. |
| Normalization Failed: Over 10% dead Rx channels. |  |  |
| Normalization Failed: Magnitude Error. |  |  |
| Normalization Failed: Phase Error. |  |  |
| Normalization Error: Normalization might be invalid. Repeat the normalization. |  |  |

Warning Messages

NOTE

The warning messages listed in this section are accurate up to the document release date. For a complete list of up-to-date warning messages, contact Teledyne RESON Customer Support (reson-support@teledyne.com).

Warnings are marked with a yellow  in the Messages pane.

Configuration Warnings

| Message | Description | Action |
| --- | --- | --- |
| 7K Bite File cannot be Read | Software installation is corrupt. | Restore factory defaults. Reinstall software. |
| Cannot Read UDP Configuration file <x> Error |  |  |
| Cannot Open or Create UDP Configuration file |  |  |
| UDP Configuration file not available |  |  |

Data Processing Warnings

| Message | Description | Action |
| --- | --- | --- |
| ***** (xxx) Dropping out-of-order ping #<x> | The data processing chain has been delayed and stale data has been discarded. | Check sonar processor hardware and resources. |
| PING <x> :( Lost FPGA sync...sample | One or more pings have been lost due to an error on the uplink between the receiver and the software. The system is becoming unstable and data transfer is corrupted on the PCI bus. | Reboot. Check sonar processor hardware. |
| DMA Transfer Timeout. <x> of <x> bytes received |  |  |

BITE Data Warnings

| Message | Description | Action |
| --- | --- | --- |
| ERROR : NSD | An error has occurred in reading of BITE data. The data buffer has overflowed. | No action needed. |
| Non Sonar Data FIFO overflow (x packet) |  |  |

Input Signal Warnings

| Message | Description | Action |
| --- | --- | --- |
| Roll Compensation is not available. | The input data for motion is not available or the feature is not enabled. | Check device inputs and cabling. |

System Stability Warnings

| Message | Description | Action |
| --- | --- | --- |
| WARNING: PPS Time-stamp error | The internal clock may not be disciplined properly resulting in incorrect time stamp of data. | Check 1PPS connections and GPS output. If uncorrected sonar data cannot be time stamped to 1 millisecond. |
| WARNING: Reduce Ping Rate or Range - DROPPING ping | The system is overloaded processing data. A ping has been lost because the processing queue is full. | Restart the software. Reboot PC. |

Data Subscription Warnings

| Message | Description | Action |
| --- | --- | --- |
| Subscriber UDP sendto() failed | The software was unable to send data to a connected client. | Check network resources and user permissions. |
| Subscriber send() failed |  |  |
| Subscribers outbound MMF buffer memory overwrite error! | The link to the user interface running on the sonar processor has detected a non-critical error. | Check system resources, shutdown all non-critical software applications. Reboot processor. |
| Subscribers outbound MMF buffer read error! |  |  |

Network Components Warnings

| Message | Description | Action |
| --- | --- | --- |
| TCP Server - Unknown packet type | A connected client has sent an unrecognized command to the software. | Review client software source code or definitions for correct implementation of Teledyne RESON 7k Data Format. |
| UDP Server - Unknown packet type |  |  |
| TCP Server - Warning max connections reached | The network module has become saturated with client connections and will refuse any further connections. | Close some of the connected clients. |
| UDP Server - Warning max connection reached |  |  |

External Trigger Warnings

| Message | Description | Action |
| --- | --- | --- |
| Trigger out length (x) is invalid | An unsupported value was received for trigger setup. | Check settings in user interface. |
| Trigger out offset (x) is invalid |  |  |
| Trigger out offset + length (x) exceed limit of 20ms |  |  |

Error Messages – Master List in Alphabetical Order

| Message | Description | Action | Type |
| --- | --- | --- | --- |
| Could not create PCI server | Software detected a failure to configure the main data acquisition and distribution module (PCI). This can be caused by errors in the configuration files, or corrupted installation. | Reboot the PC, call Teledyne RESON support. | General System Startup |
| Could not create TCP server | Indicates a network resource was unavailable to the software from the operating system. | Check user permissions, network drivers. | General System Startup |
| Could not create UDP server | Indicates a network resource was unavailable to the software from the operating system. | Check user permissions, network drivers. | General System Startup |
| Create TCP socket failed, code… | Operating system did not allow creating of network resources. | Check Windows installation, user settings, and network resources. | General System Startup |
| DMA Transfer Timeout. <x> of <x> bytes received | This error generally indicates that the uplink from the receiver is interrupted at some point. | Check the BITE page for receiver BITE values. Restore factory defaults. Reboot the PC. | Data Reading |
| Error creating non-sonar data reader thread | BITE module has failed to initialize. | Check Windows installation and resources. | General System Startup |
| Error creating PCI Copier thread | Software detected a failure to configure the main data acquisition and distribution module (PCI). This can be caused by errors in the configuration files, or corrupted installation. | Reboot the PC, call Teledyne RESON support. | General System Startup |
| Error creating PCI Reader thread | Software detected a failure to configure the main data acquisition and distribution module (PCI). This can be caused by errors in the configuration files, or corrupted installation. | Reboot the PC, call Teledyne RESON support. | General System Startup |
| Error starting <name> thread | A module has failed to start correctly. | Reboot the system. | System Resource |
| External triggering has failed. Check/update software installation. | External Trigger Firmware cannot be initialized or is not supported | Reboot the system. Contact Teledyne RESON Customer Support if the problem persists. | Input Signal (Timing and Triggers) |
| Failed to register external trigger, ERROR = | External Trigger Firmware cannot be initialized or is not supported. | Reboot the system. Contact Teledyne RESON Customer Support if the problem persists. | Input Signal (Timing and Triggers) |
| Failed to unregister external trigger, ERROR = | External Trigger Firmware cannot be initialized or is not supported. | Reboot the system. Contact Teledyne RESON Customer Support if the problem persists. | Input Signal (Timing and Triggers) |
| FATAL ERROR: PPS Timestamp error Check 1 PPS input | The pulse is not being detected. Time stamping will be done using system time only. 1 millisecond accuracy not possible. | Check input signal cabling. | Input Signal (Timing and Triggers) |
| Inbound memory initialization failed. | PC resources are inadequate. | Check RAM or virtual memory resources | General System Startup |
| Internal 7k Center error | One or more startup conditions have failed. | Check previous alarms or warnings and verify software installation. | General System Startup |
| Internal error, 7kCenter NOT fully operational. | One or more startup conditions have failed. | Check previous alarms or warnings and verify software installation. | General System Startup |
| Lost sync with DMA read | A general error occurred while transferring data to the software. | Reboot the system. | Data Reading |
| MMF Outbound memory creation … | PC resources are inadequate. | Check RAM or virtual memory resources | General System Startup |
| Normalization Error: Normalization might be invalid. Repeat the normalization. | Normalization has failed. | Check BITE page and hardware and cable installation to the receiver. Check environmental noise. Contact Teledyne RESON Customer Support. | Normalization |
| Normalization Failed: Magnitude Error | Normalization has failed. | Check BITE page and hardware and cable installation to the receiver. Check environmental noise. Contact Teledyne RESON Customer Support. | Normalization |
| Normalization Failed: Over 10% dead Rx channels. | Normalization has failed. | Check BITE page and hardware and cable installation to the receiver. Check environmental noise. Contact Teledyne RESON Customer Support. | Normalization |
| Normalization Failed: Phase Error | Normalization has failed. | Check BITE page and hardware and cable installation to the receiver. Check environmental noise. Contact Teledyne RESON Customer Support. | Normalization |
| Normalization Failed: Too much noise | Normalization has failed. | Check BITE page and hardware and cable installation to the receiver. Check environmental noise. Contact Teledyne RESON Customer Support. | Normalization |
| Required file: "7k_0.xml" or "7k.xml" Not found. | A critical startup configuration file was missing or too many were found. | Reinstall the software. | General System Startup |
| Required file: "7k_Bite.htm" Not found. | A critical startup configuration file was missing or too many were found. | Reinstall the software. | General System Startup |
| SET UP: FAILED TO LOAD BIT FILE "<x>" for Device <x>, Enumerator <x> | A critical startup configuration file was missing or too many were found. | Reinstall the software. | General System Startup |
| Socket initialization failed, error code <x> | Windows has failed to allocate a network resource. | Check network permissions and resources or reinstall operating system. | General System Startup |

Warning Messages – Master List in Alphabetical Order

| Message | Description | Action | Type |
| --- | --- | --- | --- |
| ***** (xxx) Dropping out-of-order ping #<x> | The data processing chain has been delayed and stale data has been discarded. | Check sonar processor hardware and resources. | Data Processing |
| 7K Bite File cannot be Read | Software installation is corrupt. | Restore factory defaults. Reinstall software. | Configuration |
| Cannot Open or Create UDP Configuration file | Software installation is corrupt. | Restore factory defaults. Reinstall software. | Configuration |
| Cannot Read UDP Configuration file <x> Error | Software installation is corrupt. | Restore factory defaults. Reinstall software. | Configuration |
| DMA Transfer Timeout. <x> of <x> bytes received | One or more pings have been lost due to an error on the uplink between the receiver and the software. The system is becoming unstable and data transfer is corrupted on the PCI bus. | Reboot. Check sonar processor hardware. | Data Processing |
| ERROR : NSD | An error has occurred in reading of BITE data. The data buffer has overflowed. | No action needed. | BITE Data |
| Non Sonar Data FIFO overflow (x packet) | An error has occurred in reading of BITE data. The data buffer has overflowed. | No action needed. | BITE Data |
| PING <x> :( Lost FPGA sync...sample | One or more pings have been lost due to an error on the uplink between the receiver and the software. The system is becoming unstable and data transfer is corrupted on the PCI bus. | Reboot. Check sonar processor hardware. | Data Processing |
| Roll Compensation is not available. | The input data for motion is not available or the feature is not enabled. | Check device inputs and cabling. | Input Signal |
| Subscriber send() failed | The software was unable to send data to a connected client. | Check network resources and user permissions. | Data Subscription |
| Subscriber UDP sendto() failed | The software was unable to send data to a connected client. | Check network resources and user permissions. | Data Subscription |
| Subscribers outbound MMF buffer memory overwrite error! | The link to the user interface running on the sonar processor has detected a non-critical error. | Check system resources, shutdown all non-critical software applications. Reboot processor. | Data Subscription |
| Subscribers outbound MMF buffer read error! | The link to the user interface running on the sonar processor has detected a non-critical error. | Check system resources, shutdown all non-critical software applications. Reboot processor. | Data Subscription |
| TCP Server - Unknown packet type | A connected client has sent an unrecognized command to the software. | Review client software source code or definitions for correct implementation of Teledyne RESON 7k Data Format. | Network Components |
| TCP Server - Warning max connections reached | The network module has become saturated with client connections and will refuse any further connections. | Close some of the connected clients. | Network Components |
| Trigger out length (x) is invalid | An unsupported value was received for trigger setup. | Check settings in user interface. | External Trigger |
| Trigger out offset (x) is invalid | An unsupported value was received for trigger setup. | Check settings in user interface. | External Trigger |
| Trigger out offset + length (x) exceed limit of 20ms | An unsupported value was received for trigger setup. | Check settings in user interface. | External Trigger |
| UDP Configuration file not available | Software installation is corrupt. | Restore factory defaults. Reinstall software. | Configuration |
| UDP Server - Unknown packet type | A connected client has sent an unrecognized command to the software. | Review client software source code or definitions for correct implementation of Teledyne RESON 7k Data Format. | Network Components |
| UDP Server - Warning max connection reached | The network module has become saturated with client connections and will refuse any further connections. | Close some of the connected clients. | Network Components |
| WARNING: PPS Time-stamp error | The internal clock may not be disciplined properly resulting in incorrect time stamp of data. | Check 1PPS connections and GPS output. If uncorrected sonar data cannot be time stamped to 1 millisecond | System Stability |
| WARNING: Reduce Ping Rate or Range - DROPPING ping | The system is overloaded processing data. A ping has been lost because the processing queue is full. | Restart the software. Reboot PC. | System Stability |

Installation Instructions

Components List

Table 8: SeaBat T51-R Components

| Part Number | Description |
| --- | --- |
| 1010498-DK | Projector, Bathy, DF400/800kHz, TC2186, TDY logo, max depth 400m |
| 1010473-DK | Receiver, EM7222, 512Ch 400/800kHz, Ti |
| 85160033 | Mounting Bracket Standard, T50, EM7218/TC2181, Composite |
| 1016786-DK | T_ASSY,RSP+ Dual head |
| 1008776-DK | CABLE ASSY, RSP/PSP to Rx, 25m, 7K |
| 85002229 | CABLE ASSY, RSP/PSP to DF-Tx, 25m, 7K |
| 1014438-DK | SHIP CASE, KIT, STENCILED, WET END,T50 |
| 1011057-DK | SHIP CASE, KIT, STENCILED, RSP, INS, BRACKET |
| 1008978-DK | Memory Stick, USB, >=16 GB, Crosair V3, tube, silver |
| 1008979-DK | Strap, black, Teledyne Marine white logo |
| 1016061-DK | Quick Reference Guide, SeaBat T51-R |
| 1008951-DK | Quick Installation Guide, SeaBat UI for T-Series |
| 1011088-DK | COTS, Cable, LAN Straight RJ45 to Straight RJ45 CAT6 Shielded Cable, 7.5m |
| 60000002 | Cable, power, 3-Cond |
| 88100610 | Cable, power, 230 volt UK |
| SB9999/00057 | Cable, power, 230 volt EU |
| 1008075-DK | COTS, Cable,Conf,RF,3m,BNC,M, BNC,M |
| 1010933-DK | Kit, screws, spacers and allen keys for T20/T50 series mounting bracket |
| 1007252-DK | Wet connector lubrication kit |
| Optional components |  |
| 85160040 | Fairing Add-on, Mounting Bracket Standard, T50 |
| Options for configurations with integrated INS |  |
| 1016787-DK | T_ASSY,RSP+ Dual head with INS |
| 1010790 | T_ASSY,IMU-20 for T-Series sonar |
| 1010763 | T_ASSY,IMU-30 for T-Series sonar |
| 1010795 | Cable assy TCS-10003 PMCIL-8-MP X=10m |
| 1010794 | Cable assy TCS-10003 PMCIL-8-MP X=25m |
| 1010679 | Cable assy TCS-10003 PMCIL-8-MP X=50m |
| 1010889 | Cable assy TCS-10003 PMCIL-8-MP X=100m |
| 1010898 | COTS, Antenna Trimble, GNSS |
| 1010752 | Cable Assy, Antenna cable, TNC to TNC on LMR-240 Cable, XX=10m |
| 1010756 | Cable Assy, Antenna cable, TNC to TNC on LMR-240 Cable, XX=20m |
| 1010757 | Cable Assy, Antenna cable, TNC to TNC on LMR-240 Cable, XX=40m |

Installation Checklist

NOTE

This checklist only covers the most basic, standard system installation. Depending on the system configuration or specific customer requirements, this list may or may not apply.

Follow these basic steps to install and configure your new SeaBat sonar system:

Install the dry-end components:

Rack-mounted sonar processor

Install the wet-end components:

Receiver unit

Projector unit

IMU (optional)

Install and connect cables:

Rack-mounted sonar processor to receiver cable

Rack-mounted sonar processor to projector cable

Auxiliary sensor cables

Rack-mounted sonar processor power cable

Rack-mounted sonar processor to laptop/PC Ethernet cable

Power up the system

Test the system

For information on system configuration and operation, refer to the Sonar UI User Manual (see Appendix I Reference Documentation).

CAUTION

Before turning the system on, make sure that there is no significant difference between the temperatures of the equipment and environment in order to prevent condensation.

System Interfacing

Interfacing Diagram for Standard Configurations

Please refer to the interfacing diagram in Figure 45: T51-R, interfacing diagram.

For further details, please refer to:

Appendix B.4 Dry-End Installation

Appendix B.7 Wet-End Installation, incl.

B.7.7 Hardware

Section 3.3 Starting the Sonar User Interface

Figure 45: T51-R, interfacing diagram

Interfacing Diagram for Configurations with Integrated INS

Please refer to one of the interfacing diagrams:

Figure 46: T51-R with integrated Applanix AP+ INS, interfacing diagram

Figure 47: T51-R with integrated INS, interfacing diagram

For further details, please refer to:

Appendix B.4 Dry-End Installation

Appendix B.7 Wet-End Installation, incl.

B.7.7 Hardware

Section 3.3 Starting the Sonar User Interface

Figure 46: T51-R with integrated Applanix AP+ INS, interfacing diagram

Figure 47: T51-R with integrated INS, interfacing diagram

Dry-End Installation

Rack-mounted Sonar Processor (RSP)

Table 9: RSP Technical Specifications

| Specification | Value |  |
| --- | --- | --- |
| Power requirements | 100-230VAC, 50/60Hz |  |
| Power consumption | Single: 400W max. Dual: 600W max. |  |
| Data communication, 2x | Gigabit Ethernet |  |
| Mounting | 19-inch rack |  |
| Dimensions | Height: 88.0mm Width: 478.0mm Depth: 461.6mm |  |
| Weight (excl. cables) | 12.3kg (without INS) 13.8kg (with INS) |  |
| Temperature | Operating: -5° to +45°C Storage: -30° to +70°C |  |
| Humidity | Max. 95% RH noncondensing |  |
| Vibration | ±1mm at 7m/s2 |  |

CAUTION

The RSP is designed and assembled specifically to accommodate the software packages provided by Teledyne RESON and installed at the factory. Installing additional software on this machine can result in decreased performance and/or system malfunction. Do not open, change, add, or remove anything (incl. software) on the RSP. Noncompliance may affect warranty situation (see Appendix K Standard Warranty Information)

After the RSP is installed, the only interaction required should be to press the power switch and energize the system.

Sonar Processor Mounting

The sonar processor should be rack mounted. It must be in an area that is within easy reach of the operator workstation connections and the interconnect cables. Allow clearance around the unit to provide adequate air circulation and easy access to the power switch.

CAUTION

Do not mount the sonar processor on the side, as this can damage the internal workings of the unit. Always mount with the bottom down. Noncompliance may affect warranty situation (see Appendix K Standard Warranty Information).

Sonar Processor Cable Connections

All dry-end cables provided by Teledyne RESON come with factory-installed connectors to be attached at the appropriate port(s) on the rear panel of the rack-mounted sonar processor (RSP).

CAUTION

Do not disconnect any cables from the rear of the RSP while the unit is running. This can damage the internal workings of the unit.

The critical connections are clearly marked on the RSP and shown in Figure 48, Figure 49, and Figure 50.

| (1) Mains power input 100-230VAC (2) Protective earth screw for equipotential connection (3) 1PPS out (4) Trigger in (5) Trigger out (6) Projector connection(s) (Use TX1) (7) Receiver connection(s) (Use RX1) | (8) Alternate SVP serial port w/LED (COM 3) (9) Main SVP serial port with power (10) Ethernet for laptop/PC Ethernet for NAS (a1) Time in serial port UTC/ZDA w/LED (a2) Motion in serial port w/LED (a3) COM 4 serial port w/LED (RS-232) (a4) 1PPS in |
| --- | --- |
| Figure 48: RSP Standard Configuration – Cable Connections |  |
| (1) Mains power input 110-230VAC (2) Protective earth screw for equipotential connection (3) 1PPS out (4) Trigger in (5) Trigger out (6) Projector connection(s) (Use TX1) (7) Receiver connection(s) (Use RX1) | (8) Alternate SVP serial port w/LED (COM 3) (9) Main SVP serial port with power (10) Ethernet for laptop/PC Ethernet for NAS (b1) Antennas 1 and 2 (b2) IMU (b3) Motion out serial port w/LED (b4) COM serial port w/LED (RS-232) |
| Figure 49: RSP Applanix AP+ INS Configuration (optional) – Cable Connections |  |
| (1) Mains power input 110-230VAC (2) Protective earth screw for equipotential connection (3) 1PPS out (4) Trigger in (5) Trigger out (6) Projector connection(s) (Use TX1) (7) Receiver connection(s) (Use RX1) (8) Alternate SVP serial port w/LED (COM 3) | (9) Main SVP serial port with power (10) Ethernet for laptop/PC Ethernet for NAS (b1) Antennas 1 and 2 (b2) IMU (b3) Time out serial port UTC/ZDA w/LED (b4) Motion out serial port w/LED (b5) COM 3-5 configurable serial ports w/LED (b6) GNSS serial port |
| Figure 50: RSP INS Configuration (optional) – Cable Connections |  |

WARNING!

For safety reasons the SeaBat™ must be connected to protective earth. The protective earth pin in the mains appliance inlet is the protective conductor terminal of SeaBat™.

NOTE

The LED above the serial port will blink in accordance with the registered data repetition rate. Green for data output, yellow for data input.

NOTE

The DB9 COM port LEDs are powered by the RS-232 lines and will therefore be lit, when these lines are active, even if the RSP is not connected to power.

NOTE

If a power failure occurs on the mains input the sonar may shut down. To recover from the shutdown, press the power button on the RSP.

| WARNING! |  |  |
| --- | --- | --- |
|  | The projector connection must be handled with great care, as the output power voltage is hazardous to human safety. Do not pull out the cable with power on. |  |

| CAUTION |  |
| --- | --- |
| To disconnect the projector/receiver/SVP cable from the sonar processor, hold on the groove part of the connector (at the arrow) and pull gently outwards. Do not pull directly on the cord, as this will damage both cable and the RSP. Do not bend the cables beyond the recommended limits when positioning the cables. Flexing the cable to less than the specified bend radius will damage the cable. |  |

Rack-mounted Sonar Processor Power Button and BITE LED

The power button and a BITE LED are located on the front of the RSP.

Power button light indicator:

Blue: when turned on

Red: on standby

No light: not connected to power

BITE LED:

No light: 7kCenter not started
(system not booted/applications not started)

Green: Normal (everything in order)

Yellow: BITE warning (investigate UI on surveyor laptop/PC to find the cause)

Red: BITE error (investigate UI on surveyor laptop/PC to find the cause)

Red flashing: 7kCenter not responding (e.g. not started/not running/crashed)

Auxiliary Sensors

NOTE

The auxiliary sensors required depends on the application for which the sonar is used; as a general rule an accurate online sound velocity value is required.

The rack-mounted sonar processor uses DB9 connectors as inputs for COM ports. Handshake and control wires have been looped to ensure that sensor devices with flow control/handshake can send data to the sonar processor.

Standard SeaBat T51-R Configurations

Figure 51: SVP, Motion, and Time Standard Connections

SeaBat T51-R Configurations with Integrated Applanix AP+ INS

Figure 52: SVP and Motion Connections with Integrated AP+ INS (optional)

SeaBat T51-R Configurations with Integrated INS

Figure 53: SVP, Motion, and Time Connections with Integrated INS (optional)

Sound Velocity Input

| To connect a sound velocity probe to the RSP: | Table 10: RSP Port “SVP Com 3” DB9 Pin; RS-232. 1-4-6; Shorted internally. 2; RxD. 3; TxD. 5; GND. 7-8; Shorted internally. 9; -. |
| --- | --- |
| - Using a DB9 female connector, connect the SVP data to pin 2 and ground to pin 5. - Connect the DB9 to the port marked SVP Com 3 on the rear panel of the RSP (see item 8 in Figure 48, Figure 49, or Figure 50). or - Connect a Teledyne RESON standard SVP cable with a Fischer connector to the port marked SVP (see item 9 in Figure 48, Figure 49, or Figure 50). |  |

Time Input

| To connect an external clock to the RSP: | Table 11: RSP Port “Time In Com 1” DB9 Pin; RS-232. 1-4-6; Shorted internally. 2; RxD. 3; TxD. 5; GND. 7-8; Shorted internally. 9; -. |
| --- | --- |
| - Using a DB9 female connector, connect the time data to pin 2 and ground to pin 5. - Connect the DB9 to the port marked Time In Com 1 on the rear panel of the RSP (see item a1 in Figure 48). - Using a BNC male connector, connect the 1PPS pulse to the port marked PPS IN on the rear panel of the RSP (see item a4 in Figure 48) for setting the internal clocks of the sonar. |  |

Time Output (for optional INS configurations)

| To connect an external clock to the RSP: | Table 12: RSP Port “Time Out Com 1” for optional INS DB9 Pin; RS-232. 1-4-6; Shorted internally. 2; -. 3; TxD. 5; GND. 7-8; Shorted internally. 9; -. |
| --- | --- |
| - Using a DB9 female connector, connect the time data to pin 3 and ground to pin 5. - Connect the DB9 to the port marked Time Out Com 1 on the rear panel of the RSP (see item b3 in Figure 50). |  |

Motion Input

| To connect a motion sensor to the RSP: | Table 13: RSP Port “Motion In Com 2” DB9 Pin; RS-232. 1-4-6; Shorted internally. 2; RxD. 3; TxD. 5; GND. 7-8; Shorted internally. 9; -. |
| --- | --- |
| - Using a DB9 female connector, connect the motion data to pin 2 and ground to pin 5. - Connect the DB9 to the port marked Motion In Com 2 on the rear panel of the RSP (see item a2 in Figure 48). |  |

Motion Output (for optional INS configurations)

| To connect a motion sensor to the RSP: | Table 14: RSP Port “Motion Out Com 2” for optional INS DB9 Pin; RS-232. 1-4-6; Shorted internally. 2; -. 3; TxD. 5; GND. 7-8; Shorted internally. 9; -. |
| --- | --- |
| - Using a DB9 female connector, connect the motion data to pin 3 and ground to pin 5. - Connect the DB9 to the port marked Motion Out Com 2 on the rear panel of the RSP (see item b4 in Figure 50). For Applanix AP+ INS, use Com 1 (see item b3 in Figure 49). |  |

Standard Serial Ports

Table 15: RSP Ports “Com 4”, “Com 3-5” for optional INS

| DB9 Pin | RS-232 |
| --- | --- |
| 1-4-6 | Shorted internally |
| 2 | RxD |
| 3 | TxD |
| 5 | GND |
| 7-8 | Shorted internally |
| 9 | - |

SeaBat T51-R Cables

A standard cable set is provided to connect the SeaBat T51-R system components.

Figure 54: SeaBat T51-R System Cables

Receiver to RSP (1): This cable allows power and data to be passed between the receiver and RSP. The cable has a standard length of 25 meters and a minimum bend radius of 100mm.

Projector to RSP (2): This cable allows power and data to be passed between the projector and RSP. The cable has a standard length of 25 meters and a minimum bend radius of 73mm.

RSP to laptop/PC (3): This cable is a CAT 6 unshielded cable for data connection between the laptop/PC and the RSP. The cable has a length of 6.1 meters (two cables are supplied with the system).

RSP to IMU (optional) (4): This cable passes sensor data from the IMU to the RSP. The cable is available in lengths of 10, 25, 50, and 100 meters and has a minimum bend radius of 92mm.

RSP to antennas (optional) (5): These cables pass position and time to the RSP. The cables are available in lengths of 10, 20, and 40 meters and have a minimum bend radius of 20mm (installed) or 64mm (repeated).

CAUTION

The units must not be suspended from own outside plugs and cables. Ensure that strain relief on the cable (cable grip) is used. If the cable is very long, use an external wire or cable covered with steel flex.

For cable options, see appendix J.4 SeaBat Cables.

Wet-End Installation

EM7222 Receiver Unit (PN 1010473-DK)

Table 16: EM7222 Receiver Unit Technical Specifications

| Specification | Value | Front view Rear view |
| --- | --- | --- |
| Material | Titanium (grade 2) lid Polyurethane acoustic window |  |
| Dimensions | Height 86.8mm Length: 460.0mm Depth: 102.0mm |  |
| Weight | Air: 8.7kg Water: 5.3kg |  |
| Temperature | Operating: -2° to +36°C Storage: -30° to +70°C |  |

The EM7222 receiver unit mounts to the vessel or mounting bracket with four M8 screws already in place in the receiver. The cable can be arranged as needed in the grooves.

For ease of mounting, the ends of the housing are labeled P and S for Port and Starboard, respectively. The connector is placed on the port side.

CAUTION

Ensure free access to water inlet holes (on starboard side) for pressure equalization.

TC2186 Projector Unit (PN 1010498-DK)

Table 17: TC2186 Projector Unit Technical Specifications

| Specification | Value | Front view Rear view |
| --- | --- | --- |
| Material | Titanium (grade 2) lid Epoxy acoustic window |  |
| Dimensions | Height: 88.4mm Length: 280.0mm Depth: 79.0mm |  |
| Weight | Air: 4.2kg Water: 3.0kg |  |
| Temperature | Operating: -5° to +36°C Storage: -30° to +70°C |  |

The TC2186 projector unit mounts to the vessel or mounting bracket with M8 Hex screws already in place in the projector.

The cable connection is not exposed, as the unit’s design provides an integrated cable protection.

Wet-End Cable Connections

| CAUTION |  |
| --- | --- |
| Female connectors must be lubricated at least every 5 matings to avoid damaged male connectors (as shown here). Follow the instructions provided with the wet connector lubrication kit (see the components checklist in appendix B.1 above). Noncompliance may affect warranty situation (see Appendix K Standard Warranty Information). |  |

The cable connections to the wet-end units should be made prior to immersing the units in water.

| WARNING! |  |
| --- | --- |
|  | The projector connection must be handled with great care, as the output power voltage is hazardous to human safety. Do not pull out the cables with power on. Output voltage: 280Vrms ≈ 800Vp-p |

A light lubrication coating on the connector face will ensure sealing (see section 2.2.2.2 Wet-End Connectors).

Unused connectors should be sealed with dummy plugs.

Verify that all cables are securely fastened and that strain relief is provided to prevent cable damage.

CAUTION

Take care to avoid impact to the faces of projector and receiver units, as damage could degrade system performance.

Sonar Head Assembly Orientation

The sonar head assembly should be mounted with the faces of the arrays oriented vertically downwards.

Ideally, the receiver is mounted across track with the connector to the port side, while the projector is mounted along track and aft of the receiver.

If necessary, the projector may be oriented forward or the head rotated away from vertical with allowances made in the processor setup and data acquisition software.

CAUTION

An alternate mounting makes the projector more susceptible to impact damage and collection of trailing debris.

The relative distance between the projector and receiver face (i.e. Z-value) needs to be known with an accuracy better than 2mm.

The projector and receiver (seen from the respective reference points) must not be further than 250mm apart in the XY plane.

The actual sonar head assembly installation must be reflected in the user interface settings (see Appendix I Reference Documentation).

Sonar Reference Point

The following axis convention is applied for the sonar reference point convention:

| X = across ship, points to starboard: | + Starboard - Port |
| --- | --- |
| Y = along ship, points forward: | + Forward - Aft |
| Z = vertical, points up: | + Up - Down |

All output is related to one single sonar reference point, normally defined in the data acquisition software as a node or offset point. The point is in the center of the receiver face in both X and Z directions and in the center of the projector in the Y direction.

For details, refer to Figure 67: T51 Bracket Installation: Sonar Reference Point.

It is important to note that the sonar reference point may not be a physical point on the sonar system.

Wet-End Mounting

Take care that the sonar head assembly is electrically isolated from the structure to avoid corrosion problems.

| The system includes a mounting bracket that is intended for mounting flat or via a pole on the ship’s hull, or on a pole over the side. See also appendix I.2 SeaBat T51-R Design Documents). | - Bracket: EM7222 receiver TC2186 projector |
| --- | --- |

For straightforward positioning of the receiver, both bracket and receiver are marked with P and S for Port and Starboard, respectively.

- Plane surface installation:

- Use the 4 x Ø13.5 mounting holes to mount the bracket directly on the ship’s hull or similar (see Figure 65).

- Secure with 4 x M12 screws (from the inside of the bracket).
Length of screws depends on each individual installation.

- Mount the wet-end components and secure with the screws already in place in the units.

- Pole mount installation:

- Mount the wet-end components on the bracket and secure with the screws already in place in the units.

- Align the holes of the pole flange with the four holes around the cable ring of the bracket (see Figure 65).

- Secure with 4 x M12 screws (from the outside of the bracket).
Length of screws depends on each individual installation.

CAUTION

Make sure the screw length does not exceed 24mm from entry of bracket.

The head should be immersed in water to a depth of at least 0.8m. This ensures that noise from surface waves and turbulence generated by the mounting structure does not degrade performance.

Mounting on a surface vessel may be over the side, over the bow, hull mount, or through a moon pool.

Ensure that submersion depth is sufficient to avoid shadows from the keel, hull, or other obstacles.

NOTE

Fairings can also be used to improve hydrodynamic performance (see section 5.10 Hydrodynamic Drag and appendix I.2 SeaBat T51-R Design Documents). The sonar head is not designed to be hydrodynamic at speeds in excess of two knots. Vibration and oscillation will become evident, if the projector and receiver are not rigidly mounted. To minimize the unsupported length, place a support as close to the waterline as possible.

Hardware

| The Bracket Configuration on the Hardware pane controls the one-time transmit array and projector settings. These settings should be set to reflect the initial installation of your system and should not require modification unless the physical installation changes. For instructions on navigating the user interface, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation). | Click the Hardware tab to the right in the user interface to access the Bracket Configuration settings. |
| --- | --- |

For details on sonar head assembly orientation, see appendix B.7.4 above.

Select as appropriate either one of the preset mounting bracket configurations or enter custom values.

Preset Configurations:

Normal: To be used when the sonar head is installed with the receiver forward and projector aft, that is when the indication for port ("P") on the receiver/bracket is to the port side.

Reversed: To be used if the sonar head is rotated 180 degrees and the indication for port ("P") on the receiver/bracket is to the starboard side. This option reverses the beams thereby normalizing the image and the generated data output.

Customer Configuration:

For nonstandard installations, the XYZ offset of the receiver reference point to the projector reference point must be defined. For details on reference points, see appendix B.7.5 above.

CAUTION

Incorrect offset settings could severely affect data quality.

Head Tilt:

The Head tilt feature allows the user to normalize the display, gates, and bottom detect processing to reflect the physical orientation of the sonar head. For examples on how to use this feature, see section 4.3.4 Head Tilt.

The angular value is entered in degrees (+ Port up/- Port down) and must be between -45° and 45° in increments of 0.5°.

Remote Operation

The 7k series software architecture is built on a client/server concept. The 7k Center running on the RSP controls the sonar performs data processing, and exports data to external clients connected via TCP/IP or UPD/IP. The clients should control the 7k Center via remote SUI or use of customized control applications based on DFD. Refer to the appendix section in the Data Format Definition document for example code for writing the control application (see appendix I.1 SeaBat T51-R Documentation).

For information on system configuration and user interface navigation, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

System Check

Once the SeaBat T51-R has been installed, follow these steps to ensure that the system is working properly. For information on the settings, refer to the RESON Sonar UI User Manual (see Appendix I Reference Documentation).

Configuration

Confirming that the system has booted up in the right configuration demonstrates that communications between wet end and processors are functional, and that the ID chip from the receiver and projector are read correctly.

The configuration is presented in the top of the Sonar UI window.

The serial numbers are displayed in the Service menu under the Config tab. Find the serial numbers on the units:

  - Receiver and projector: stamped into the housings.

  - RSP: on the front label.

Figure 55: Configuration Check via Sonar UI Service Menu

Receiver and Projector Test

The following steps explain how to test the receiver and projector.

CAUTION

The projector must always be tested in water, whereas the receiver may be tested in air.

- Ensure that all system cable connections are mated properly.

- Apply power to the system. The Windows® operating system will boot up and the 7kCenter (sonar control software) will start automatically.

- Start the Sonar UI by double-clicking the icon on the operator desktop and select the correct IP address in the toolbar (see section 3.3 Starting the Sonar User Interface).

- Verify that the Power setting on the Main pane is set to OFF.

Receiver Test (performed in air or water)

CAUTION

If tested in air, ensure that the projector is submerged. The projector must not ping in air.

- Once the system has been energized, follow these steps to verify that the system is operating correctly:

  - Set Range to 150.

  - Start pinging the transmitter by setting Max Rate on the Main pane.

  - Set Gain to 30.

  - Select Blue-Yellow from the Sonar Palette drop-down menu on the Display pane.

  - Perform a receiver check by vigorously rubbing the face of the receiver unit with fingertips while monitoring the display. A noise pattern should be barely visible on the display.

Projector Test (always performed in water)

CAUTION

Transmitting into air for any period of time may result in the permanent damage of the ceramics of the projector and consequently degraded performance. Always perform this test with the receiver and projector in water. A small container will suffice, as long at the units are fully submerged.

- Once the system has been energized, follow these steps to verify that the system is operating correctly:

  - Set Range to the lowest value, 3/5m, on the Main pane.

  - Set Gain to 83dB.

  - Set Pulse Length to 150µs (CW).

  - Set Power to the lowest value.

  - Verify that echoes are present in the wedge display.

  - Set transmit Power to OFF.

  - Shut down the system.

Integrated INS – Applanix AP+

This appendix explains the setup steps for a SeaBat T51-R with Applanix AP+ GNSS-INS.

Access to the AP+

- From a browser, access the AP+ by entering its IP address.

The default is 10.11.10.11. Ensure that the IP range of the network card from the PC you are using is in the same range.

- Log in with username and password.

The serial number of the AP+ unit is the password as shown in the top right corner of the application window.

| Username | admin |  |
| --- | --- | --- |
| Password | AP+ serial number |  |

GNSS Antenna Types

- In the left-hand navigation pane, click Receiver Configuration.

- For Position Antenna (primary antenna), select the Aero Antenna AT1675-540TS under GNSS Antenna Type.

- The Antenna height should be 0.0 and the correct measurement method set to “Antenna Phase Center”. The offset (used in appendix C.4.4 below) should then also be measured to the phase center.

- Repeat antenna type selection for Vector Antenna (secondary antenna).

Set Up the GAMS Baseline

- In the left-hand navigation pane, click Receiver Configuration.

- Under Navigation Controls, keep the settings as shown in this screen shot.

- Observe for Remote IMU Type:

  - IMU Type 90 = AP+20/WaveMaster.

  - IMU Type 82 = AP+30/OceanMaster.

  - IMU Type 94 = AP+40/OceanMaster+.

- Click OK to confirm the settings.

You will be prompted to reboot the receiver to apply the new settings to the AP+.

NOTE

Observe that the Applanix AP+ uses a different axis convention than Teledyne. Table 18: Applanix AP+ and Teledyne Axis Conventions AP+; Teledyne. Positive X; Forward; To starboard. Positive Y; To starboard; Forward. Positive Z; Down; Up.

INS Remote Settings

| CAUTION |  |
| --- | --- |
| Only the INS Remote section is relevant. Ignore the INS Onboard section. |  |

This part is crucial for the definition of the output of the data. There are different ways to define the output data and offsets; the route taken depends on where you want the georeferencing of data to be done: a) at the AP+ sensor end or b) in the acquisition/processing software.

The method b) explained here is Teledyne’s preferred way.

Let the Common Reference Point (0,0,0) be the reference point of the IMU.

  - Use this definition of CRP in your acquisition software as well, so there is no confusion in-between software applications.

Let the acquisition software do the georeferencing.

  - That means that we are only concerned with the position/motion/heave data from the AP+ calculated at the IMU’s location.

  - Then in the acquisition software, we assign the IMU’s offset to the incoming data, and the MBES’ offset to the T20 (sonar reference point).

Graphic Prerequisites

| The size of the vessel is just for display purposes. Keep the Reference Location from Center of Vehicle at zero (0). |  |
| --- | --- |

Reference

| The reference offsets can be kept at zero (0.0), as we do not use the vessel lever arm. |  |
| --- | --- |

IMU Remote

| We are using the IMU as the Reference (i.e. CRP). So we need to set the offset from the IMU Reference Mark to the IMU Sensing Center. (See Table 19 below). If the angular misalignment of the IMU is dimensionally controlled or otherwise known, add the values here. If you have the misalignment angles, remember that Applanix uses a different axis convention (see Table 18). |  |
| --- | --- |
| - Also keep in mind that the Tait-Bryan order needs to be applied. This is particularly important, if there is a large rotation around Z. - Contact Teledyne support for further assistance. |  |

Table 19: Offsets from IMU Reference Mark to IMU Sensing Center

| Applanix IMU Type | Teledyne IMU | X (+fwd) | Y (+stbd) | Z (+down) |
| --- | --- | --- | --- | --- |
| 42 | 1010790 IMU-20 for T-Series sonar, Inertial Measurement Unit | N/A – not supported for AP+ |  |  |
| 64 | 1010763 IMU-30 for T-Series sonar, Inertial-Measurement Unit | N/A – not supported for AP+ |  |  |
| 82 | 1013992 IMU-20-1 for T-Series sonar, Inertial-Measurement Unit | 0.000m | 0.000m | +0.046m |
| 89 | 1016020 IMU-30-1 for T-Series sonar, Inertial-Measurement Unit | 0.000m | 0.000m | +0.046m |
| 82-30 | 1016878 IMU-30-2 for T-Series sonar, Inertial-Measurement Unit | 0.000m | 0.000m | +0.046m |
| 90 | 1017457-DK IMU+20 for T-Series sonar Inertial-Measurement Unit | -0.009m | +0.012m | +0.069m |
| 94 | 1018314-DK IMU+40 for T-Series sonar Inertial-Measurement Unit | +0.008m | -0.009m | +0.079m |
| NOTE: | Use the Applanix axis convention as described in this table. |  |  |  |

GNSS Antenna

| Enter the Offset from IMU to the Primary GNSS antenna (phase center). Enter the accuracy of the measured offsets (1σ). |  |
| --- | --- |

Sensor 1 / Sensor 2

| Depending on the selected output data format, the data will be valid at Sensor 1 or Sensor 2. In our scenario we only consider Sensor 1, and as we want to get the data valid at the IMU, we keep Sensor 1 offsets at zero (0.0). |  |
| --- | --- |

Heave Definition

| Set the heave bandwidth. - In general terms, a small vehicle and choppy conditions will require a heave bandwidth closer to 9. If you know the Center of Rotation with respect to the Reference (i.e. IMU), use it here. - Observe that in most (small) vehicles, the Center of Rotation is at the IMU/MBES location or otherwise pretty close, in which case the offsets should be set at zero (0.0). |  |
| --- | --- |

I/O Configuration

Select the outputs from the AP+.

The example below demonstrates the output of Trimble GSOF messages. The GSOF format replaces the now deemed legacy POS ICD message types.

The Port Summary page will present live Ethernet traffic as green.

Live serial data is not colored.

Set up the outputs for ZDA Time (Ethernet UDP) and Motion for Roll stabilization (currently only NMEA PASHR is available) and interface these in the Sonar UI. Motion data for Roll stabilization has always been serial output.

Select the GSOF messages (Ethernet UDP) as per your setup.

Following Teledyne’s method described in appendix C.4 above, we are outputting:

GSOF message 65 (Sensor 1)

Other options are:

GSOF 65+66  Sensor 1 Frame position and orientation + rms

GSOF 67+68  Sensor 2 Frame position and orientation + rms

GSOF 63+64  Vessel Frame position and orientation + rms

NOTE

The now deemed legacy POS ICD groups (1, 2, 102, 103, 104, 105, 111, 113) are still available, but groups 3, 10, 12, 13, 20 are currently not supported.

Time (NMEA ZDA Ethernet UDP)

|  | Ensure to use the Generic ZDA+1PPS driver. The PPS is hard wired in the sonar processor (and also available from the AUX connector). |
| --- | --- |

Motion for Roll Stab (NMEA PASHR Serial COM2)

|  |  | TSS1 is currently not supported by the AP+. Data for Roll Stab must be over a serial connection. Roll stab data are also parsed to the AUX connector. If other serial input/output is required from the AP+, Serial 1 must be used (e.g. CMR, RTCM corrections). Aux GNSS is not supported. |
| --- | --- | --- |

Trimble GSOF 65 (Sensor 1) Output for Teledyne PDS

NOTE

Optionally, interface to IO module for s7k recording or RRIO distribution.

True Heave Output for Teledyne PDS

True heave is not captured in the GSOF messages and remains available in the POS ICD groups 111+113.

GAMS Calibration

If the baseline vector in the vehicle frame can be measured to an accuracy of 0.01m or better, the parameters should be entered into the Navigation Controls window in the GUI and the 1-sigma value set to 0.002.

- In this case, the system will not attempt to calibrate the vector, but will use the entered values once they are saved.

If it is not possible or undesirable to accurately measure the baseline vector, and the vehicle is capable of sufficient dynamic motion to allow the baseline vector to be calibrated, the GAMS calibration feature as described below should be used.

Once successful calibration values are stored, this procedure does not need to be repeated, unless the physical location of the antennas is changed.

- The GAMS requires data from five or more satellites with a Positional Dilution of Precision (PDOP) of three, or less, to achieve a successful antenna installation calibration.

- Perform the antenna installation calibration at a time, when there is good satellite geometry (i.e. low PDOP and open sky).

- Ensure that GAMS is enabled under Receiver Configuration/Navigation Controls (see appendix C.3).

- Enter 0,0,0 for the X,Y,Z Primary to Secondary GNSS Baseline Vector lever arms to allow the system to estimate the values.

- Enter a 1-σ (one sigma) value equal to the approximate antenna separation (i.e. if the antennas are set 2m apart, use 2m).

- Enter the desired Target Heading Accuracy in degrees. We recommend a value of 0.5°, but do not make it too strict. The more accurate the target, the better the calibration will be, but smaller heading accuracies are more difficult and will take longer (especially with a lower position accuracy, e.g. if RTK is not available).

- Keep Auto-copy disabled.

- Maneuver to perform a number of turns (e.g. figure eights or s-turns) in an area where there is a minimum of GNSS multipath.

- Monitor the calibration progress in the Navigation Controls window to ensure the INS Heading Accuracy is getting smaller and approaching the Target Heading Accuracy. Continue until the INS Heading Accuracy meets, or exceeds, the Target Heading Accuracy. When the INS Heading Accuracy no longer gets smaller click the Calibrate button. See screenshot above (item 1).

- Now sail a straight line (i.e. constant heading) until the calibration completes.

- Once complete, copy the Current Estimate baseline vector across to the left entry fields by clicking the << button.

- Click the OK button at the bottom of the Navigation Controls page to save the settings.

- Reset the system, either by power cycling or by a soft reset via the Reset page under Receiver Configuration in the left-hand navigation pane.

Logging POSPac Data

- Go to Data Logging/Summary. Note that only internal logging is available. The data is initially stored onto the sonar processor (~10 GB available).

- Select Session/Configure.

The Enable checkbox is used to manually start and stop logging, while the Auto Log checkbox will automatically start logging data, when the system is turned on and tracking satellites.

The Schedule drop-down list allows the user to change the way the files are logged:

  - Always: data is logged in one single T04 file.

  - Continuous: data is logged in many files of size Duration.

  - Manual: only one file is logged of size Duration, after which logging is stopped.

- Select the data types and interval as below. Choose the preferred folder structure and the naming convention.

NOTE

If using a timestamp for the file names, when a new file is created within the same minute, the file will be appended. The result will be an IMUcount gap which later on requires processing around the gap in POSPac. Therefore, avoid a stop logging command right after a start logging command.

- Downloading the data:

  - Download the data from Data Logging/Data Files.

  - Select the desired files by clicking in the check box to the right of the file name and then selecting the Download Selected Files icon. The files will be downloaded to the Downloads folder on your PC. If multiple files are selected for download the web browser may prompt for confirmation, although the behavior is somewhat dependent on the web browser used.

PPS In and Trigger In

Electrical Specifications for PPS In and Trigger In

Table 20: PPS In and Trigger In Signals

| Parameter | Min. | Typical | Max. | Unit | Note |
| --- | --- | --- | --- | --- | --- |
| Absolute maximum signal voltage | 15 |  |  | V |  |
| Absolute maximum reverse signal voltage |  |  | -1 | V |  |
| Low voltage threshold | 1 | 1.6 |  | V |  |
| High voltage threshold |  | 1.6 | 1.8 | V |  |
| Input current (high) @2V @3V @5V @15V |  | 7 8 10 20 |  | mA mA mA mA |  |
| Input current (low) |  | 0 |  |  |  |
| Input delay |  | 400 | 1000 | ns | Pulse 0V to 2V, 50 Ohm |

Typical PPS In and Trigger Performance

Figure 56: Required PPS Pulse Width

Figure 57: PPS In Delay

Hints and Advice

The PPS and trigger inputs work with regular CMOS TTL outputs.

The drive capability of a standard TTL output is inadequate to drive PPS and trigger high.

When more systems use the same PPS source using BNC-T-adaptors it should be verified that its drive capability is sufficient.

For distribution of the PPS pulse, "PPS out" may be used instead of the BNC-T-adaptor.

Make sure the PPS source output is in accordance with the specifications listed in Table 20.

Operation is ensured if the source resistance is less than 470 Ohm, the drive voltage is greater than 3V, and the pulse width is greater than 1µs.

The PPS/trigger described is the externally available signal. The signal is conditioned before it is internally distributed.

Figure 58: PPS/Trigger Signal

Data Rates

Table 21: Theoretical Max. Data Rates – SeaBat T51-R Single Head

| System and number of beams | Utilized storage GB/hour | Max. data rate Mb/s | Utilized storage GB/hour |
| --- | --- | --- | --- |
| T51-R, 256 beams 34kS/s | ~130 | ~300 | ~5 |
| T51-R, 512 beams 34kS/s | ~250 | ~550 | ~6 |
| T51-R, 1024 beams 34kS/s | N/A | N/A | ~8 |
| T51-R, 256 beams 50kS/s | ~180 | ~400 | ~6 |
| T51-R, 384 beams 50kS/s | ~270 | ~600 | ~7 |
| T51-R, 1024 beams 50kS/s | N/A | N/A | ~9 |
| Records/data: Survey + interfaced sensor data: | Incl. water column |  | Excl. water column |
| Notes: |  |  |  |
| 1024 beams exceed the Ethernet data bandwidth. |  |  |  |

Network Configuration

System Network Configuration

The Rack-mounted Sonar Processor (RSP) is at delivery configured as follows:

Table 22: IP Configuration for RSP

| Network port | LAN 1 | LAN 2 | INS (optional) |
| --- | --- | --- | --- |
| IP address | 10.11.10.1 | 10.11.10.2 | 10.11.10.11 (fixed) |
| Subnet mask | 255.255.255.0 | 255.255.255.0 | 255.255.255.0 |

The CFE laptop/PC which runs the Sonar User Interface (SUI) should in standard systems be configured in the IP range 10.11.10.3-10.11.10.255 in the same subnet as the RSP.

Change IP Configuration

The IP configuration can be changed from the SUI installed on the laptop/PC:

| - Click the Application Settings button in the top right corner of the Sonar UI. |  |  |
| --- | --- | --- |
| - Unfold the Sonar network connections setting. - If the connection has been established, the Computer name is already stated; if not, enter SN+serial number of the RSP. Please find the serial number (with seven digits) on the front label on the RSP. |  |  |
| - Press Enter on the keyboard. After a few seconds the Configuration of LAN 1 and LAN 2 will be populated |  |  |

Manual or automatic configuration of LAN 1 and LAN 2:

Both LAN 1 and LAN 2 can be configured to obtain a manual fixed IP address or an IP address automatically from a DHCP (Dynamic Host Configuration Protocol) server.

  - The manual fixed IP setting is freely configurable for LAN 1, but for LAN 2 the range is limited to 10.11.10.2 to 10.11.10.9 with subnet mask 255.255.255.0.

If the IP for LAN 2 is configured outside this range it will revert to 10.11.10.2 with subnet mask 255.255.255.0 after the next reboot of the RSP.

The factory default setting is 10.11.10.1 for LAN 1 and 10.11.10.2 for LAN 2, each with subnet mask 255.255.255.0.

  - If Automatic IP address is selected, the RSP will request an IP from a DHCP on the network; if none is given, it will revert to alternative IP addresses, 10.11.10.1 for LAN 1 and 10.11.10.2 for LAN 2. The timeout before reverting is up to 1m30s.

- After configuring the desired IP addresses, click OK.

NOTE

If you change the IP addresses (within the same subnet), the connection to the RSP will be lost. Shut down the SUI and restart to re-establish the connection.

NOTE

The connection from the SUI to the RSP will be lost if the RSP moves to another SUBNET than the one the laptop/PC is located on. In that case, change the laptop/PC IP address to a valid IP address within the same subnet, and restart the SUI, and establish connection as described in section 3 System Startup and Shutdown.

NOTE

The outlined procedure should work even with an IP address configured in another subnet. In the event of a problem, connect a cable directly between the LAN 2 port on the RSP and the laptop/PC, change the laptop/PC to IP address 10.11.10.10 with subnet 255.255.255.0, and restart the SUI. It might be necessary to reboot the laptop/PC as well.

Configuration of Integrated INS (optional)

NOTE

For further details, refer also to the documents listed in Appendix I Reference Documentation.

POS MV – How to Perform a GAMS Calibration

MV-POSView – Standby Mode

GAMS calibration is performed using the MV-POSView software. Many installation settings may only be altered when the POS MV is in Standby mode. To switch to Standby, first ensure the Connect button is depressed, then click the Standby button.

10.11.10.11

GAMS Parameter Setup

Prior to a GAMS calibration any existing GAMS vector parameters should be zeroed out. Record the previous settings prior to zeroing. Choose Settings > Installation > GAMS Param. Setup.

An approximate value for the “Two Antenna Separation (m)” may be entered to assist the GAMS calibration.

The “Heading Calibration Threshold (deg)” parameter limits heading accuracy under which a GAMS calibration may be performed. Typically this should be set to 0.5.

The “Heading Correction (deg)” will typically be set at 0. This is for entering a post GAMS calibration gyro calibration parameter.

Ensure that the Baseline Vector parameters are zeroed.

Positioning Quality

The GAMS calibration is best performed with highest positional reliability. This means that Nav Status should read at best FIXED RTK or RTCM DGPS if RTK is not available.

Ensure that differential corrections are received before commencing a GAMS calibration.

GAMS Calibration – Procedure

- Maneuver the vessel into an area of clear water where a series of rapid turns and speed changes may be performed.

- If the system is in Navigate mode set it in Standby mode to clear the Kalman filter history.

10.11.10.11

- Click Navigate.

- Instruct the vessel to perform a series of sharp turns and speed changes, normally figures of 8. The greater acceleration sensed by the IMU in all directions the faster the Heading Accuracy will improve.

- During the vessel maneuvers, monitor the Heading Accuracy as indicated below. After clicking Standby it will momentarily reset to 55.000. During the vessel maneuvers it will start to decrease, at first rapidly then slowing as the accuracy reaches its limit.

10.11.10.11

- The GAMS calibration cannot be started unless the Heading Accuracy becomes lower than the Heading Threshold value set in the GAMS parameter setup dialog. When it does, the following message will appear: GAMS Ready Offline in the Status section in the top left. Continue vessel maneuvers until the Heading Accuracy is as low as it will go. Aim for 0.1 degrees.

- When the GAMS Ready Offline message appears and the Heading Accuracy is as low as it will go, instruct the vessel to come to a standstill (or transit slowly) on a constant heading.

- Choose Settings > GAMS Calibration Control > Start. The message CAL in Progress will appear.

- Continue to remain on standstill (or slow speed) on constant heading until the calibration is complete. The process will take less than 5 minutes.

- When the calibration has been successfully completed the CAL Complete message will appear for a brief period.

- Following this the GAMS Online message will appear and the Heading Accuracy will jump to a very low value.

10.11.10.11

- Choose Settings – Save Settings.

- Choose File – Save POS Config.

- Observe the results of the calibration in GAMS Parameter Setup dialog. The fields will have been automatically populated with the results.

- Note down the results.

- The process should be repeated twice and the results from each calibration compared. They should agree to better than 5mm for each baseline vector component.

POS MV – How to Set Up for a Multibeam Survey

This section briefly explains the POS MV installation parameters for a typical multibeam survey with reference to the different types of outputs and acquisition software settings.

Conventions

All sensor offsets and rotations are applied in the POS MV with respect to the Tate/Bryant co-ordinate system as illustrated below.

| Figure 59: Tate/Bryant Co-ordinate/Rotation System |  |
| --- | --- |
| Where: X axis is positive to bow Y axis is positive to starboard Z axis is positive down | X rotation is positive starboard down Y rotation is positive bow up Z rotation is positive to starboard |

Coordinate Frames

There are three coordinate frames to consider:

Reference: The reference frame is the frame in which all sensor offsets are measured, usually by dimensional control, whose origin is an arbitrary point on the vessel convenient to the survey method employed. The reference frame is assumed to be defined in the fore/aft direction by the lubber line.

Vessel: The vessel and reference frames are assumed by the POS MV to be co-aligned and differ only in the location of their origins. The origin of the vessel frame may be any arbitrary point, but for a multibeam survey the most rigorous approach is to make it the acoustic center of the sonar.

IMU: The IMU frame is the frame defined by the axes of the 3 mutually orthogonal accelerometers inside the IMU whose origin is the IMU reference point as marked on the top. The accelerometers are assumed to be co-aligned with the machined housing of the IMU.

Offsets – Example

The offsets marked in bold in Table 23 must be measured as a minimum (refer to Figure 60).

Table 23: Sensor Offsets

| Common reference point to: | X | Y | Z |
| --- | --- | --- | --- |
| Primary POS GPS | 19.984 | 4.691 | -25.199 |
| IMU reference mark | 13.708 | -2.261 | -4.033 |
| Sonar acoustic center | 25.580 | 3.972 | 5.535 |
| Center of rotation | 10.824 | 2.487 | 0.000 |
| Auxiliary GPS | 21.593 | 1.375 | -25.213 |
| Secondary POS GPS | 18.689 | -1.575 | -25.199 |

Figure 60 shows how the sensors may be located on a vessel and the dialogs presented in the following sections show how they would be entered in the POS MV.

Figure 60: Vessel Outline – Example

In the example in Figure 60 the common reference point (CRP) is an arbitrary point on the aft deck from which all sensor offsets and center of rotation are measured in the reference frame.

NOTE

Under ideal circumstances the IMU should be mounted on the center line of the vessel close to the center of rotation/gravity. The IMU reference mark may then become the CRP.

Alignments – Example

For rigorous setup the following alignments must be observed by dimensional survey. Only for temporary installations with short lever arms is it acceptable not to measure the IMU frame with respect to reference frame alignments.

Table 24: Sensor Alignments

| Common reference point to: | X | Y | Z |
| --- | --- | --- | --- |
| IMU frame w.r.t. reference frame |  |  |  |
| Sonar frame w.r.t. reference frame |  |  |  |

In the event that dimensional survey is not available (small or temporary installations), the IMU frame with respect to reference frame alignment fields may be left as zeros with the assumption that the IMU and reference frames are co-aligned.

Note that with large lever arms, positional errors due to the offset measurements being measured in a frame whose relationship to the IMU frame is not accounted for may be large. Essentially the offset measurements will be incorrect with respect to the frame that the POS MV assumes they are measured in. Therefore, making the assumption that the IMU and reference frames are co-aligned should only be used for small installations and as a last resort for large, temporary Installations.

Similarly, the sonar frame with respect to reference frame heading alignment may be left zero until the patch test has been completed unless measured by dimensional survey.

MV-POSView Software

The POS MV is configured by running the MV-POSView software.

10.11.10.11

Installation Parameters – Lever Arms & Mounting Angles

The offsets listed in Table 24 are entered in the MV-POSView configuration program by choosing Settings > Installation > Lever Arms & Mounting Angles. In the dialog, the term Ref refers to the reference point from which all offsets were measured in the reference frame (CRP in this case). The IMU and Primary POS GPS offsets are entered simply as they were measured from the CRP.

In the example given in this document, the sonar acoustic center has been chosen as the origin of the vessel reference frame. In the screen capture above the sonar acoustic center position has been entered in the “Ref. to Vessel Lever Arm” fields. This is essentially telling the POS MV where the origin of the vessel frame is – in the reference frame. In other words, its translation from the CRP. The NMEA position output from the POS is now valid for the vessel reference frame origin which is in this case the sonar acoustic center.

If the alignment of the IMU frame has been measured with respect to the reference frame by dimensional survey then these values are entered in the “IMU Frame w.r.t. Ref. Frame” fields. Care must be taken that they are entered correctly. X is a rotation about the X axis, i.e. X is the roll alignment, Y is the pitch alignment and Z is the heading alignment in the Tate/Bryant system. If the IMU and reference frames are perfectly co-aligned e.g. by means of an adjustable mounting plate then these values will all be zeros (as in the above example).

| Figure 61: IMU misalignment – Example | The diagram gives an illustration of a heading misalignment of the IMU frame with respect to reference frame in order to clarify the sign convention. Since the Tate/Bryant system dictates that a positive rotation from the reference frame is clockwise then, if the IMU frame is anti-clockwise of the reference frame, the IMU alignment is negative with respect to the reference frame. |
| --- | --- |

Installation Parameters – Sensor Mounting

Auxiliary sensors for which motion data is critical are configured in the Sensor Mounting tab. In every case the sonar offset from CRP in the reference frame should be entered in the Sensor 1 fields. If the orientation of the sonar has been measured by dimensional survey then the alignments should be entered in the “Sensor 1 Frame w.r.t. Ref. Frame” fields in the same way as described for the “IMU Frame w.r.t Ref. Frame” in appendix G.2.6 above. If this is done, then the multibeam calibration is essentially solving for any residual misalignments between reference and sonar frames. Following the patch test these values may be updated with the patch test results.

If an external differential correction source is supplying corrections to the POS, then no offsets are required for this. If the POS is receiving an external differentially corrected position message (e.g. GGA) as aiding, then the offsets for this antenna are required and are configured under Auxiliary GPS.

Installation Parameters – Tags, Multipath & 	AutoStart

Ensure that UTC time stamps together with AutoStart.

Installation Parameters – Heave Filter

Set the heave filter appropriately for the vessel. Choose Settings > Heave…

For a small vessel enter 10 seconds. For a large vessel enter 20 seconds.

POS MV Output– Serial Data

Serial data may be input or output from COMs 1 to 5 and may be configured using Settings > Input > Output Ports… from the MV POS View software.

COM ports 1 and 2 are preconfigured on delivery (output only).

COM ports 3-5 are available to the customer for other serial output or input.

Table 25: POS MV Output Serial Data – COM Ports 1-2

Input/Output Ports Set-up

COM 1 – ZDA Time Sync, NMEA (PRECONFIGURED)

To synchronize Teledyne PDS and/or a SeaBat 7k sonar it is necessary to generate an NMEA ZDA (UTC time) output message from one of the serial ports. Select the desired baud rate and communication parameters, choose NMEA from Output Select and click the $INZDA checkbox. Ensure that the update rate is set at 1Hz. COM 1 is only for output.

COM 2 –Attitude/Motion, Binary (PRECONFIGURED)

To synchronize Teledyne PDS and/or a SeaBat 7k sonar it is necessary to generate a binary motion output from one of the serial ports. Select the desired baud rate and communication parameters, choose Binary from Output Select and select the TSS1 format from the Formula Select drop-down menu. Ensure that the update rate is set at 200Hz. COM 2 is only for output.

POS MV Output– Ethernet Realtime

Ethernet Realtime Output Control

The preferred option for generating output data from the POS MV to the acquisition software is via Ethernet. The screen capture below shows the configuration for generating navigation, attitude, and quality data from Sensor 1 (sonar) to the acquisition package.

Clearly the data is valid for the Sensor 1 location and the Sensor 1 Frame. It is important to realize this as this has an impact on the computation of offset locations in the acquisition software. The heave is valid for the Sensor 1 Frame Origin.

Acquisition Software

In order that the acquisition software is configured correctly it is important to be sure which frame and which position the data exiting the POS MV is valid for.

Patch Test Results

On completion of a patch test a value will have been computed for the heading misalignment between the frame for which heading data is valid and the actual sonar frame itself. In order to compensate for any pitch and roll cross-talk this heading misalignment must be incorporated in the POS MV. How this is incorporated depends on what heading values are entered into the POS prior to the patch test and what heading data was output from the POS during the calibration.

Patch Test Results

Serial Data

With serial data the NMEA INHDT heading output from the POS is always valid for the Vessel frame. In this case a patch test has solved for the heading misalignment between Vessel and Sonar frames. All motion serial data, including TSS1 message, is valid for the position and mounting angles entered for Sensor 1. Generally these will all be set to zero prior to the patch test. In this scenario the patch test heading misalignment must be entered in the “Sensor 1 Frame w.r.t. Ref. Frame” Z field in order to remove any pitch and roll cross-talk. The heading misalignment must also be entered in the acquisition package because the heading output from the POS has not changed – it is still true heading of Vessel frame. The roll and pitch patch test results may be entered either in the X and Y fields of the “Sensor 1 Frame w.r.t. Ref. Frame” respectively, or in the acquisition package but NOT both as the pitch and roll output from the POS will change as a result of entering them. If any alignment values are in present in the “Sensor 1 Frame w.r.t. Ref. Frame” fields prior to the patch test then the heading value must be replaced by the patch test result. Any roll and pitch values will be residual errors and therefore must be summed with whatever is present in the X and Y fields respectively.

Ethernet Data

All Ethernet data will be valid for Sensor 1 by choosing message 102. In order to remove any pitch and roll cross-talk the patch test heading misalignment must be entered in the “Sensor 1 Frame w.r.t. Ref. Frame” Z field. In doing so the heading output of the POS is changed so the heading misalignment must NOT be entered in the acquisition package. The roll and pitch patch test results may be entered either in the X and Y fields of the “Sensor 1 Frame w.r.t. Ref. Frame” respectively, or in the acquisition package but NOT both as the pitch and roll output from the POS will change as a result of entering them. If any alignment values are in present in the “Sensor 1 Frame w.r.t. Ref. Frame” fields prior to the patch test then all patch test results will be residual errors and therefore must be summed with whatever is present in the fields.

POS Statistics

Ensure that a screen capture is taken of the Statistics page for record keeping purposes. This displays information such as software versions, serial numbers and enabled options. Choose View – Statistics…

Interfacing Sensors To Teledyne PDS

Establishing a Connection to the RSP

- Connect the laptop/PC to the Rack-mounted Sonar Processor (RSP) using an Ethernet cable. The cable should be connected to the LAN 1 or LAN 2 port of the RSP, or through a switch.

- Ensure that the IP address of the RSP is on the same range as the address of the acquisition laptop/PC.

For example, if 10.11.10.2 is assigned to the RSP, the laptop/PC should use 10.11.10.3. Both systems should use the same subnet mask.

NOTE

If the RSP is configured to acquire an automatic IP address (assigned by a DHCP server), but fails, the two Ethernet ports will revert to the static IP addresses 10.11.10.1 (LAN 1) and 10.11.10.2 (LAN 2), both with subnet mask 255.255.255.0, after typically up to 1m30s of trying to acquire DHCP contact (see also Appendix F Network Configuration). This does not apply to the built-in INS option, as it always has a fixed IP address (see appendix F.1 System Network Configuration).

- To assign an IP address to the laptop/PC, navigate to Control Center > Network and Internet > Network and Sharing Center.

- Select Change adapter settings in the navigation pane.

- Open the Local Area Connection.

- Select Properties.

- Select Internet Protocol Version 4 (TCP/IPv4).

- Select Properties.

- Select Use the following IP address and configure the IP address and subnet mask accordingly.

- Start the RSP.

- Launch the Sonar UI on the acquisition laptop/PC.

In the toolbar at the top of the user interface, select the IP address of the RSP from the drop-down list.

Interfacing to Teledyne PDS

Configuring the Sonar UI to Output Data

- Configure the sonar’s supporting sensors (Clock, Position, Sound Velocity, etc.) in the I/O Module tab.

- From the Service menu, select the IO Module QC tab to verify that all sensors are interfaced correctly.

- Expand the Input distribution (RRIO) options.

- Enter the IP address of the laptop/PC and define a Base Port Number.

The Base Port Number can be any number, e.g. 2020.

- The Base Port Number automatically determines the port numbers of all other sensors:

| Device ID | Port number | Example |
| --- | --- | --- |
| VRU | Base port number | 2020 |
| Positioning System | Base port number + 2 | 2022 |
| Compass | Base port number + 4 | 2024 |
| Sound Velocity | Base port number + 6 | 2026 |
| External Clock | Base port number + 8 | 2028 |
| Pan and Tilt | Base port number + 10 | 2030 |

Configuring Teledyne PDS

- Open your project in Teledyne PDS Control Center and go to Vessel Configuration > Equipment.

- Add a Multibeam Device and click the I/O Port button.

  - Add a New Network Port and give it a name (e.g. T51-R).

  - Then enter the IP address of the RSP (10.11.10.2 in this example).

- Similarly, add a Positioning sensor and click the I/O Port button.

  - Add a New Network Port and select the RESON Remote IO option.

  - Give it a name (e.g. Position) and select Position as the Sensor type.

Ensure that 2020 is entered as the Base Port Number. The Base Port Number must match the Base Port Number defined in the Sonar UI Raw Data Distribution setting (see appendix H.2.1 above).

- Test both devices to ensure that that valid data is being transferred,

- Repeat Step 3 for any additional supporting sensors, e.g. Attitude and Compass.

NOTE

For further information on Teledyne PDS, please refer to the appropriate manual or http://www.teledynemarine.com/pds.

Reference Documentation

SeaBat T51-R Documentation

In addition to this document, the following SeaBat documentation is available in Adobe Portable Document Format (.pdf) for printing.

| Document Title | Document Number | Document Description |
| --- | --- | --- |
| SeaBat T51-R Operator's Manual | OM22087 | Version 7 or higher |
| SeaBat T51-R Quick Reference Guide | QG22106 | Version 2 or higher |
| SeaBat UI for T-Series Quick Installation Guide | QG15552 | Version 3 or higher |
| SeaBat Updater Quick Installation Guide | QG18505 | Version 3 or higher |
| Sonar UI User Manual | OM14950 | Version 1.6.0 or higher |
| Data Format Definition | DFD20723 | Version 3.17 or higher |
| With integrated INS option |  |  |
| POS MV V5 Installation and Operation Guide | PUBS-MAN-004291 | Revision 13 or higher |
| POS MV V5 Release Notes | PUBS-CINS-004279 | Revision 19 or higher |
| POS MV V5 User Interface Control Document | PUBS-ICD-004089 | Revision 7 or higher |

Each system contains a link in the Start menu to system documentation (as shown in the example in Figure 62.

Figure 62: System Documentation Link (Example Only)

SeaBat T51-R Design Documents

The following design documents and drawings are provided for reference purposes.

Table 26: SeaBat T51-R Design Documents

| Document Title | Document Number |  |
| --- | --- | --- |
| Receiver, EM7222, Outline | 22252 |  |
| Outline dimension 400 to 800kHz projector TC2186 | 22253 |  |
| Mounting bracket standard, T51, Outline | 22256 (p. 1) | (new date) |
| Fairing add-on, T51, Outline | 22256 (p. 2) | (new date) |
| Sonar reference point, T51 Std. bracket | 22256 (p. 3) | (new date) |
| IMU reference point, T51 Std. bracket | 22256 (p. 4) | updated |
| CABLE ASSY, Conf., Power + GB ethernet, Xm, DIL13F, FISCHER SS105A110, With Lubrication and Dummy Plug | PARTD15415 |  |
| Cable Assembly, Transceiver to DF Transmitter | 85002228B011 |  |
| RSP + Dual head, Outline | 17535 |  |
| RSP + Dual head with INS option, Outline | 17534 |  |
| Outline, RSP+ dual head with AP+ INS option | 25568 |  |
| IMU-20 & IMU-30 for T-Series sonar, Outline | 17537 (p. 1) |  |
| IMU-20 & IMU-30 for T-Series sonar, Outline (motion reference point) | 17537 (p. 2) |  |
| Cable assy, Cable TCS-10003 with PMCIL-8- MP connector and Fischer SS105A124, With Pull Grib, Lubrication and Dummy Plug (IMU cable) | PARTD17145 |  |
| GPS Antenna, GNSS (outline drawing) | AT1675-540T (COTS1010898) |  |
| Cable Assy, Antenna cable, TNC to TNC on LMR-240 Cable | PARTD17193 |  |
| Mounting bracket, T51 dual head. Outline | 26030 (p. 1) | new |
| Mounting bracket, T51 dual head. Outline with fairing | 26030 (p. 2) | new |
| Sonar reference point, T51 dual head bracket | 26030 (p. 3) | new |
| IMU reference point, T51 dual head bracket | 26030 (p. 4) | new |
| Cable Assembly, SVP70 to Transceiver | 85002222B011 |  |

Receiver, EM7222 Outline

Figure 63: Receiver, EM7222 Outline

Projector, TC2186 Outline

Figure 64: Projector, TC2186 Outline

T51 Bracket Installation: Receiver EM7222 and Projector TC2186

Figure 65: T51 Bracket Installation: Receiver EM7222 and Projector TC2186

T51 Bracket Installation with Fairing (optional)

Figure 66: T51 Bracket Installation with Fairing (optional)

T51 Bracket Installation: Sonar Reference Point

Figure 67: T51 Bracket Installation: Sonar Reference Point

T51 Bracket Installation with IMU Reference Point

Figure 68: T51 Bracket Installation with IMU Reference Point

Cable Assembly, Rack-mounted Sonar Processor to Receiver

Figure 69: Cable Assembly, Rack-mounted Sonar Processor to Receiver

Cable Assembly, Rack-mounted Sonar Processor to Projector

Figure 70: Cable Assembly, Rack-mounted Sonar Processor to Projector

Rack-mounted Sonar Processor

Figure 71: Rack-mounted Sonar Processor

Rack-mounted Sonar Processor with INS (optional)

Figure 72: Rack-mounted Sonar Processor with INS (optional)

Rack-mounted Sonar Processor with Applanix AP+ INS (optional)

Figure 73: Rack-mounted Sonar Processor with Applanix AP+ INS (optional)

IMU Outline (optional)

Figure 74: IMU Outline (optional)

IMU Motion Reference Points on T-Series Brackets (optional)

Figure 75: IMU Motion Reference Points on T-Series Brackets (optional)

IMU Cable

Figure 76: IMU Cable

Antenna Outline

Note: Phase center is 58mm above the base of the antenna.

Figure 77: Antenna Outline

Antenna Cable

Figure 78: Antenna Cable

T51 Dual Head Bracket Installation (optional)

Figure 79: T51 Dual Head Bracket Installation (optional)

T51 Dual Head Bracket Installation with Fairing (optional)

Figure 80: T51 Dual Head Bracket Installation with Fairing (optional)

T51 Dual Head Bracket Installation: Sonar Reference Point (optional)

Figure 81: T51 Dual Head Bracket Installation: Sonar Reference Point (optional)

T51 Dual Head Bracket Installation with IMU Reference Point (optional)

Figure 82: T51 Dual Head Bracket Installation with IMU Reference Point (optional)

Cable Assembly, Sonar Processor to SVP 70 (optional)

Figure 83: Cable Assembly, Sonar Processor to SVP 70 (optional)

7kCenter Licenses – Copyright Information

Table 27: LAPACK – Copyright Information

Copyright (c) 1992-2013 The University of Tennessee and The University of Tennessee Research Foundation. All rights reserved. Copyright (c) 2000-2013 The University of California Berkeley. All rights reserved. Copyright (c) 2006-2013 The University of Colorado Denver. All rights reserved. $COPYRIGHT$ Additional copyrights may follow $HEADER$ Redistribution and use in source and binary forms, with or without modification, are permitted provided that the following conditions are met: - Redistributions of source code must retain the above copyright notice, this list of conditions and the following disclaimer. - Redistributions in binary form must reproduce the above copyright notice, this list of conditions and the following disclaimer listed in this license in the documentation and/or other materials provided with the distribution. - Neither the name of the copyright holders nor the names of its contributors may be used to endorse or promote products derived from this software without specific prior written permission. The copyright holders provide no reassurances that the source code provided does not infringe any patent, copyright, or any other intellectual property rights of third parties. The copyright holders disclaim any liability to any recipient for claims brought against recipient by any third party for infringement of that parties intellectual property rights. THIS SOFTWARE IS PROVIDED BY THE COPYRIGHT HOLDERS AND CONTRIBUTORS "AS IS" AND ANY EXPRESS OR IMPLIED WARRANTIES, INCLUDING, BUT NOT LIMITED TO, THE IMPLIED WARRANTIES OF MERCHANTABILITY AND FITNESS FOR A PARTICULAR PURPOSE ARE DISCLAIMED. IN NO EVENT SHALL THE COPYRIGHT OWNER OR CONTRIBUTORS BE LIABLE FOR ANY DIRECT, INDIRECT, INCIDENTAL, SPECIAL, EXEMPLARY, OR CONSEQUENTIAL DAMAGES (INCLUDING, BUT NOT LIMITED TO, PROCUREMENT OF SUBSTITUTE GOODS OR SERVICES; LOSS OF USE, DATA, OR PROFITS; OR BUSINESS INTERRUPTION) HOWEVER CAUSED AND ON ANY THEORY OF LIABILITY, WHETHER IN CONTRACT, STRICT LIABILITY, OR TORT (INCLUDING NEGLIGENCE OR OTHERWISE) ARISING IN ANY WAY OUT OF THE USE OF THIS SOFTWARE, EVEN IF ADVISED OF THE POSSIBILITY OF SUCH DAMAGE.

Table 28: WinPCap – Copyright Information

Copyright (c) 1999-2005 NetGroup, Politecnico di Torino (Italy). Copyright (c) 2005-2010 CACE Technologies, Davis (California). All rights reserved. Redistribution and use in source and binary forms, with or without modification, are permitted provided that the following conditions are met: 1. Redistributions of source code must retain the above copyright notice, this list of conditions and the following disclaimer. 2. Redistributions in binary form must reproduce the above copyright notice, this list of conditions and the following disclaimer in the documentation and/or other materials provided with the distribution. 3. Neither the name of the Politecnico di Torino, CACE Technologies nor the names of its contributors may be used to endorse or promote products derived from this software without specific prior written permission. THIS SOFTWARE IS PROVIDED BY THE COPYRIGHT HOLDERS AND CONTRIBUTORS “AS IS” AND ANY EXPRESS OR IMPLIED WARRANTIES, INCLUDING, BUT NOT LIMITED TO, THE IMPLIED WARRANTIES OF MERCHANTABILITY AND FITNESS FOR A PARTICULAR PURPOSE ARE DISCLAIMED. IN NO EVENT SHALL THE COPYRIGHT OWNER OR CONTRIBUTORS BE LIABLE FOR ANY DIRECT, INDIRECT, INCIDENTAL, SPECIAL, EXEMPLARY, OR CONSEQUENTIAL DAMAGES (INCLUDING, BUT NOT LIMITED TO, PROCUREMENT OF SUBSTITUTE GOODS OR SERVICES; LOSS OF USE, DATA, OR PROFITS; OR BUSINESS INTERRUPTION) HOWEVER CAUSED AND ON ANY THEORY OF LIABILITY, WHETHER IN CONTRACT, STRICT LIABILITY, OR TORT (INCLUDING NEGLIGENCE OR OTHERWISE) ARISING IN ANY WAY OUT OF THE USE OF THIS SOFTWARE, EVEN IF ADVISED OF THE POSSIBILITY OF SUCH DAMAGE. This product includes software developed by the University of California, Lawrence Berkeley Laboratory and its contributors. This product includes software developed by the Kungliga Tekniska Högskolan and its contributors. This product includes software developed by Yen Yen Lim and North Dakota State University. Portions Copyright (c) 1990, 1991, 1992, 1993, 1994, 1995, 1996, 1997 The Regents of the University of California. All rights reserved. Redistribution and use in source and binary forms, with or without modification, are permitted provided that the following conditions are met: 1. Redistributions of source code must retain the above copyright notice, this list of conditions and the following disclaimer. 2. Redistributions in binary form must reproduce the above copyright notice, this list of conditions and the following disclaimer in the documentation and/or other materials provided with the distribution. 3. All advertising materials mentioning features or use of this software must display the following acknowledgement: "This product includes software developed by the University of California, Berkeley and its contributors." 4. Neither the name of the University nor the names of its contributors may be used to endorse or promote products derived from this software without specific prior written permission. THIS SOFTWARE IS PROVIDED BY THE INSTITUTE AND CONTRIBUTORS “AS IS” AND ANY EXPRESS OR IMPLIED WARRANTIES, INCLUDING, BUT NOT LIMITED TO, THE IMPLIED WARRANTIES OF MERCHANTABILITY AND FITNESS FOR A PARTICULAR PURPOSE ARE DISCLAIMED. IN NO EVENT SHALL THE REGENTS OR CONTRIBUTORS BE LIABLE FOR ANY DIRECT, INDIRECT, INCIDENTAL, SPECIAL, EXEMPLARY, OR CONSEQUENTIAL DAMAGES (INCLUDING, BUT NOT LIMITED TO, PROCUREMENT OF SUBSTITUTE GOODS OR SERVICES; LOSS OF USE, DATA, OR PROFITS; OR BUSINESS INTERRUPTION) HOWEVER CAUSED AND ON ANY THEORY OF LIABILITY, WHETHER IN CONTRACT, STRICT LIABILITY, OR TORT (INCLUDING NEGLIGENCE OR OTHERWISE) ARISING IN ANY WAY OUT OF THE USE OF THIS SOFTWARE, EVEN IF ADVISED OF THE POSSIBILITY OF SUCH DAMAGE. Portions Copyright (c) 1983 Regents of the University of California. All rights reserved. Redistribution and use in source and binary forms are permitted provided that the above copyright notice and this paragraph are duplicated in all such forms and that any documentation, advertising materials, and other materials related to such distribution and use acknowledge that the software was developed by the University of California, Berkeley. The name of the University may not be used to endorse or promote products derived from this software without specific prior written permission. THIS SOFTWARE IS PROVIDED “AS IS” AND WITHOUT ANY EXPRESS OR IMPLIED WARRANTIES, INCLUDING, WITHOUT LIMITATION, THE IMPLIED WARRANTIES OF MERCHANTIBILITY AND FITNESS FOR A PARTICULAR PURPOSE. Portions Copyright (c) 1995, 1996, 1997 Kungliga Tekniska Högskolan (Royal Institute of Technology, Stockholm, Sweden). All rights reserved. Redistribution and use in source and binary forms, with or without modification, are permitted provided that the following conditions are met: 1. Redistributions of source code must retain the above copyright notice, this list of conditions and the following disclaimer. 2. Redistributions in binary form must reproduce the above copyright notice, this list of conditions and the following disclaimer in the documentation and/or other materials provided with the distribution. 3. All advertising materials mentioning features or use of this software must display the following acknowledgement: "This product includes software developed by the Kungliga Tekniska Högskolan and its contributors." 4. Neither the name of the University nor the names of its contributors may be used to endorse or promote products derived from this software without specific prior written permission. THIS SOFTWARE IS PROVIDED BY THE INSTITUTE AND CONTRIBUTORS “AS IS” AND ANY EXPRESS OR IMPLIED WARRANTIES, INCLUDING, BUT NOT LIMITED TO, THE IMPLIED WARRANTIES OF MERCHANTABILITY AND FITNESS FOR A PARTICULAR PURPOSE ARE DISCLAIMED. IN NO EVENT SHALL THE INSTITUTE OR CONTRIBUTORS BE LIABLE FOR ANY DIRECT, INDIRECT, INCIDENTAL, SPECIAL, EXEMPLARY, OR CONSEQUENTIAL DAMAGES (INCLUDING, BUT NOT LIMITED TO, PROCUREMENT OF SUBSTITUTE GOODS OR SERVICES; LOSS OF USE, DATA, OR PROFITS; OR BUSINESS INTERRUPTION) HOWEVER CAUSED AND ON ANY THEORY OF LIABILITY, WHETHER IN CONTRACT, STRICT LIABILITY, OR TORT (INCLUDING NEGLIGENCE OR OTHERWISE) ARISING IN ANY WAY OUT OF THE USE OF THIS SOFTWARE, EVEN IF ADVISED OF THE POSSIBILITY OF SUCH DAMAGE. Portions Copyright (c) 1997 Yen Yen Lim and North Dakota State University. All rights reserved. Redistribution and use in source and binary forms, with or without modification, are permitted provided that the following conditions are met: 1. Redistributions of source code must retain the above copyright notice, this list of conditions and the following disclaimer. 2. Redistributions in binary form must reproduce the above copyright notice, this list of conditions and the following disclaimer in the documentation and/or other materials provided with the distribution. 3. All advertising materials mentioning features or use of this software must display the following acknowledgement: "This product includes software developed by Yen Yen Lim and North Dakota State University" 4. The name of the author may not be used to endorse or promote products derived from this software without specific prior written permission. THIS SOFTWARE IS PROVIDED BY THE AUTHOR “AS IS” AND ANY EXPRESS OR IMPLIED WARRANTIES, INCLUDING, BUT NOT LIMITED TO, THE IMPLIED WARRANTIES OF MERCHANTABILITY AND FITNESS FOR A PARTICULAR PURPOSE ARE DISCLAIMED. IN NO EVENT SHALL THE AUTHOR BE LIABLE FOR ANY DIRECT, INDIRECT, INCIDENTAL, SPECIAL, EXEMPLARY, OR CONSEQUENTIAL DAMAGES (INCLUDING, BUT NOT LIMITED TO, PROCUREMENT OF SUBSTITUTE GOODS OR SERVICES; LOSS OF USE, DATA, OR PROFITS; OR BUSINESS INTERRUPTION) HOWEVER CAUSED AND ON ANY THEORY OF LIABILITY, WHETHER IN CONTRACT, STRICT LIABILITY, OR TORT (INCLUDING NEGLIGENCE OR OTHERWISE) ARISING IN ANY WAY OUT OF THE USE OF THIS SOFTWARE, EVEN IF ADVISED OF THE POSSIBILITY OF SUCH DAMAGE. Portions Copyright (c) 1993 by Digital Equipment Corporation. Permission to use, copy, modify, and distribute this software for any purpose with or without fee is hereby granted, provided that the above copyright notice and this permission notice appear in all copies, and that the name of Digital Equipment Corporation not be used in advertising or publicity pertaining to distribution of the document or software without specific, written prior permission. THE SOFTWARE IS PROVIDED “AS IS” AND DIGITAL EQUIPMENT CORP. DISCLAIMS ALL WARRANTIES WITH REGARD TO THIS SOFTWARE, INCLUDING ALL IMPLIED WARRANTIES OF MERCHANTABILITY AND FITNESS. IN NO EVENT SHALL DIGITAL EQUIPMENT CORPORATION BE LIABLE FOR ANY SPECIAL, DIRECT, INDIRECT, OR CONSEQUENTIAL DAMAGES OR ANY DAMAGES WHATSOEVER RESULTING FROM LOSS OF USE, DATA OR PROFITS, WHETHER IN AN ACTION OF CONTRACT, NEGLIGENCE OR OTHER TORTIOUS ACTION, ARISING OUT OF OR IN CONNECTION WITH THE USE OR PERFORMANCE OF THIS SOFTWARE. Portions Copyright (C) 1995, 1996, 1997, 1998, and 1999 WIDE Project. All rights reserved. Redistribution and use in source and binary forms, with or without modification, are permitted provided that the following conditions are met: 1. Redistributions of source code must retain the above copyright notice, this list of conditions and the following disclaimer. 2. Redistributions in binary form must reproduce the above copyright notice, this list of conditions and the following disclaimer in the documentation and/or other materials provided with the distribution. 3. Neither the name of the project nor the names of its contributors may be used to endorse or promote products derived from this software without specific prior written permission. THIS SOFTWARE IS PROVIDED BY THE PROJECT AND CONTRIBUTORS “AS IS” AND ANY EXPRESS OR IMPLIED WARRANTIES, INCLUDING, BUT NOT LIMITED TO, THE IMPLIED WARRANTIES OF MERCHANTABILITY AND FITNESS FOR A PARTICULAR PURPOSE ARE DISCLAIMED. IN NO EVENT SHALL THE PROJECT OR CONTRIBUTORS BE LIABLE FOR ANY DIRECT, INDIRECT, INCIDENTAL, SPECIAL, EXEMPLARY, OR CONSEQUENTIAL DAMAGES (INCLUDING, BUT NOT LIMITED TO, PROCUREMENT OF SUBSTITUTE GOODS OR SERVICES; LOSS OF USE, DATA, OR PROFITS; OR BUSINESS INTERRUPTION) HOWEVER CAUSED AND ON ANY THEORY OF LIABILITY, WHETHER IN CONTRACT, STRICT LIABILITY, OR TORT (INCLUDING NEGLIGENCE OR OTHERWISE) ARISING IN ANY WAY OUT OF THE USE OF THIS SOFTWARE, EVEN IF ADVISED OF THE POSSIBILITY OF SUCH DAMAGE. Portions Copyright (c) 1996 Juniper Networks, Inc. All rights reserved. Redistribution and use in source and binary forms, with or without modification, are permitted provided that: (1) source code distributions retain the above copyright notice and this paragraph in its entirety, (2) distributions including binary code include the above copyright notice and this paragraph in its entirety in the documentation or other materials provided with the distribution. The name of Juniper Networks may not be used to endorse or promote products derived from this software without specific prior written permission. THIS SOFTWARE IS PROVIDED “AS IS” AND WITHOUT ANY EXPRESS OR IMPLIED WARRANTIES, INCLUDING, WITHOUT LIMITATION, THE IMPLIED WARRANTIES OF MERCHANTABILITY AND FITNESS FOR A PARTICULAR PURPOSE. Portions Copyright (c) 2001 Daniel Hartmeier All rights reserved. Redistribution and use in source and binary forms, with or without modification, are permitted provided that the following conditions are met: - Redistributions of source code must retain the above copyright notice, this list of conditions and the following disclaimer. - Redistributions in binary form must reproduce the above copyright notice, this list of conditions and the following disclaimer in the documentation and/or other materials provided with the distribution. THIS SOFTWARE IS PROVIDED BY THE COPYRIGHT HOLDERS AND CONTRIBUTOR “AS IS” AND ANY EXPRESS OR IMPLIED WARRANTIES, INCLUDING, BUT NOT LIMITED TO, THE IMPLIED WARRANTIES OF MERCHANTABILITY AND FITNESS FOR A PARTICULAR PURPOSE ARE DISCLAIMED. IN NO EVENT SHALL THE COPYRIGHT HOLDERS OR CONTRIBUTORS BE LIABLE FOR ANY DIRECT, INDIRECT, INCIDENTAL, SPECIAL, EXEMPLARY, OR CONSEQUENTIAL DAMAGES (INCLUDING, BUT NOT LIMITED TO, PROCUREMENT OF SUBSTITUTE GOODS OR SERVICES; LOSS OF USE, DATA, OR PROFITS; OR BUSINESS INTERRUPTION) HOWEVER CAUSED AND ON ANY THEORY OF LIABILITY, WHETHER IN CONTRACT, STRICT LIABILITY, OR TORT (INCLUDING NEGLIGENCE OR OTHERWISE) ARISING IN ANY WAY OUT OF THE USE OF THIS SOFTWARE, EVEN IF ADVISED OF THE POSSIBILITY OF SUCH DAMAGE. Portions Copyright 1989 by Carnegie Mellon. Permission to use, copy, modify, and distribute this program for any purpose and without fee is hereby granted, provided that this copyright and permission notice appear on all copies and supporting documentation, the name of Carnegie Mellon not be used in advertising or publicity pertaining to distribution of the program without specific prior permission, and notice be given in supporting documentation that copying and distribution is by permission of Carnegie Mellon and Stanford University. Carnegie Mellon makes no representations about the suitability of this software for any purpose. It is provided “as is” without express or implied warranty.

SeaBat T51-R System Options

The following options are available with the SeaBat T51-R system. If you require more detailed information about any of these options, please contact Teledyne RESON Sales for assistance.

| Extended Warranty Contract Fairings Integrated INS SeaBat Cables | Service Level Agreement Sound Velocity Probe System Integration and Training Teledyne PDS Data Acquisition Software |
| --- | --- |

Extended Warranty Contract

All SeaBat T51-R systems are provided with a standard three-year warranty (see Appendix K Standard Warranty Information). This option provides the user with the possibility of purchasing additional annual support.

Fairings

Fairings for the wet-end brackets are available for purchase.

Front fairings help reduce drag and keep a laminar water flow around the transducer arrays. Teledyne RESON generally recommends the use of fairings, especially on moon pool and over-the-side pole installations where the fairing will reduce the drag on the sonar head by up to 35%.

Integrated Inertial Navigation System

The Rack-mounted Sonar Processor can also be configured with a fully integrated Inertial Navigation System for accurate sensor time tagging and motion stabilization. The INS uses GPS for position and time, and IMU for measuring roll, heave, and acceleration. The INS generates true heave output and is fully compliant with Applanix POSPac MMS.

| Component | INS Type-20 | INS Type-30 |  |
| --- | --- | --- | --- |
| Internal INS |  |  |  |
| External IMU Type-20 |  |  |  |
| External IMU Type-30 |  |  |  |
| IMU wet-end cable |  |  |  |
| 2* GPS antennas |  |  |  |
| 2* antenna cables |  |  |  |

SeaBat Cables

Additional cables in various lengths are available for purchase.

Service Level Agreement

A Service Level Agreement (SLA) runs for annual terms of 12 months at a time, and covers a wide range of services as listed here:

| Annual system health check Backup system (availability depending on system) Factory refurbishment (preventive maintenance procedure) | Product and software defect repair services Software and firmware updates Spare part allowance Telephone and email support |
| --- | --- |

Sound Velocity Probe

To ensure proper beam steering, the sonar requires the input of the local sound velocity in the area of the transmitter and receiver arrays. The sound velocity probe (SVP) is used to continuously report this value to the sonar processor unit.

The unit described below may be purchased from Teledyne RESON as an option, or a SVP of similar specifications may be supplied by the customer.

Table 29: SVP Technical Specifications

| Specification | Value |
| --- | --- |
| Type | SVP 70 |
| Sound velocity | Range: 1350-1800m/s Resolution: 0.01m/s Accuracy: ±0.05m/s (0-50m) Accuracy: ±0.25m/s (0-6000m) |
| Sampling rate | 20Hz and lower, programmable |
| Depth | 6000m |
| Baud rate | 2400-115200 |
| Supply | 9-55VDC |
| Power | 2W |
| Housing | Titanium |
| Dimensions | Diameter: 44mm (max) Length: 165mm (excl. connector) End cap height: 69mm (max) |
| Weight | 1.0kg (excl. cable) |
| Temperature | -20° to +55°C |

For further information, please refer to the appropriate manual.

System Integration and Training

Onsite installation, integration with a bathymetric data acquisition system, calibration, and training of operating personnel are available from Teledyne RESON.

Teledyne PDS Data Acquisition Software

Teledyne PDS software provides a wealth of displays and data processing options. The SeaBat T-series rack-mounted sonar processor is available with the software preinstalled. The software can also be configured to run on a separate computer with a TCP/IP or UDP interface. Teledyne PDS also handles all commonly interfaced sensors.

For further information, please refer to the appropriate manual or http://www.teledynemarine.com/pds.

Standard Warranty Information

Limited Warranty

Teledyne RESON warrants that our SeaBatTM T-Series systems shall be free from defects in materials and workmanship for a period of thirty-six (36) months from the date of Teledyne RESON’s original shipment, and that all our other systems and/or auxiliary items shall be free from defects in materials and workmanship for a period of twelve (12) months from the date of Teledyne RESON’s original shipment. These warranty periods apply, unless Teledyne RESON has specified a longer standard warranty period for a particular product and/or offered an extended warranty.

During the warranty period, Teledyne RESON will, at its sole option, either repair, replace, or issue a credit for the original price of the defective system. Such repair, replacement, or credit shall be the sole remedy for a defective system. For full information about the warranty provisions, please refer to the section “Limited Warranty” in our General Terms and Conditions of Sale.

Teledyne RESON systems must be serviced by Teledyne RESON or one of Teledyne RESON’s authorized service providers. Any return of nonconforming or defective systems is subject to Teledyne RESON’s current return authorization process and procedures (see appendix L.2 Equipment Service Procedure).

Warranty Exclusions

The warranty on Teledyne RESON systems does not apply to defects arising from:

Improper installation, operation, or maintenance.

Improper handling, storage, or transportation.

Unauthorized modifications, alterations, or repairs.

Failure to comply with Teledyne RESON’s safety precautions.

Accidental damage.

Normal wear and tear.

Warranty Disclaimer

No warranties of merchantability or fitness for a particular purpose is intended or given.

Servicing During Warranty Period

If your system should encounter technical issues during the warranty period, please contact Teledyne RESON Support (see appendix L.1 Teledyne RESON Support) to protect your warranty rights.

Support and Service

Teledyne RESON Support

If you experience difficulties with your Teledyne RESON system, please contact Teledyne RESON Support for further instructions:

| EUROPE Tel: +45 20 999 088 e-mail: Marine-supportDK@Teledyne.com (Denmark) Marine-supportGE@Teledyne.com (Germany) Marine-supportNL@Teledyne.com (The Netherlands) Marine-supportUK@Teledyne.com (United Kingdom) |  |
| --- | --- |
| UNITED STATES Tel: +1 805 233 3900 e-mail: Marine-supportUS@Teledyne.com |  |
| Teledyne PDS support Tel: +31 10 245 15 00 e-mail: Marine-supportNL@Teledyne.com |  |

Equipment Service Procedure

Policy for Servicing Equipment

NOTE

Serviced items will be returned with the latest released firmware/software.

Teledyne RESON continuously releases new firmware to mitigate any stability, functionality, or other issue, and to improve system performance.

We ensure that firmware/software releases are backwards compatible with older hardware.

We can’t always support new hardware on older firmware/software releases; thus we can’t guarantee that a repaired component is compatible with other than the latest released firmware/software.

Returning Equipment for Service

NOTE

Please do not return any equipment without prior authorization from Teledyne RESON, as evidenced by a Return Authorization (RMA) Number.

Before returning any equipment for service, you must follow the Teledyne RESON equipment return authorization procedure as outlined here:

- Contact Teledyne RESON Support (see appendix L.1 above) to obtain an approved Return Material Authorization (RMA) number.

- Follow the instructions in the supplied document and pack the equipment in the original shipping containers.

- Ship the equipment, transportation and insurance prepaid, according to the instructions issued by Teledyne RESON.

- Ensure that the RMA number is included on all shipping documents and, most importantly, marked on the shipping container’s address label.

- Include a note which identifies the model or part number, and serial number (if applicable) along with a brief, but thorough, description of the problem.

SeaBat T51-R Operator's Manual

Produced by

Teledyne RESON A/S

Fabriksvangen 13

3550 Slangerup

Denmark

Tel: +45 47 38 00 22

Fax: +45 47 38 00 66

Document Number: OM22087-7B

Part Number: 1016039

Includes FlexMode.

For details on this option, please contact Teledyne PDS at pds-support@teledyne.com

FlexMode is an optional add-on feature (see Table 1).

As highlighted in the report “Backscatter measurements by seafloor-mapping sonars” by the GeoHab Backscatter Working Group.

Multi-detect is an optional add-on feature (see Table 1).

Note, that absorption and attenuation of sound in the water column caused by microbubbles and species such as, e.g., plankton and fish is not included in this consideration.

UDP is also available.

Seabed will not be 100% ensonified when exceeding approx. 140° in equi-distant mode.

This is a depth range within which the system is normally operated, from the minimum depth to a depth value corresponding to the max. swath -50%.

This is a single value corresponding to the depth at which the swath is reduced to 10% of its max. value.

Max. 5ms at 66kHz for T51-R at 400kHz.

Receive center beam width mean values from far-field simulation. All beam widths measured at -3dB, center beam unsteered, with a sound velocity of 1440m/s, 30dB Chebyshev shading applied.

FlexMode is an optional add-on feature (see Table 1). FlexMode requires Max. Beam license.

Pipe detection and tracking is an optional add-on feature (see Table 1).

Multi-detect is an optional add-on feature (see Table 1).

X-Range is an optional add-on feature (see Table 1).

TC2181 projector and EM7218 receiver.

[1] The ping rate is limited to 50p/s (see Table 4).

(a1)-(a3) are I/O module referred COM port numbers.

(a4) is split for hardware PPS input; the PPS COM port inputs to COM port 1 (a1).

NMEA PASHR @115200 8,1,n serial data.

This COM port can generate various data formats and interface serial corrections.

See also Appendix C Integrated INS – Applanix AP+.

See also Appendix G Configuration of Integrated INS (optional).

Not applicable for Applanix AP+ INS.

For Applanix AP+ INS, it is Com 1.

Measured at the BNC terminals.

Voltages beyond maximum limits may cause permanent damage to the input.

Completely unconnected input guaranteed to be low.

LAN 2 does not have connectivity to Applanix AP+ INS.

Connected via LAN 1. For details on INS configuration, see Appendix G Configuration of Integrated INS (optional).

Also available here: https://github.com/Teledyne-Marine/7k

The drawing shows a backplane for dual-head operation.

The drawing shows a backplane for dual-head operation.

The drawing shows a backplane for dual-head operation.

Please refer to Teledyne RESON’s General Terms and Conditions of Sale, available at http://www.teledynemarine.com/reson/ > Terms and Conditions.

Refer also to http://www.teledynemarine.com/support.