## Slide 1
MBARI WISSL Mechanical Housing Design Review
December 19, 2023

## Slide 2: Background / Topics
Requirements /Goals
Mechanical Housing with Connector
Mechanical Analysis
MBARI Feedback

## Slide 3: PROPRIETARY INFORMATION
NOTICE: This presentation contains 3D at Depth Confidential and Proprietary information.
In order to better describe some of the design decisions and options, this presentation contains details about the internal architecture and design decisions for 3D at Depth’s current 3D laser scanning products. Therefore, this information should ONLY be shared with MBARI employees who need to read and evaluate the presentation. 3D at Depth appreciates our customers taking the appropriate safety measures when distributing this information internally.

## Slide 4: Requirements – WiSSL Gen 2
Reference WiSSL Upgrades Wishlist_R1_1.PDF
Highlighted Requirements | Must | Goal | Currently | 6000m Option
Depth Rating | 4000m |  | 4000m | 6000m
Size | < 20” diameter, < 3ft long | 6.25” diameter, < 20” long, 26.6” tip to tip discussed | 6.25” Diameter, 22” Tube Length, 28.1” Tip to tip | 6.375” Diameter, 22” Tube Length, 28.1” Tip to tip
Field of View (FoV) | 80° | 90° (360° possible) | 360° | 360°

## Slide 5: Design Goals
Primary Design Goal
  Single housing capable of full 90° FOV (360° possible) that includes both optical components and electronics. This bottle connects directly to the ROV/AUV.
ROV /AUV Connection: 1x 13 pin connector (DFCR2013M)
4000m and 6000m consideration

## Slide 6: Comparison to WiSSL Gen 1 Optical Sensor
OD: 6.25”
Endcaps + Tube Length: 22.37”
Tip-to-tip Length: 28.10” (28.5”)
Weight of tube + endcaps: 44.4 pounds
Tube material: TI-6AL-4V Grade 5
Output Endcap material: 17-4 PH H10
Connector Endcap material: TI-6AL-4V Grade 5
Add .4 inches OAL for 6000m capability
OD: 6.94”
Endcaps + Tube Length: 16.625”
Tip-to-tip Length: 20.29”
Weight of tube + endcaps: 42.41 pounds
Old Optical Bottle (x2)
New All In One
Old Optical Bottle (x2)
Old Optical Bottle (x2)

## Slide 7: Housing Analysis

## Slide 8
Simulation Parameters: Material
Ti-6Al-4V (Gr 5) [all simulations]
  Yield strength: 128 ksi
  Young’s Modulus: 16.5 Mpsi
  Density: 0.166lb/in3
  Poisson’s Ratio: 0.3
Under Pressure Material Properties:
(used as reference check)
SolidWorks Material Properties:
CONTAINS CONFIDENTIAL AND PROPRIETARY INFORMATION – FOR 3D AT DEPTH INTERNAL USE ONLY

## Slide 9
Simulation Conditions: Depth
Factor of Safety (FOS): 1.15
Initial target depth: 4000m
  Actual pressure: 5,846 psi
  4000m x 1.15 FOS = 4600m
  Initial test pressure: 6,723 psi
Goal target depth: 6000m
  Actual pressure: 8,735 psi
  6000m x 1.15 FOS = 6900m
  Goal test pressure: 10,045 psi
Window tube end fixed
Hydrostatic pressure applied to outer shell walls
CONTAINS CONFIDENTIAL AND PROPRIETARY INFORMATION – FOR 3D AT DEPTH INTERNAL USE ONLY

## Slide 10
Results Summary
Note: Buckling FOS is relative to test pressure, not actual pressure. Test pressures include an additional 15% safety margin.
Parameters | 4000m | 6000m
ID (in) | 5.600 | 5.625
OD (in) | 6.25 | 6.375
Wall (in) | 0.3125 | 0.375
Length (in) | 22 | 22
Test Pressure (psi) | 6,723 | 10,045
End area: | 30.68 in^2 | 31.92 in^2
End force: | 206,255 | 320,627.1 LBF
Max Stress (psi) [SW] | 59,675 | 79,069
Max Displacement (mm) [SW] | 0.699 | 0.699
Buckling Failure at (psi) [UP] | 7,464.9 | 10,875
Maximum Depth Rating (m)[UP] | 5067 | 7346
Yield FOS | 2.14 | 1.54
Buckling FOS [SW] | 1.45 | 1.33
Recommended
for 4000m:
CONTAINS CONFIDENTIAL AND PROPRIETARY INFORMATION – FOR 3D AT DEPTH INTERNAL USE ONLY
Recommended
for 6000m:

## Slide 11
Case 1 (4000m): SolidWorks Simulation Results
Air Weight: 15.984 lbs.
Min stress: 46517 psi
Max stress: 59,675 psi
Maximum deflection: 0.510mm
Result: PASS
Case 1 Static Stress Plot:
Case 1 Displacement Plot:
CONTAINS CONFIDENTIAL AND PROPRIETARY INFORMATION – FOR 3D AT DEPTH INTERNAL USE ONLY
Case 1 Buckling mode 1 Plot:
Results: Simulations show pass

## Slide 12
Case 1: Under Pressure Results
Buckling Failure Pressure: 7,464.9 psi
Equivalent Depth: 5,067.6m [16,626FT]
Air Weight: 22.093lbs (TUBE ONLY)
Water Weight: -2.9073lbs
4000m Result: PASS
6000m Result: FAIL
Test 1 Results:
CONTAINS CONFIDENTIAL AND PROPRIETARY INFORMATION – FOR 3D AT DEPTH INTERNAL USE ONLY
Results: Simulations show
Fail for 4000 (underpressure)

## Slide 13
Case 2 (6000m): SolidWorks Simulation Results
Air Weight: 25.814 lbs
Min stress: 60,368 psi
Max stress: 79,069 psi
Maximum deflection: 0.699mm
Result: PASS
Case 2 Static Stress Plot:
Case 2 Displacement Plot:
CONTAINS CONFIDENTIAL AND PROPRIETARY INFORMATION – FOR 3D AT DEPTH INTERNAL USE ONLY
Case 2 Buckling mode 1 Plot:
Results: Simulations show pass

## Slide 14
Case 2: Under Pressure Results
Buckling Failure Pressure: 10,875 psi
Equivalent Depth: 7345.6m
Air Weight: 25.814lbs
Water Weight: -0.196lbs
4000m Result: PASS
6000m Result: PASS
Test 5 Results:
Note: Under Pressure output is the same for Test 4 and Test 5
CONTAINS CONFIDENTIAL AND PROPRIETARY INFORMATION – FOR 3D AT DEPTH INTERNAL USE ONLY
Results: Simulations and underpressure show pass for 6000m

## Slide 15
Conclusions for housing tube
Pressure vessel passes stress and bucking for 4000m with tube of OD from 6.25in and ID of 5.60
Pressure vessel passes stress and bucking for 6000m with tube of OD from 6.375in and ID of 5.60
Increasing the length of the tube has a negligible effect on maximum stress and displacement
CONTAINS CONFIDENTIAL AND PROPRIETARY INFORMATION – FOR 3D AT DEPTH INTERNAL USE ONLY

## Slide 16
Output Endcap Features (by Product Line)
CONTAINS CONFIDENTIAL AND PROPRIETARY INFORMATION – FOR 3D AT DEPTH INTERNAL USE ONLY
Parameter | MBARI (Next Gen) 4000m | MBARI (Next Gen) 6000m
Part number | M000404 revE | 3DDINT-ME-PA-0640 revA
Front face/tube OD (in) * | 6.375 | 6.375
Front face thickness (in) | .7 | .9
Rear boss OD (in) * | 5.597 | 5.597
Rear boss thickness (in) | 2.8 | 2.8
Total thickness (in) | 3.5 | 3.5
Material | 17-4 PH H1025 | 17-4 PH H1150
Yield strength (ksi) | 165 | 165
Rotating cube | √ | √
Safety light | √ | √
Counterbores | √ | √
Front gasket | √ | √
Fasteners | #10-32 (qty 6) | #10-32 (qty 6)
Depth rating (m) | 4000 | 6000

## Slide 17
Part number: 3DDINT-ME-PA-0640 revA
Front face OD: 6.375in
Front face thickness: 0.700in
Rear boss OD: 5.597in
Rear boss thickness: 2.8 in
Total thickness: 3.5in
Material: 17-4 PH H1025
Yield strength: 165 ksi
Rotating Cube: Yes
Safety Light: Yes
Front Gasket: Yes
Fasteners: #10-32 (qty 6)
Next Gen MBARI Output Endcap Features
CONTAINS CONFIDENTIAL AND PROPRIETARY INFORMATION – FOR 3D AT DEPTH INTERNAL USE ONLY
Safety Light

## Slide 18
MBARI Output Endcap Simulation Parameters
CONTAINS CONFIDENTIAL AND PROPRIETARY INFORMATION – FOR 3D AT DEPTH INTERNAL USE ONLY
Parameter | Case 1 (4000m) | Case 2 (6000m)
Front face/tube OD (in) | 6.375 | 6.375
Front face thickness (in) | 0.700 | .900
Boss 1 OD (in) | 5.597 | 5.597
Boss 1 thickness (in) | 2.800 | 2.800
Total thickness (in) * | 3.500 | 3.500
Distance to 1st o-ring (in) | .375 | .375
Material | 17-4 PH H1025 | 17-4 PH H1025
Yield strength (ksi) | 165 | 165

## Slide 19: Scanner Assembly + Pressure Housing
*(built and tested with non-MBARI funding)

## Slide 20: Window - Pressure Testing Results
Pressure test at 3,000m
  Two 10-minute holds and one 1hour hold
Pressure test at 4,000m
  Two 10-minute holds and one 1hour hold
Pressure test at 10,500psi
  Two 10-minute holds and one 1hour hold
Pressure test with motor running
  No increase in motor torque at 10.5 ksi
  See table
Pressure (PSI) | Motor Current (mA) | Notes
0 | 0.42 | 
1,000 | 0.41 | 
2,000 | 0.42 | 
3,000 | 0.42 | 
4,000 | 0.42 | 
5,000 | 0.41 | 
6,000 | 0.41 | 
7,000 | 0.42 | 
8,000 | 0.42 | 
9,000 | 0.41 | 
10,000 | 0.41 | Initial
10,000 | 0.40 | After 30-minute hold
10,500 | 0.40 | Initial
10,500 | 0.40 | After 30-minute hold
0 | 0.41 | After de-pressure
0 | 0.41 | After removed from vessel
*(built and tested with non-MBARI funding)

## Slide 21
Parameter | Case1 (4000m) | Case2 (6000m)
Front face/tube OD (in) | 6.375 | 6.375
Front face thickness (in) | 0.700 | 0.900
Boss 1 OD (in) | 5.597 | 5.597
Boss 1 thickness (in) | 2.800 | 2.800
Boss 2 OD (in) | N/A | N/A
Boss 2 thickness (in) | N/A | N/A
Total thickness (in) | 3.500 | 3.500
Distance to 1st o-ring (in) | 0.161 | 0.250
Material | 17-4 PH H1150 | 17-4 PH H1150
Yield strength (ksi) | 110 | 110
Test pressure (psi) | 10,045 | 10,045
Air weight (lbs) | 21.8 | 21.8
Max stress (ksi) | 105.4 | 84.6
Yield FOS | 1.04 | 1.30
Max displacement (in) | 0.00300 | 0.00289
Window ledge Δ (in) | 0.00077 | 0.00077
Output Endcap Results Summary
CONTAINS CONFIDENTIAL AND PROPRIETARY INFORMATION – FOR 3D AT DEPTH INTERNAL USE ONLY
*: Specified Wide Swath yield strength was later revised to 110 ksi. Yield FOS indicated on table is taken relative to the 125 ksi yield strength used in Brandon’s legacy simulations.

## Slide 22
Connector Endcap Simulation Failure Criteria
CONTAINS CONFIDENTIAL AND PROPRIETARY INFORMATION – FOR 3D AT DEPTH INTERNAL USE ONLY
Front face OD
Rear boss OD
*
*
Point: Near side displacement
Point: Far side displacement
Local stresses shall not exceed material yield strength
Maximum permissible resultant displacement (URES) across connector o-ring seal (discrepancy between “near side displacement” and “far side displacement” values): ≤0.006in (with ≤0.005in preferred)
  “Near side” and “far side” locations refer to the approximate locations of the face seal o-ring affiliated with the connector

## Slide 23
MBARI Endcap Simulation Conditions
Model: 3DDINT-ME-PA-0656.sldprt (revA)
Design requirements:
  Rear boss OD: 5.597in (per Parker)
  O-rings: 2-253 (per Parker)
  Connector: DFCR2013M (Macartney)
Material: Ti-6Al-4V (Gr 5)
  Yield strength: 128 ksi
  Young’s Modulus: 16.5 Mpsi
  Density: 0.166lb/in3
  Poisson’s Ratio: 0.3
SolidWorks Material Properties:
CONTAINS CONFIDENTIAL AND PROPRIETARY INFORMATION – FOR 3D AT DEPTH INTERNAL USE ONLY

## Slide 24
MBARI Endcap Simulation Conditions Case 2
Depth Factor of Safety (FOS): 1.15
Goal target depth:
  Case 1: 4000m
    4000m x 1.15 FOS = 4600m
    Goal test pressure: 6,723 psi
  Case 2: 6000m
    Actual pressure: 8,735 psi
    6000m x 1.15 FOS = 6900m
    Goal test pressure: 10,045 psi
Bulkhead inner edge is fixed
CONTAINS CONFIDENTIAL AND PROPRIETARY INFORMATION – FOR 3D AT DEPTH INTERNAL USE ONLY
Simulation Loading:

## Slide 25
Parameters
Face OD: 6.375in
Face thickness: 0.7in
Boss OD: 5.597in
Boss thickness: 0.875in
Total thickness: 1.575in
Connector distance: 1.775in
MBARI Endcap: Test 1 4000m
CONTAINS CONFIDENTIAL AND PROPRIETARY INFORMATION – FOR 3D AT DEPTH INTERNAL USE ONLY
Test 1 Stress Results:
Test 1 Displacement Results:
Results
Air weight: 6.981 lbs
Max stress: 94.9 ksi (PASS)
Yield FOS: 1.34
Max displacement: 0.0078in
Displacement across seal: 0.0023in (PASS)

## Slide 26
MBARI Endcap: Test 2 6000m
CONTAINS CONFIDENTIAL AND PROPRIETARY INFORMATION – FOR 3D AT DEPTH INTERNAL USE ONLY
Test 2 Stress Results:
Test 2 Displacement Results:
Max stress shows artifact of fixturing. Stresses below yield for part.
Parameters
Face OD: 6.375in
Face thickness: 0.9in
Boss OD: 5.597in
Boss thickness: 0.875in
Total thickness: 1.775in
Connector distance: 1.875in
Results
Air weight: 7.253 lbs
Max stress: 119.7 ksi (PASS)
Yield FOS: 1.06
Max displacement: 0.0111in
Displacement across seal: 0.0043in (PASS)

## Slide 27
Conclusions
Stress levels are well below the yield stress limit for the 4000m connector end cap.
The 6000m connector end cap is below yield stress limit, but .2 inches thicker.
Face seal displacement is acceptable for connector o-ring.
An endcap OD of 6.375in is viable with the caveat that the fasteners are not centered relative to the shell walls.
CONTAINS CONFIDENTIAL AND PROPRIETARY INFORMATION – FOR 3D AT DEPTH INTERNAL USE ONLY

## Slide 28: MBARI Feedback