## Slide 1: i2MAP Camera and Lens Testing
Mark Chaffey
MBARI
chma@mbari.org
July 13, 2014
© MBARI 2014

## Slide 2: Distance Requirements for the Identification of Representative Organisms in the Camera Field-of View
Chart: 
| Aegina | Aegina citrea | Apolemia | Atolla | Atolla wyvillei | Bargmannia elongata | Beroe | Chuniphyes | Clione | Doliolula | Forskalia | Gigantocypris | Hastigerinella digitata | Kephyes ovata | Lampocteis cruentiventer | Melanostigma pammelas | Mysidae | Nanomia bijuga | Pasiphaeidae | Periphylla | Periphylla periphylla | Poeobius meseres | Pseudothecosomata | Salpa | Salpa fusiformis | Solmissus | Solmundella bitentaculata
Calculated First I.D. Distance (cm) | 183.3 | 220 | 275 | 110 | 165 | 165 | 165 | 137.5 | 110 | 110 | 165 | 110 | 100.8 | 110 | 110 | 165 | 110 | 187 | 220 | 165 | 165 | 180 | 110 | 220 | 165 | 192.5 | 137.5
Front of Camera Dome
Measured First I.D. Distance (cm) (Rob Sherlock)
Species
Required I.D. Distance

## Slide 3: Close Focus Test (~10cm from dome)

## Slide 4: Resolution Testing (~ 0.5 m from dome)

## Slide 5: Far Focus (~2.5m from dome)

## Slide 6: Camera Test Setup

## Slide 7: Camera Test Work Flow
AJA Quad Pro ProRes 10 bit 422 (HQ) Recording
Sony XAVC Recording (F-5 only)
Sony 16 bit Raw Recording (F-5 only)
CAMERAS
Ikegami
HDL 50/HS
Sony
HDCP1,
NEX-FS700U
Capture Still Frame in .tiff Format
Export to ProRes 10 bit 422 HQ
Export to DPX Frames
Open in PhotoShop
Export to ProRes 10 bit 4444
2k
4k
1.9k
1.9k
Pixel Resolutions
1.9k = 1920 x 1080
2k = 2048 x 1080
4k = 4096 x 2160
Export to DPX Frames
Open in PhotoShop
1.9k
CAMERA
Sony F-5
CODEC Comparison Reference

## Slide 8: Camera Detail Resolution Comparison Locations Target Distance at “Sweet Spot” ~ 0.5 Meters From Dome
Sony F-5, SCL-P11X15 Lens, 400 lux, 70 IRE, Target 36 inches
Detail Areas
Center
Edge

## Slide 9: Camera Comparison Sony HDCP1 vs. Ikegami HDL50/HS
The B4 Mount Lens Winner Was Used for The Camera Test
Zeiss Digi Prime T1,9/5mm

## Slide 10: Camera Comparison Sony NEX-FS700U vs. Ikegami HDL50/HS
The PL and B4 mount Lens Winners Were Used for The Camera Test
Sony SCL-P11X15 Lens and Zeiss Digi Prime T1,9/5mm

## Slide 11: Camera Comparison Sony F-5 vs. Ikegami HDL50/HS
The PL and B4 mount Lens Winners Were Used for The Camera Test
Sony SCL-P11X15 Lens and Zeiss Digi Prime T1,9/5mm

## Slide 12: Finalist #1 Sony F-5 and Sony SCL-P11X15 Lens
Sony F-5, SCL-P11X15 Lens, 400 lux, 70 IRE, Target 36 inches File: Test_13X.tiff

## Slide 13: Finalist #2, Ikegami HDL-50/H and Zeiss Digi Prime T1,9/5mm Lens
Ikegami HDL50/HS, 5mm Lens, 140 lux, 80 IRE, Target 39 inches File: Test_77.tiff

## Slide 14: Sony F-5 vs. Ikegami HDL50/HS Center Resolution
Sony F-5, Sony 11mm Lens, 400 lux,
ISO 800, Target 36 inches, 1920x 1080 P60
File: Test_13X.tiff
Ikegami HDL-50/HS, Zeiss 5mm Lens, 140 lux, 80 IRE, Target 38 inches
File: Test_77.tiff

## Slide 15: Sony F-5 vs. Ikegami HDL50/HS Edge Resolution
Ikegami HDL-50/HS, Zeiss 5mm Lens, 140 lux, 80 IRE, Target 38 inches
File: Test_77.tiff
Sony F-5, Sony 11mm Lens, 400 lux,
ISO 800, Target 36 inches, 1920x 1080 P60
File: Test_13X.tiff

## Slide 16: Sony F-5 1920 X 1080 vs. Sony F-5 4096 X 2160 (4K) Center Resolution
Sony F-5, Zeiss 12mm Lens, 400 lux,
IE 800, Target 36 inches, 1920x 1080 P60
(no color correction)
File: Test_54_1920_4444.tiff
Sony F-5, Zeiss 12mm Lens, 800 lux,
IE 2000, Target 34 inches, 4K P60
(no color correction)
File: Test_54_4K4444.tiff

## Slide 17: Sony F-5 1920 X 1080 vs. Sony F-5 4096 X 2160 (4K) Center Resolution - Detail
Sony F-5, Zeiss 12mm Lens, 400 lux, ISO 800, Target 34 inches, 1920x 1080 P60
(no color correction)
File: Test_54_1920_4444.tiff
Sony F-5, Zeiss 12mm Lens, 800 lux, IE 2000, Target 34 inches, 4K P60
(no color correction)
File: Test_54_4K4444.tiff

## Slide 18: Sony F-5 1920 X 1080 vs. Sony F-5 4096 X 2160 (4K) Edge Resolution
Sony F-5, Zeiss 12mm Lens, 400 lux, ISO 800, Target 34 inches, 1920x 1080 P60
(no color correction)
File: Test_54_1920_4444.tiff
Sony F-5, Zeiss 12mm Lens, 800 lux, IE 2000, Target 34 inches, 4K P60
(no color correction)
File: Test_54_4K4444.tiff

## Slide 19: Sony F-5 vs. Ikegami HDL50/HS Color Reproduction (4:2:2 Format)
Sony F-5, Sony 11mm Lens, 400 lux, ISO 800, Target 36 inches, 1920x 1080 P60 (no color correction)
Ikegami HDL-50/HS, Zeiss 5mm Lens, 140 lux, 80 IRE, Target 38 inches
(no color correction)

## Slide 20: Test Showing the Sony F-5 and the Sony SCL-P11X15 Lens Meets the i2MAP Depth of Field Requirements (at F 5.6)
Sony F-5, Sony SCL-P11X15 lens set to 11mm, 560 lux, 75 IRE, Target 18 inches (10cm from dome equiv.), F 5.6, focus distance 31.5 inches
File: Test_14.tiff
Sony F-5, Sony SCL-P11X15 lens set to 11mm, 560 lux, 75 IRE, Target 113 inches (2.5m from dome equiv.), F 5.6, focus distance 31.5 inches
File: Test_45.tiff

## Slide 21: Sony F-5 1920 X 1080 vs. Sony F-5 4096 X 2160 (4K) Resolution Example
Sony F-5, Zeiss 12mm Lens, 500 lux, IE 2000, Target 113 inches, 2K P60
(no color correction)
File: Test_53_2K.tiff
Detail Area

## Slide 22: Sony F-5 1920 X 1080 vs. Sony F-5 4096 X 2160 (4K) Resolution Example 2A
Sony F-5, Zeiss 12mm Lens, 500 lux, IE 2000, Target 113 inches, 2K P60 (Lightened 10%) File: Test_53_2K.tiff

## Slide 23: Sony F-5 1920 X 1080 vs. Sony F-5 4096 X 2160 (4K) Resolution Example 2B
Sony F-5, Zeiss 12mm Lens, 500 lux, IE 2000, Target 113 inches,
4K P60 (Lightened 20%) File: Test_53_4K444.tiff

## Slide 24: Sony F-5 1920 X 1080 vs. Sony F-5 4096 X 2160 (4K) Resolution Example 1
Sony F-5, Zeiss 12mm Lens, 400 lux,
IE 800, 1920x 1080 P60
(+20% brightness correction)
File: Test_53_2K4444.tiff
Sony F-5, Zeiss 12mm Lens, 800 lux,
IE 2000, 4K P60
(+20% brightness correction)
File: Test_53_4K4444.tiff

## Slide 25: Temporal Resolution Testing
Rotating Test Target Reproduced Object Angular Velocities Simulating AUV Speeds of 0.55 and 1.0 meters per second

## Slide 26: Translation of Vehicle Motion to Focal-Plane
x_3p – focal plane to particle
x_3t – focal plane to disk
Eqn. 3 => 4 if they’re equal.
2/3” Camera Distance
Super-35 Camera Distance
(adjusted for 0.5 meters from dome front)
1.0 m/sec. = 11 RPM
0.55 m/sec. = 6 RPM

## Slide 27: Temporal Resolution, Can We Get Good Image Quality at 1 Meter Per Second? Enter The Wheel of Death
Ikegami HDL50/HS, Zeiss 5mm Lens, 100 lux, 1920x 1080i 30
Disk 6 RPM (~0.5 m/sec), File: Test_73_White.tiff (DeInterlaced)
Detail Test Area

## Slide 28: Comparison of Current (0.5 m/sec.) and Planned (1.0 m/sec.) Temporal Resolution at the 0.5 Meter from Dome “Sweet Spot”.
Ikegami HDL50/HS, Zeiss 5mm Lens,
110 lux, 1920x 1080i 30
Disk 6 RPM (~0.5 m/sec)
File: Test_73_White.tiff (DeInterlaced)
Ikegami HDL50/HS, Zeiss 5mm Lens,
170 lux, 1920x 1080 P60
Disk 11 RPM (~1 m/sec)
File: : Test_76_White.tiff

## Slide 29: The Effect of Camera Shutter at The Planned (1.0 m/sec.) AUV Velocity tested at the 0.5 Meter from Dome “Sweet Spot”.
Sony HDCP1, Zeiss 5mm Lens,
800 lux, 1920x 1080 P60
Disk 11 RPM (~1 m/sec), Shutter 1/250
File: Test_68_White.tiff
Sony HDCP1, Zeiss 5mm Lens,
120 lux, 1920x 1080 P60
Disk 11 RPM (~1 m/sec), No Shutter
File: Test_66_White.tiff

## Slide 30: Camera Testing Conclusions
Camera
  The Ikegmai 2/3” HDL-50HS camera is significantly better in resolution than the Sony HDCP1.
  The F-5 camera is about 2 stops slower when depth of field control is taken into account.
  The NEX-FS700U is about as fast as the 2/3” cameras and has good resolution but the color reproduction is poor.
  Shutter of 1/240 is needed to freeze the motion of 1 meter per second transect.
  Ikegami has black adjustment issues on power-up boot. Suggest putting 100% filter in one position in filter wheel – would require camera control.
  The Ikegami HDL-50/HS would completely meet our present requirements.
  The Sony F-5 (or F-55) would allow possible future migration:
    To 4K imaging + $12k & new recorder.
    Possibly to higher frame rates.
    Still need to test the F-5 with the Sony lens at 4K to make sure it will hold up.
Video Recorder
  The Apple ProRes 422(HQ) CODEC is visually lossless.
We should have adequate illumination.
  Need ~ 3200 lux maximum
  4 lights measure at ~ 3500 – 5000 lux at 0.5 meters from dome.

## Slide 31: Which is the best camera for us?
Ikegami HDL-50/HS 2/3” box camera
  Box camera has a great picture but is limited to 1920 X 1080 60 frames per second.
  Box camera is the smallest and would be the easiest to integrate – except for auto BB on boot-up.
Sony F-5 super 35 format camera.
  F-5 Cine camera has 4K sensor and offers a migration path to higher resolution formats, 2K and 4K.
    Current recording limitation is 2048 X 1080 60 frames per second.
    AJA has recording hardware in development that supports 4K at 60 frames per second that may be available in the future.
    Sony F-5 can be upgraded to F-55 with 4K output for $12,000.
Cost is very similar – Go with the F-5 (F-55).

## Slide 32: Estimated Camera Cost Comparison
Budget is ~ $57,592

## Slide 33: Extra Slides

## Slide 34: Comparison of the current ROV standard of 1920x 1080i 29.97 at simulated AUV Speed of 0.5 m/sec. and 1920x 1080 P59.4 (Progressively Scanned)
Ikegami HDL45, Fujinon 5.5x10 Lens,
110 lux, 1920x 1080i 29.97
Disk 6 RPM (~0.5 m/sec)
File: Test_103_White.tiff (not DeInterlaced)
Ikegami HDL50/HS, Zeiss 5mm Lens,
170 lux, 1920x 1080 P60
Disk 6 RPM (~0.5 m/sec)
File: : Test_74_White.tiff

## Slide 35: Comparison of the current ROV standard of 1920x 1080i 29.97 at simulated 0.5 m/sec. and the same image de-interlaced using Photoshop
Ikegami HDL45, Fujinon 5.5x10 Lens,
110 lux, 1920x 1080i 29.97
Disk 6 RPM (~0.5 m/sec)
File: Test_103_White.tiff (not DeInterlaced)
Ikegami HDL45, Fujinon 5.5x10 Lens,
110 lux, 1920x 1080i 29.97
Disk 6 RPM (~0.5 m/sec)
File: Test_103_White.tiff (DeInterlaced)

## Slide 36: Comparison of the current ROV standard of 1920x 1080i 29.97 at simulated 0.5 m/sec. De-Interlaced and 1920x 1080 P59.4 (Progressively Scanned)
Ikegami HDL45, Fujinon 5.5x10 Lens,
110 lux, 1920x 1080i 29.97
Disk 6 RPM (~0.5 m/sec)
File: Test_103_White.tiff (DeInterlaced)
Ikegami HDL50/HS, Zeiss 5mm Lens,
170 lux, 1920x 1080 P60
Disk 6 RPM (~0.5 m/sec)
File: : Test_74_White.tiff