## Slide 1: i2MAP Camera Selection (working draft)

## Slide 2: Distance Requirements for the Identification of Representative Organisms in the Camera Field-of View
Chart: 
| Aegina | Aegina citrea | Apolemia | Atolla | Atolla wyvillei | Bargmannia elongata | Beroe | Chuniphyes | Clione | Doliolula | Forskalia | Gigantocypris | Hastigerinella digitata | Kephyes ovata | Lampocteis cruentiventer | Melanostigma pammelas | Mysidae | Nanomia bijuga | Pasiphaeidae | Periphylla | Periphylla periphylla | Poeobius meseres | Pseudothecosomata | Salpa | Salpa fusiformis | Solmissus | Solmundella bitentaculata
Calculated First I.D. Distance (cm) | 183.3 | 220 | 275 | 110 | 165 | 165 | 165 | 137.5 | 110 | 110 | 165 | 110 | 100.8 | 110 | 110 | 165 | 110 | 187 | 220 | 165 | 165 | 180 | 110 | 220 | 165 | 192.5 | 137.5
Front of Camera Dome
Measured First I.D. Distance (cm) (Rob Sherlock)
Species
Required I.D. Distance

## Slide 3: 2/3” 3-chip Sensor Cameras
Use f 1.4 optical prism and three arrays – need special lens design.
Use B4 lens mount.
Lenses are compatible across same type of camera.
Not a large selection of prime lenses.
Known for good sensitivity and resolution.
Hard to find faster frame rates.
Small array has large depth of field (DOF).

## Slide 4: Sony HDCP1
Max frame rate: 1920 x 1080 at 60P
Internal recorder: None
Recorder interface: 1 x 14-bit HD-SDI 4:2:2
Lens: Sony 2/3-type bayonet mount (B4).
Control interface : LAN based Sony system (CCU compatible).
Budget ~ $76k

## Slide 5: 2/3” Prism Camera with Zeiss Digi Prime T1,9/5 mm @ f 2.5 Calculated Depth of Field Focused at 0.9 meter
Chart: 
| Aegina | Aegina citrea | Apolemia | Atolla | Atolla wyvillei | Bargmannia elongata | Beroe | Chuniphyes | Clione | Doliolula | Forskalia | Gigantocypris | Hastigerinella digitata | Kephyes ovata | Lampocteis cruentiventer | Melanostigma pammelas | Mysidae | Nanomia bijuga | Pasiphaeidae | Periphylla | Periphylla periphylla | Poeobius meseres | Pseudothecosomata | Salpa | Salpa fusiformis | Solmissus | Solmundella bitentaculata
Calculated First I.D. Distance (cm) | 183.3 | 220 | 275 | 110 | 165 | 165 | 165 | 137.5 | 110 | 110 | 165 | 110 | 100.8 | 110 | 110 | 165 | 110 | 187 | 220 | 165 | 165 | 180 | 110 | 220 | 165 | 192.5 | 137.5
Imager to dome distance 39cm
Measurement point | Imager | Dome ~
Imager to dome distance (meters) | 0.39 | 
Focus Distance d (meters) | 0.9 | 
Focal length f (mm) | 5 | 
Aperture N (f number) | 2.5 | 
Circle of confusion c (mm) | 0.009 | 
Hyperfocal distance H (meters) | 1.116 | 0.726
Front Focus (meters) | 0.497 | 0.107
Rear Focus (meters) | 4.627 | 4.237
DOF (Meters) | 4.130 | 
Relative exposure to current lens : 0.7
Notes: For the prime lens on a 2/3” prism camera the distance from the imager to the end of the dome is much less – the lens is shorter. This makes it more difficult to focus close to the dome. To compensate we need to move the close focus point closer and stop down the iris to 3.2 to increase the depth of field to get a similar viewing area.

## Slide 6: Zeiss Digi Prime T1,9/5 mm on 2/3” Prism camera
Volume = 2.26 cubic meters

## Slide 7
Notes: For camera temporal resolution testing purposes, we used the “sweet spot” at 0.5 meters from the dome.

## Slide 8: Depth of Field Test Requirement (all lens and camera combinations)

## Slide 9: Camera Temporal Resolution Test Target
Viewing Area 16:9 Aspect Ratio
Test Target ~ .66 meters from Camera Lens (approx. 0.5 m from dome)
Rotation Speed to Simulate 55cm and 100cm per second transect speeds
25%
50%
75%

## Slide 10: Temporal Resolution Test – 2/3” imager
Test point 25%
Test point 50%
Test point 75%
Maximum diagonal
Entrance Pupil
Estimated Target Height
Estimated Target Width
Estimated Image Plane
Test Conditions:
Target rotates at a speed equivalent to the angular motion of particles in the actual field transect.
The two rotational speeds are tested equivalent to advance speeds of 55cm and 100cm per second.
The target distance from the lens entrance pupil is 66.3 cm (89 cm from imager plane).
The test points are a percent of total diagonal field of view.
Lens: Zeiss Digi Prime T1,9/5 mm
Imager: 2/3" 3CCD 8.8 x 6.6 mm (B4 Mount)

## Slide 11: Temporal Resolution Test – super 35 imager
Test point 25%
Test point 50%
Test point 75%
Maximum diagonal
Entrance Pupil
Estimated Target Height
Estimated Target Width
Estimated Image Plane
Test Conditions:
Target rotates at a speed equivalent to the angular motion of particles in the actual field transect.
The two rotational speeds are tested equivalent to advance speeds of 55cm and 100cm per second.
The target distance from the lens entrance pupil is 66.3 cm (89 cm from imager plane).
The test points are a percent of total diagonal field of view.
Lens: Arri/Zeiss Ultra Prime UP 12 mm/T2
Imager: DIN Super 35 24 x 12.7 mm Bayer Array (PL mount)

## Slide 12: Sony HDCP1 Dimensions

## Slide 13: Sony HDCP1 Housing Dia. Mockup

## Slide 14: Ikegami HDL-50/HS
Max frame rate: 1920 x 1080 at 60P
Internal recorder: None
Recorder interface: 2 x HD-SDI 4:2:2 (dual link)
Lens: B4 2/3-type bayonet mount.
Control interface : RM-50, RM-51, RM-11, OCP-200 & MCP-200
Field of view 87 degrees with Zeiss Digi Prime T1,9/5 mm lens.
Very good picture (3CCD).
Good form factor for housing.
No extra peripherals.
Mechanical shutter.
High quality lens available.
Good DOF
+ Side
- Side
Max frame rate: 1920 x 720 at 60P
Same $ as Sony P1.
Ikegami does not make there own imagers.
Poor sensitivity.

## Slide 15: Large sensor cameras
Known as “cine type”
Single array with Bayer pattern has about 70% of the resolution as 3-chip array.
Single array creates film-like optical path, can use almost any large format lens, i.e. PL mount, Canon EF mount, Nikon, etc.
Cameras often have internal storage and the best compression engines (CODECS).
Large sensor reduces the depth of field for a given focal length.

## Slide 16: Sony PMW-F5
Max frame rate: 1920 x 1080 at 60P with XAVC codec, 240 FPS 2K RAW (full res. 60P max.)
Internal recorder: XAVC codec
Recorder interface: 3G-SDI with 2K and 4K capability or HDMI 1.4
Lens: Sony FZ-mount with PL adapter.
Control interface : LANC ??
Field of view 80 degrees with 15mm Zeiss CP.2 15mm/T2.9 E-mount prime lens.

## Slide 17: DIN Super 35 Camera with Arri Ultra Prime UP 10 mm/T2.1 @ f 4.5 Calculated Depth of Field Focused at 0.85 meter
Chart: 
| Aegina | Aegina citrea | Apolemia | Atolla | Atolla wyvillei | Bargmannia elongata | Beroe | Chuniphyes | Clione | Doliolula | Forskalia | Gigantocypris | Hastigerinella digitata | Kephyes ovata | Lampocteis cruentiventer | Melanostigma pammelas | Mysidae | Nanomia bijuga | Pasiphaeidae | Periphylla | Periphylla periphylla | Poeobius meseres | Pseudothecosomata | Salpa | Salpa fusiformis | Solmissus | Solmundella bitentaculata
Calculated First I.D. Distance (cm) | 183.3 | 220 | 275 | 110 | 165 | 165 | 165 | 137.5 | 110 | 110 | 165 | 110 | 100.8 | 110 | 110 | 165 | 110 | 187 | 220 | 165 | 165 | 180 | 110 | 220 | 165 | 192.5 | 137.5
Imager to dome distance 37cm
Measurement point | Imager | Dome ~
Imager to dome distance (meters) | 0.37 | 
Focus Distance d (meters) | 0.85 | 
Focal length f (mm) | 10 | 
Aperture N (f number) | 4.5 | 
Circle of confusion c (mm) | 0.02 | 
Hyperfocal distance H (meters) | 1.121 | 0.751
Front Focus (meters) | 0.481 | 0.111
Rear Focus (meters) | 3.484 | 3.114
DOF (Meters) | 3.002 | 
Relative exposure to current lens : 0.9

## Slide 18: DIN Super 35 Camera with SONY SCL-P11X15 @ f 5.6 Calculated Depth of Field Focused at 0.8 meter
Chart: 
| Aegina | Aegina citrea | Apolemia | Atolla | Atolla wyvillei | Bargmannia elongata | Beroe | Chuniphyes | Clione | Doliolula | Forskalia | Gigantocypris | Hastigerinella digitata | Kephyes ovata | Lampocteis cruentiventer | Melanostigma pammelas | Mysidae | Nanomia bijuga | Pasiphaeidae | Periphylla | Periphylla periphylla | Poeobius meseres | Pseudothecosomata | Salpa | Salpa fusiformis | Solmissus | Solmundella bitentaculata
Calculated First I.D. Distance (cm) | 183.3 | 220 | 275 | 110 | 165 | 165 | 165 | 137.5 | 110 | 110 | 165 | 110 | 100.8 | 110 | 110 | 165 | 110 | 187 | 220 | 165 | 165 | 180 | 110 | 220 | 165 | 192.5 | 137.5
Imager to dome distance 34cm
Measurement point | Imager | Dome ~
Imager to dome distance (meters) | 0.34 | 
Focus Distance d (meters) | 0.8 | 
Focal length f (mm) | 11 | 
Aperture N (f number) | 5.6 | 
Circle of confusion c (mm) | 0.02 | 
Hyperfocal distance H (meters) | 1.091 | 0.751
Front Focus (meters) | 0.460 | 0.120
Rear Focus (meters) | 2.966 | 2.626
DOF (Meters) | 2.507 | 
Relative exposure to current lens : 0.7

## Slide 19: DIN Super 35 Camera with Arri/Zeiss Ultra Prime UP 12 mm/T2 @ f 6.7 Calculated Depth of Field Focused at 0.8 meter
Chart: 
| Aegina | Aegina citrea | Apolemia | Atolla | Atolla wyvillei | Bargmannia elongata | Beroe | Chuniphyes | Clione | Doliolula | Forskalia | Gigantocypris | Hastigerinella digitata | Kephyes ovata | Lampocteis cruentiventer | Melanostigma pammelas | Mysidae | Nanomia bijuga | Pasiphaeidae | Periphylla | Periphylla periphylla | Poeobius meseres | Pseudothecosomata | Salpa | Salpa fusiformis | Solmissus | Solmundella bitentaculata
Calculated First I.D. Distance (cm) | 183.3 | 220 | 275 | 110 | 165 | 165 | 165 | 137.5 | 110 | 110 | 165 | 110 | 100.8 | 110 | 110 | 165 | 110 | 187 | 220 | 165 | 165 | 180 | 110 | 220 | 165 | 192.5 | 137.5
Imager to dome distance 36cm
Relative exposure to current lens : 0.57
Measurement point | Imager | Dome ~
Imager to dome distance (meters) | 0.36 | 
Focus Distance d (meters) | 0.8 | 
Focal length f (mm) | 12 | 
Aperture N (f number) | 6.7 | 
Circle of confusion c (mm) | 0.02 | 
Hyperfocal distance H (meters) | 1.087 | 0.727
Front Focus (meters) | 0.459 | 0.099
Rear Focus (meters) | 2.999 | 2.639
DOF (Meters) | 2.540 | 

## Slide 20: Sony NEX-FS700UK
Max frame rate: 1920 x 1080 at 60P with AVCHD codec, 240 FPS 2K RAW (full res. 60P max.)
Internal recorder: AVCHD codec
Recorder interface: 3G-SDI with 2K and 4K capability or HDMI 1.4
Lens: Sony E-mount hot shoe.
Control interface : LANC
Field of view 80 degrees with 15mm Zeiss CP.2 15mm/T2.9 E-mount prime lens.
High frame rate.
High sensitivity.
Can change lens.
Low cost ($7700).
Low cost lens ($5700)
+ Side
- Side
Unknown picture quality in our application.
Not needed features, LCD screen, etc.

## Slide 21: Arri Ultra Prime UP 10 mm/T2.1 on DIN Super 35 camera
Volume = 3.40 cubic meters

## Slide 22: Sony PMW-F5 Dimensions

## Slide 23: Sony PMW-F5 Housing Dia. Mockup

## Slide 24: Rob’s distance to resolve data
Aegina | 3.3 | 183.3
Aegina citrea | 4.0 | 220.0
Apolemia | 5.0 | 275.0
Atolla | 2.0 | 110.0
Atolla wyvillei | 3.0 | 165.0
Bargmannia elongata | 3.0 | 165.0
Beroe | 3.0 | 165.0
Chuniphyes | 2.5 | 137.5
Clione | 2.0 | 110.0
Doliolula | 2.0 | 110.0
Forskalia | 3.0 | 165.0
Gigantocypris | 2.0 | 110.0
Hastigerinella digitata | 1.8 | 100.8
Kephyes ovata | 2.0 | 110.0
Lampocteis cruentiventer | 2.0 | 110.0
Melanostigma pammelas | 3.0 | 165.0
Mysidae | 2.0 | 110.0
Nanomia bijuga | 3.4 | 187.0
Pasiphaeidae | 4.0 | 220.0
Periphylla | 3.0 | 165.0
Periphylla periphylla | 3.0 | 165.0
Poeobius meseres | 3.3 | 180.0
Pseudothecosomata | 2.0 | 110.0
Salpa | 4.0 | 220.0
Salpa fusiformis | 3.0 | 165.0
Solmissus | 3.5 | 192.5
Solmundella bitentaculata | 2.5 | 137.5
Grand mean: | 3.0 | 165.0
Time Calc. Distance (cm)
Time Calc. Distance (cm)

## Slide 25: 2/3” Prism Camera with Fujinon HA 10X5.2 @ f 2.2 Calculated Depth of Field Focused at 1.0 meter
Chart: 
| Aegina | Aegina citrea | Apolemia | Atolla | Atolla wyvillei | Bargmannia elongata | Beroe | Chuniphyes | Clione | Doliolula | Forskalia | Gigantocypris | Hastigerinella digitata | Kephyes ovata | Lampocteis cruentiventer | Melanostigma pammelas | Mysidae | Nanomia bijuga | Pasiphaeidae | Periphylla | Periphylla periphylla | Poeobius meseres | Pseudothecosomata | Salpa | Salpa fusiformis | Solmissus | Solmundella bitentaculata
Calculated First I.D. Distance (cm) | 183.3 | 220 | 275 | 110 | 165 | 165 | 165 | 137.5 | 110 | 110 | 165 | 110 | 100.8 | 110 | 110 | 165 | 110 | 187 | 220 | 165 | 165 | 180 | 110 | 220 | 165 | 192.5 | 137.5
Measurement point | Imager | Dome ~
Imager to dome distance (meters) | 0.46 | 
Focus Distance d (meters) | 1 | 
Focal length f (mm) | 5.2 | 
Aperture N (f number) | 2.2 | 
Circle of confusion c (mm) | 0.009 | 
Hyperfocal distance H (meters) | 1.371 | 0.911
Front Focus (meters) | 0.577 | 0.117
Rear Focus (meters) | 3.682 | 3.222
DOF (Meters) | 3.105 | 
Imager to dome distance 46cm
Notes: Current Ikegami setup. This is a 2/3” imager with a beam splitting prism between the imager and the zoom lens. The lens is in the full wide position (5.2mm focal length)
Focus distance calculation is from the imager. I adjusted to also be from the dome end to compare with field data.
This is with the iris wide open.
Note that the far focal distance is about 3.2 meters – just past the graph.

## Slide 26: Sony PMW-F55
Max frame rate: 1920 x 1080 at 60P with AVCHD codec, 240 FPS 2K RAW (full res. 60P max.)
Internal recorder: XAVC codec
Recorder interface: 3G-SDI with 2K and 4K capability or HDMI 1.4
Lens: Sony FZ-mount with PL adapter.
Control interface : LANC ??
Field of view 80 degrees with 15mm Zeiss CP.2 15mm/T2.9 PL-mount prime lens.
Very high frame rate.
High sensitivity.
Can change lens.
Low cost lens ($5700)
Better 4k performance.
+ Side
- Side
Unknown picture quality in our application.
Not needed features, LCD screen, etc.
High cost ($29k).

## Slide 27: Blackmagic Cinema Camera EF
Max frame rate: 1920 x 1080 at 30P
Internal recorder: Apple ProRes. All compressed recording in 1920x1080 10-bit YUV with choice of Film or Video Dynamic Range.
Recorder interface: 1 x 10-bit HD-SDI 4:2:2
Lens: EF (Canon) and ZE (Zeiss) mount compatible with electronic iris control.
Control interface: LANC
Lens: EF (Canon) and ZE (Zeiss) mount with electronic iris control (or micro 4/3)
Very good picture.
Low cost. ($2k)
+ Side
- Side
Max frame rate: 1920 x 1080 at 30P
Not the best form factor for housing.
Has unnecessary LCD display interface.

## Slide 28: Sony NEXFS100UK
Max frame rate: 1920 x 1080 at 60P
Internal recorder: AVCHD codec
Recorder interface: HD-SDI, 8-bit HDMI 1.4 4:2:2
Lens: Sony E-mount hot shoe.
Control interface : LANC
Field of view 80 degrees with 15mm Zeiss CP.2 15mm/T2.9 E-mount prime lens.
High frame rate.
High sensitivity.
Low cost ($5k).
Can change lens.
Low cost lens ($5700)
+ Side
- Side
Unknown picture quality in our application.
Not needed features, LCD screen, etc.

## Slide 29: Canon EOS C500 4K
Max frame rate: 1920 x 1080 at 60P with MPG 4 codec, 120 FPS 2K RAW (full res. 60P max.)
Internal recorder: AVCHD codec
Recorder interface: 3G-SDI with 2K and 4K capability
Lens: EF or PL mount.
Control interface : LANC Remote control through WFT-E6A accessory is also possible
Field of view 85+? degrees with 14mm Canon CN-E14mm T3.1 L F EF-mount prime lens.
High frame rate.
High sensitivity.
Can change lens.
Low cost lens ($5500)
Better 4k performance.
+ Side
- Side
Unknown picture quality in our application.
Not needed features, viewfinder, etc.
Have to disassemble viewfinder for housing.
High cost ($20k).

## Slide 30: Canon EOS-1D C
Max frame rate: 1920 x 1080 at 60P
Internal recorder: MOV, Motion JPEG, MPEG-4 AVC (H.264)
Recorder interface: Does not appear to be available except 1080 60i or 1080 24p.
Lens: EF mount.
Control interface : LANC Remote control through WFT-E6A accessory is also possible
Field of view 85+? degrees with 14mm Canon CN-E14mm T3.1 L F EF-mount prime lens.
High resolution.
Can change lens.
Low cost lens ($5500)
+ Side
- Side
Unknown picture quality in our application.
Not needed features, viewfinder, etc.

## Slide 31: Cathx Ocean M-12 Wide angle HD
Max frame rate: 1920 x 1080 at 30P, 1920 x 720 at 60p
Internal recorder: 256GB or 512GB(?)
Recorder interface: SDI
Internal recorder: H.264 MPEG4
Lens: Included zoom
Control interface : 485 serial and Ethernet.
Field of view 88 degrees with 11.5-46mm f/2.8 zoom lens.
COTs system
Compact housing
+ Side
- Side
Unknown picture quality in our application.
High cost, $58k (may not be too much with housing).
Not great at 60P frame rate.

## Slide 32: Ikegami HDL-50/HS
Max frame rate: 1920 x 1080 at 60P
Internal recorder: None
Recorder interface: 2 x HD-SDI 4:2:2 (dual link)
Lens: B4 2/3-type bayonet mount.
Control interface : RM-50, RM-51, RM-11, OCP-200 & MCP-200
Field of view 87 degrees with Fujinon HAeF5 5mm prime cine lens.
Very good picture (3CCD).
Good form factor for housing.
Good sensitivity.
No extra peripherals.
Mechanical shutter.
High quality lens available.
Good DOF
+ Side
- Side
Max frame rate: 1920 x 720 at 60P
Expensive ($?k).
Lens expensive ($35k).

## Slide 33: Cerberus LMP HD1200 (single Bayer array)
Max frame rate: 1920 x 1080 at 60P
Internal recorder: No
Recorder interface: 1080p Dual Link SDI or HDSDI RAW Data in 12/10 bit Cineform DFX.
Lens: C, B4 mount.
Control interface : 232C serial and USB.
Field of view ?? degrees with ? prime lens.
Small.
Can change lens.
Low cost lens.
+ Side
- Side
Unknown picture quality in our application.
Single 2/3” sensor.