Shallow Water FOCE GUI

Controls and Variables Document, v0.1

Chad Kecy

Monterey Bay Aquarium Research Institute

Dec 3, 2014

Table of Contents

Overview	5

Central Node Signal Descriptions	5

Digital Outputs	5

Digital Inputs	6

Analog Inputs	7

Main Electronics Housing Signal Descriptions	9

Serial Outputs	9

Serial Inputs	9

Gateway Node Signal Descriptions	9

Digital Outputs	9

Digital Inputs	11

Analog Inputs	11

Sensor Node Signal Descriptions	13

Digital Outputs	13

Digital Inputs	14

Analog Inputs	14

Valve Assembly	15

Digital Outputs	15

Digital Inputs	15

Scientific Instruments	15

SeaBird SBE18S pH Probe	15

SeaBird SBE43 CTD	15

Satlantic PAR Sensor	15

Omega xxx Flowmeter	16

Appendices	16

No table of figures entries found.

No table of figures entries found.

No table of figures entries found.

# Overview

# Central Node Signal Descriptions

## Digital Outputs

FAN_01_CONT

This is an active high TTL level digital signal which turns on a cooling fan for use with Vicor module #1. The Vicor module converts the incoming +375Vdc into the +48Vdc bus for swFOCE chamber #1.

FAN_02_CONT

This is an active high TTL level digital signal which turns on a cooling fan for use with Vicor module #2. The Vicor module converts the incoming +375Vdc into the +48Vdc bus for swFOCE chamber #2.

VICOR_01_CONT

This is an active low TTL level digital signal which turns on the Vicor module #1.

VICOR_02_CONT

This is an active low TTL level digital signal which turns on the Vicor module #2.

NODE_01_CONT

This is an active high TTL level digital signal which activates a DPDT PRD Series relay, turning on swFOCE chamber #1.

NODE_02_CONT

This is an active high TTL level digital signal which activates a DPDT PRD Series relay, turning on swFOCE chamber #2.

48V_01_GF

This is an active high TTL level digital signal which is used to test the swFOCE chamber #1's 48V bus for seawater shorts.

48V_01_RET_GF

This is an active high TTL level digital signal which is used to test the swFOCE chambers #1's 48V return line for seawater shorts.

48V_02_GF

This is an active high TTL level digital signal which is used to test the swFOCE chamber #2's 48V bus for seawater shorts.

48V_02_RET_GF

This is an active high TTL level digital signal which is used to test the swFOCE chambers #2's 48V return line for seawater shorts.

TEST_GF

This is an active high TTL level digital signal which is used to test the ground fault circuitry.

TEST_RET_GF

This is an active high TTL level digital signal which is used to test the ground fault circuitry.

XPORT_INT

This is an active high TTL level digital signal which is used to reset the Lantronix Matchport Ethernet-to-serial server.

## Digital Inputs

FAN_01_STATE

This is a TTL level signal which indicates the state of cooling fan #1. High indicates the fan is on, low is off.

FAN_02_STATE

This is a TTL level signal which indicates the state of cooling fan #2. High indicates the fan is on, low is off.

VICOR_01_STATE

This is a TTL level signal which indicates the state of the Vicor module #1. High indicates the module is on, low is off.

VICOR_02_STATE

This is a TTL level signal which indicates the state of the Vicor module #2. High indicates the module is on, low is off.

NODE_01_STATE

This is a TTL level signal which indicates the state of the swFOCE chamber #1. High indicates the chamber is on, low is off.

NODE_02_STATE

This is a TTL level signal which indicates the state of the swFOCE chamber #2. High indicates the chamber is on, low is off.

48V_READY

This is a TTL level signal which indicates both delay circuits have fired and that it is safe to energize the Vicor modules. High indicates it safe, low it is still in the delay period.

375V_READY

This is a TTL level signal that indicates the 375Vdc is present. High indicates the presence of 375Vdc, low it is not.

H2O_SENSE_01

This is a TTL level signal that indicates the presence of water in the housing. High indicates water detected, low is normal operation.

H2O_SENSE_02

This is a TTL level signal that indicates the presence of water in the housing. High indicates water detected, low is normal operation.

## Analog Inputs

375V_VOLT

This is an analog signal that represents the incoming +375Vdc.  The measurement range is 0 to 5A and the ADC range is 0 to +3.0Vdc. The defining equation is Y (volts) = 141.8x (volt/volt) - 0.96.

375V_CURR

This is an analog signal that represents the current consumption of the +375V line. The measurement range is 0 to 425Vdc and the ADC range is 0 to +3.0Vdc. The defining equation is Y (amps) = 1.67x (amp/volt)

VICOR_TEMP_01

This is an analog signal that represents the temperature of the Vicor 375V:48V Maxi module #1 which feeds the swFOCE Experimental chamber. The measurement range is 0 to 90 degrees Celsius and the ADC range is 0 to +3.0Vdc. The defining equation is Y (degC) = 33.33x (degC/volt).

VICOR_TEMP_02

This is an analog signal that represents the temperature of the Vicor 375V:48V Maxi module #2 which feeds the swFOCE Control chamber. The measurement range is 0 to 90 degrees Celsius and the ADC range is 0 to +3.0Vdc. The defining equation is Y (degC) = 33.33x (degC/volt).

VICOR_TEMP_03

This is an analog signal that represents the temperature of the Vicor 375V:24V Micro module. The measurement range is 0 to 90 degrees Celsius and the ADC range is 0 to +3.0Vdc. The defining equation is Y (degC) = 33.33x (degC/volt).

HEATSINK_TEMP

This is an analog signal that represents the temperature of the heatsink of the Central Node. The measurement range is 0 to 90 degrees Celsius and the ADC range is 0 +3.0Vdc. The defining equation is Y (degC) = 33.33x (degC/volt).

48V_01_CURR

This is an analog signal that represents the current consumption of the swFOCE Experimental chamber.  The measurement range is 0 to 12.5 amps and the ADC range is 0 to +3.0Vdc. The defining equation is Y (amps) = 4.17x (amp/volt).

48V_01_VOLT

This is an analog signal that represents the voltage level of the swFOCE Experimental chamber bus voltage. The measurement range is 0 to 60 volts and the ADC range is 0 to +3.0Vdc. The defining equation is Y (amps) = 20.0x (volt/volt).

48V_02_CURR

This is an analog signal that represents the current consumption of the swFOCE Control chamber.  The measurement range is 0 to 12.5 amps and the ADC range is 0 to +3.0Vdc.

The defining equation is Y (amps) = 4.17x (amp/volt).

48V_02_VOLT

This is an analog signal that represents the current consumption of the swFOCE Control chamber.  The measurement range is 0 to 12.5 amps and the ADC range is 0 to +3.0Vdc.

The defining equation is Y (amps) = 4.17x (amp/volt).

24V_CURR

This is an analog signal that represents the current consumption of the +24Vdc used on the MV board. The measurement range is 0 to 3.125 amps and the ADC range is 0 to +3.0Vdc. The defining equation is Y (volts) = 1.11x (amp/volt).

24V_VOLT

This is an analog signal that represents the voltage level of the +24Vdc used on the MV board. The measurement range is 0 to 30Vdc and the ADC range is 0 to +3.0Vdc. The defining equation is Y (volts) = 10x (volt/volt).

3.3V_CURR

This is an analog signal that represents the current consumption of the +3.3Vdc used on the MV board. The measurement range is 0 to 4 amps and the ADC range is 0 to +3.0Vdc. The defining equation is Y (volts) = 1.33x (amp/volt).

3.3V_VOLT

This is an analog signal that represents the voltage level of the +3.3Vdc used on the MV board. The measurement range is 0 to 3.6Vdc and the ADC range is 0 to +3.0Vdc. The defining equation is Y (volts) = 1.20x (volt/volt).

AMB_HUM

This is an analog signal that represents the ambient humidity level within the Central Node housing. The measurement range is 0 to 100 percent relative humidity (%RH) and the ADC range is 0 to +5.0Vdc. The defining equation is Y (%RH) = 37.5x (%RH/volt) - 37.7

AMB_TEMP

This is an analog signal that represents the ambient temperature within the Central Node housing. The measurement range is -40 degC to +80 degC and the ADC range is 0 to +5.0Vdc.

The defining equation is Y (degC) = -2.43x^3 + 20x^2 -74.19x + 123.32 (degC/volt)

GF_CURR

This is an analog signal that represents a potential conduction path between the seawater ground and the bus voltages and return lines that feed the swFOCE chambers. This is a bipolar signal centered around 2.5Vdc. Signals lower than 2.5Vdc represent a conduction path to the return line while those between 2.5Vdc and 5.0Vdc represent a conduction path to the positive rails.The measurement range is -1ma to +1ma and the ADC range is 0 to +5Vdc. The defining equation is Y (ma) = 0.4x (ma/volt) -1.

# Main Electronics Housing Signal Descriptions

## Serial Outputs

Motor Control

Motor Speed Command

Gateway Node Control

Light #1 Control

Camera Control

Light #2 Control

Light #3 Control

## Serial Inputs

Motor Speed Feedback

Ambient Temperature

Ambient Humidity

H2O Sense #1

H2O Sense #2

# Gateway Node Signal Descriptions

## Digital Outputs

NODE_1_ON

NODE_1_ON is used to apply power to Sensor Node #1; it does this by activating the reset coil of relay K1. The default state is high and its active level is a negative going pulse 5 msec in duration.

NODE_1_OFF

NODE_1_OFF is used to remove power from Sensor Node #1; it does this by activating the set coil of relay K1. The default state is high and its active level is a negative going pulse 5 msec in duration.

NODE_2_ON

NODE_2_ON is used to apply power to Sensor Node #2; it does this by activating the reset coil of relay K3. The default state is high and its active level is a negative going pulse 5 msec in duration.

NODE_2_OFF

NODE_2_OFF is used to remove power from Sensor Node #2; it does this by activating the set coil of relay K3. The default state is high and its active level is a negative going pulse 5 msec in duration.

NODE_3_ON

NODE_3_ON is used to apply power to Sensor Node #3; it does this by activating the reset coil of relay K5. The default state is high and its active level is a negative going pulse 5 msec in duration.

NODE_3_OFF

NODE_3_OFF is used to remove power from Sensor Node #3; it does this by activating the set coil of relay K5. The default state is high and its active level is a negative going pulse 5 msec in duration.

NODE_4_ON

NODE_4_ON is used to apply power to Sensor Node #4; it does this by activating the reset coil of relay K7. The default state is high and its active level is a negative going pulse 5 msec in duration.

NODE_4_OFF

NODE_4_OFF is used to remove power from Sensor Node #4; it does this by activating the set coil of relay K7. The default state is high and its active level is a negative going pulse 5 msec in duration.

48V_GF

48V_GF is used to test the source 48Vdc for conductivity to the seawater reference; it does this by activating opto-relay K4. The default state is low and its active level is a steady state high.

48V_RET_GF

48V_RET_GF is used to test the source 48Vdc Return for conductivity to the seawater reference; it does this by activating opto-relay K1. The default state is low and its active level is a steady state high.

TEST_GF

TEST_GF is used to test a reference voltage for conductivity to the seawater reference; it does this by activating opto-relays K3 and K6. This signal allows for the calibration and verification of the ground fault detection circuitry. The default state is low and its active level is a steady state high.

TEST_RET_GF

TEST_RET_GF is used to test a reference voltage return for conductivity to the seawater reference; it does this by activating opto-relays K2 and K5. This signal allows for the calibration and verification of the ground fault detection circuitry. The default state is low and its active level is a steady state high.

## Digital Inputs

NODE_1_STATE

NODE_1_STATE indicates the power state of Sensor Node #1. Opto relay K2 drives this signal high when Sensor Node #1 is energized. It is low when Sensor Node #1 is off.

NODE_2_STATE

NODE_2_STATE indicates the power state of Sensor Node #2. Opto relay K4 drives this signal high when Sensor Node #2 is energized. It is low when Sensor Node #2 is off.

NODE_3_STATE

NODE_3_STATE indicates the power state of Sensor Node #3. Opto relay K6 drives this signal high when Sensor Node #3 is energized. It is low when Sensor Node #3 is off.

NODE_4_STATE

NODE_4_STATE indicates the power state of Sensor Node #4. Opto relay K8 drives this signal high when Sensor Node #4 is energized. It is low when Sensor Node #4 is off.

H2O_ALARM

H2O_ALARM indicates the presence of seawater in the Gateway Node subsea housing; it does this by driving an opto-isolator, U13, high if seawater is detected. A high reading indicates water in the housing while a low indicates a dry interior.

ENVR_ALARM

ENVR_ALARM indicates the either the ambient temperature or humidity has gone out of the programmed range. Its default state is high; a low indicates an alarm. The actual values can then be read back over the SPI bus.

## Analog Inputs

NODE_1_CURR

NODE_1_CURR is a voltage which is representative of the current drawn by Sensor Node #1. The maximum current drawn and measured is 2 amps, which corresponds to an output of 1.6Vdc.

NODE_1_CURR (volts) = ILoad (amps) * 0.1 (ohms) * 8

Once this value is read by the ADC on the BeagleBone, the actual current may be found:

ILoad = NODE_1_CURR (volts) / 0.8 (ohms)

NODE_2_CURR

NODE_2_CURR is a voltage which is representative of the current drawn by Sensor Node #2. The maximum current drawn and measured is 2 amps, which corresponds to an output of 1.6Vdc.

NODE_2_CURR (volts) = ILoad (amps) * 0.1 (ohms) * 8

Once this value is read by the ADC on the BeagleBone, the actual current may be found:

ILoad = NODE_2_CURR (volts) / 0.8 (ohms)

NODE_3_CURR

NODE_3_CURR is a voltage which is representative of the current drawn by Sensor Node #3. The maximum current drawn and measured is 2 amps, which corresponds to an output of 1.6Vdc.

NODE_3_CURR (volts) = ILoad (amps) * 0.1 (ohms) * 8

Once this value is read by the ADC on the BeagleBone, the actual current may be found:

ILoad = NODE_3_CURR (volts) / 0.8 (ohms)

NODE_4_CURR

NODE_4_CURR is a voltage which is representative of the current drawn by Sensor Node #4. The maximum current drawn and measured is 2 amps, which corresponds to an output of 1.6Vdc.

NODE_4_CURR (volts) = ILoad (amps) * 0.1 (ohms) * 8

Once this value is read by the ADC on the BeagleBone, the actual current may be found:

ILoad = NODE_4_CURR (volts) / 0.8 (ohms)

GF_CURR

GF_CURR is a voltage which is representative of the electrical current between seawater reference and the ground fault channel under test.

[add more after GF circuit is calibrated and tested – CDK 08/21/2014]

48V_VOLT

48V_VOLT is a scaled voltage representative of the actual 48Vdc source. For 48Vdc, the scaled voltage should be 1.58Vdc.

48V_VOLT = 0.0329 * Vin

Once this value is read by the ADC on the BeagleBone, the actual voltage may be found:

V = 30.4 * 48V_VOLT

48V_CURR

48V_CURR is a voltage which is representative of the current drawn by the entire load of the 48Vdc source. The maximum current drawn and measured is 6 amps, which corresponds to an output of 1.6Vdc.

48V_CURR (volts) = ILoad (amps) * 0.033 (ohms) * 8

Once this value is read by the ADC on the BeagleBone, the actual current may be found:

ILoad = NODE_4_CURR (volts) / 0.264 (ohms)

AMB_TEMP

AMB_HUM

VICOR_TEMP

# Sensor Node Signal Descriptions

## Digital Outputs

INSTR_1_ON

This is a logic high level positive going pulse which turns on the power relay to the instrument on port #1. The minimum pulse width is 10msec and should not exceed 50msec.

INSTR_1_OFF

This is a logic high level positive going pulse which turns off the power relay to the instrument on port #1. The minimum pulse width is 10msec and should not exceed 50msec.

INSTR_2_ON

This is a logic high level positive going pulse which turns on the power relay to the instrument on port #2. The minimum pulse width is 10msec and should not exceed 50msec.

INSTR_2_OFF

This is a logic high level positive going pulse which turns off the power relay to the instrument on port #2. The minimum pulse width is 10msec and should not exceed 50msec.

INSTR_3_ON

This is a logic high level positive going pulse which turns on the power relay to the instrument on port #3. The minimum pulse width is 10msec and should not exceed 50msec.

INSTR_3_OFF

This is a logic high level positive going pulse which turns off the power relay to the instrument on port #3. The minimum pulse width is 10msec and should not exceed 50msec.

INSTR_4_ON

This is a logic low level positive going pulse which turns on the power relay to the instrument on port #4. The minimum pulse width is 10msec and should not exceed 50msec.

INSTR_4_OFF

This is a logic low level positive going pulse which turns off the power relay to the instrument on port #4. The minimum pulse width is 10msec and should not exceed 50msec.

## Digital Inputs

INSTR_1_STATE

This is a logic level feedback signal which corresponds to the state of the relay on port #1. Logic level high indicates the relay is on.

INSTR_2_STATE

This is a logic level feedback signal which corresponds to the state of the relay on port #2. Logic level high indicates the relay is on.

INSTR_3_STATE

This is a logic level feedback signal which corresponds to the state of the relay on port #3. Logic level high indicates the relay is on.

INSTR_4_STATE

This is a logic level feedback signal which corresponds to the state of the relay on port #4. Logic level high indicates the relay is on.

CONFIG_BIT_x (0-7)

These are 8 digital bits which can be read back by the user to verify the hardware configuration of the SN module. CONFIG_BIT_7 is the MSB.

## Analog Inputs

INSTR_CURR

This is a 0-3V analog voltage signal which corresponds to the total current usage of the DC-DC converter. The conversion factor is Y(amps) = 2.0 x INSTR_CURR.

INSTR_VOLT

This is a 0-3V analog voltage signal which corresponds to the output voltage of the DC-DC converter. The conversion factor is Y(volts) = 16.0 x INSTR_VOLT.

INSTR_1_CURR

This is a 0-3V analog voltage signal which corresponds to the current usage of the instrument being monitored on port #1. The conversion factor is Y(amps) = 0.667 x INST_1_CURR.

INSTR_2_CURR

This is a 0-3V analog voltage signal which corresponds to the current usage of the instrument being monitored on port #2. The conversion factor is Y(amps) = 0.667 x INST_2_CURR.

INSTR_3_CURR

This is a 0-3V analog voltage signal which corresponds to the current usage of the instrument being monitored on port #3. The conversion factor is Y(amps) = 0.667 x INST_3_CURR.

INSTR_4_CURR

This is a 0-3V analog voltage signal which corresponds to the current usage of the instrument being monitored on port #4. The conversion factor is Y(amps) = 2.0 x INST_4_CURR.

4_20MA_VOLT

This is a 0-3V analog voltage signal which corresponds to the 4-20ma circuit, if used. The conversion factor is Y(milliamps) = 6.1 x 4_20MA_VOLT + 4.0. The scaling for 4 to 20mA ends at 2.62V. A reading of 2.97 volts (22.1ma) indicates an error in the measurement.

# Valve Assembly

## Digital Outputs

VALVE_1_CONT

VALVE_2_CONT

## Digital Inputs

VALVE_1_FDBK

VALVE_2_FDBK

# Scientific Instruments

## SeaBird SBE18S pH Probe

## SeaBird SBE43 CTD

## Satlantic PAR Sensor

## Omega xxx Flowmeter

# Appendices

Signal name

System name

Signal type

Active level

Description of function

Measurement range

Measurement units

Sample rate

Conversion equation

Display range

Display units

Low alarm

High alarm

SIAM instrument number

GUI grouping

GUI access