Shallow Water FOCE Gateway Node Prototype

Interface Control Document  v0.1

Chad Kecy

04-10-2013

System Level Description

The Gateway Node (GN) serves two primary purposes in the swFOCE experiment: it collects data from all of the instrumentation and acts as the main computer in the control loop. The GN interfaces to a maximum of up to four Sensor Nodes (SN). It provides 48Vdc and communicates via serial comms. The GN has a BeagleBone (BB) computer module onboard and it handles all decision making in swFOCE; it is the central computer in the control loop.

Function Description

BeagleBone

The GN Carrier board interfaces to a BeagleBone computer board. The BB uses a TI AM3358 720MHz ARM Cortex-A8 based microprocessor. It has an Ethernet interface, USB port, microSD card slot, analog inputs, (4) serial ports and (2) 46-pin expansion headers. The GN Carrier board also has a BB suggested EEPROM in the event a board ID is required.

DC-DC Conversion

The GN has a single 250W 48V:48V Vicor Maxi; it is used primarily to electrically isolate the SNs from the rest of the system. The Vicor Maxi is mounted off board and 48V is brought onto the GN Carrier board. The 48V is split among the four SNs and is used to create logic power on the GN Carrier board.

Load Paths

For the prototype board, there are two families of electromechanical relays that are being evaluated for the final design. Therefore, there are two pairs of load paths: one pair (1&2) utilizing an onboard DK family of relays and one pair (3&4) utilizing an off board chassis relay. Each load path can provide up to 6A to each SN. A current sense amplifier (LT1787) is used to measure the current delivered to each SN. A proportional voltage output is sent through a first order low pass filter prior to being sent to an analog in port on the BeagleBone (BB) module.

Load Paths 1&2

The DK series relays from Panasonic PEWA are pc mounted dual coil latching 2 Form C relays; they are internally wired in such a way that high side switching to the coils is required. Individual on and off signals from the BB drive a dual P-channel mosfet for coil activations. Relay state feedback is also provided to the BB.

Load Paths 3&4

The SP series relays from Panasonic PEWA are chassis mounted dual coil latching 2 Form C relays; they can be activated using low side switching. Individual on and off signals from the Beaglebone drive a dual N-channel mosfet for coil activation. Relay state feedback is also provided to the BB.

PIC and Serial Communication

The four available serial ports from BB are being used to communicate with the SNs. Separate one-channel transceivers are used in the event of a failure only affecting one channel.

Analog Signals

The BB has seven analog inputs; the signals need to be scaled to 1.8V max. The resolution of the onboard ADC is 12 bits. The measured analog signals are the four SN output currents, the input 48Vdc, and the measured ground fault current.

Ground Faulting

The GN Carrier board has the ability to check for ground faults on the +48Vdc line and Return that go to the SNs. Digital signals from the BB test the positive and negative rails of the 48Vdc against a seawater reference return. The analog signal GF_CURR will register a fault in either direction. There is also a test channel that places a known intentional short to ground on an isolated voltage source in order to test and calibrate the GF circuitry.

Environmental Sensors

The GN Carrier board will eventually be housed in a PVC housing and placed underwater, therefore standard sensors are included. These include ambient temperature, humidity, and a water presence detection circuit. There is also a thermocouple conditioning circuit for monitoring the temperature of the Vicor Mini module (type K).

Figure One: Possible Gateway Node to Sensor Node configuration on Shallow Water FOCE

| BeagleBone P8 Expansion Signals |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Pin | Proc | BB Name | Signal Name | Pin | Proc | BB Name | Signal Name |
| 1 |  | GND | Gnd | 2 |  | GND | Gnd |
| 3 | R9 | GPIO1_6 |  | 4 | T9 | GPIO1_7 |  |
| 5 | R8 | GPIO1_2 |  | 6 | T8 | GPIO1_3 |  |
| 7 | R7 | TIMER4 |  | 8 | T7 | TIMER7 |  |
| 9 | T6 | TIMER6 |  | 10 | U6 | TIMER6 |  |
| 11 | R12 | GPIO1_13 |  | 12 | T12 | GPIO1_12 |  |
| 13 | T10 | EHRPWM2B |  | 14 | T11 | GPIO0_26 |  |
| 15 | U13 | GPIO1_15 |  | 16 | V13 | GPIO1_14 |  |
| 17 | U12 | GPIO0_27 |  | 18 | V12 | GPIO2_1 |  |
| 19 | U10 | EHRPWM2A |  | 20 | V9 | GPIO1_31 | ENVR_ALARM |
| 21 | U9 | GPIO1_30 | EEPROM_WP | 22 | V8 | GPIO1_5 | H2O_ALARM |
| 23 | U8 | GPIO1_4 | 48V_RET_GF | 24 | V7 | GPIO1_1 | TEST_RET_GF |
| 25 | U7 | GPIO1_0 | 48_GF | 26 | V6 | GPIO1_29 | TEST_GF |
| 27 | U5 | GPIO2_22 | NODE_1_STATE | 28 | V5 | GPIO2_24 | NODE_3_STATE |
| 29 | R5 | GPIO2_23 | NODE_2_STATE | 30 | R6 | GPIO2_25 | NODE_4_STATE |
| 31 | V4 | UART5_CTSN |  | 32 | T5 | UART5_RTSN |  |
| 33 | V3 | UART4_RTSN |  | 34 | U4 | UART3_RTSN |  |
| 35 | V2 | UART4_CTSN |  | 36 | U3 | UART3_CTSN |  |
| 37 | U1 | UART5_TXD | UART_5_TTL_TXD | 38 | U2 | UART5_RXD | UART_5_TTL_RXD |
| 39 | T3 | GPIO2_12 | NODE_4_ON | 40 | T4 | GPIO2_13 | NODE_4_OFF |
| 41 | T1 | GPIO2_10 | NODE_3_ON | 42 | T2 | GPIO2_11 | NODE_3_OFF |
| 43 | R3 | GPIO2_8 | NODE_2_ON | 44 | R4 | GPIO2_9 | NODE_2_OFF |
| 45 | R1 | GPIO2_6 | NODE_1_ON | 46 | R2 | GPIO2_7 | NODE_1_OFF |

| BeagleBone P9 Expansion Signals |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Pin | Proc | BB Name | Signal Name | Pin | Proc | BB Name | Signal Name |
| 1 |  | GND | Gnd | 2 |  | GND | Gnd |
| 3 |  | VDD_3V3EXP | 3.3V_EXP | 4 |  | VDD_3V3EXP | 3.3V_EXP |
| 5 |  | VDD_5V | 5V_BB | 6 |  | VDD_5V | 5V_BB |
| 7 |  | SYS_5V | 5V_SYS | 8 |  | SYS_5V | 5V_SYS |
| 9 |  | PWR_BUT* |  | 10 | A10 | SYS_RESETn |  |
| 11 | T17 | UART4_RXD | UART_4_TTL_RXD | 12 | U18 | GPIO1_28 |  |
| 13 | U17 | UART4_TXD | UART_4_TTL_TXD | 14 | U14 | EHRPWM1A |  |
| 15 | R13 | GPIO1_16 |  | 16 | T14 | EHRPWM1B |  |
| 17 | A16 | I2C1_SCL |  | 18 | B16 | I2C1_SDA |  |
| 19 | D17 | I2C2_SCL | I2C2_SCL | 20 | D18 | I2C2_SDA | I2C2_SDA |
| 21 | B17 | UART2_TXD | UART_2_TTL_TXD | 22 | A17 | UART2_RXD | UART_2_TTL_RXD |
| 23 | V14 | GPIO1_17 |  | 24 | D15 | UART1_TXD | UART_1_TTL_TXD |
| 25 | A14 | GPIO3_21 |  | 26 | D16 | UART1_RXD | UART_1_TTL_RXD |
| 27 | C13 | GPIO3_19 |  | 28 | C12 | SPI1_CS0 | SPI1_CS0 |
| 29 | B13 | SPI1_DO | SPI1_MOSI | 30 | D12 | SPI1_DI | SPI1_MISO |
| 31 | A13 | SPI1_SCLK | SPI1_SCLK | 32 |  | VDD_ADC |  |
| 33 | C8 | AIN4 | GF_CURR | 34 |  | GNDA_ADC |  |
| 35 | A5 | AIN6 |  | 36 | A5 | AIN5 | 48V_MON |
| 37 | B7 | AIN2 | NODE_3_CURR | 38 | A7 | AIN3 | NODE_4_CURR |
| 39 | B6 | AIN0 | NODE_1_CURR | 40 | C7 | AIN1 | NODE_3_CURR |
| 41 | D14 | CLKOUT2 |  | 42 | C18 | SPI1_CS1 | SPI1_CS1 |
| 43 | GND | GND | Gnd | 44 |  | GND | Gnd |
| 45 | GND | GND | Gnd | 46 |  | GND | Gnd |

| Ref # | Color | Signal | Value |
| --- | --- | --- | --- |
| TP1 | Red | Vin_Pos | 48V |
| TP2 | Black | Vin_Neg | Ground |
| TP3 | Black | Vin_Neg | Ground |
| TP4 | Black | Vin_Neg | Ground |
| TP5 | Black | Vin_Neg | Ground |
| TP6 | Yellow | Vref_1.8V | 1.8V |
| TP7 | Orange | 5V_BB | 5V |
| TP8 | Orange | 5V_INT | 5V |
| TP9 | Yellow | 3.3V_INT | 3.3V |
| TP10 | Blue |  | +/- 0.2V |
| TP11 | White |  | +/- 0.2V |
| TP12 | Purple |  |  |
| TP13 | Green |  | +/- 0.2V |
| TP14 | Green |  |  |
| TP15 | Blue |  |  |
| TP16 | White |  |  |
| TP17 | White |  | 2.5V |
| TP18 | Purple |  |  |
| TP19 | Orange | +5V_ISO | 5V |
| TP20 | Yellow | -5V_ISO | -5V |
| TP21 | Brown | GND_ISO | Ground |
| TP22 | Orange | 5V_GF | 5V |
| TP23 | Yellow | -5V | -5V |
| TP24 | Black | Ground | Ground |

| Name | Signal Type | Dir (w/r to BB) | BB Expan Signal | BB Expan Port | Active (Outputs) |
| --- | --- | --- | --- | --- | --- |
| NODE_1_OFF | Digital | Out | GPIO2_7 | P8-46 | Low Pulse |
| NODE_1_ON | Digital | Out | GPIO2_6 | P8-45 | Low Pulse |
| NODE_1_STATE | Digital | In | GPIO2_22 | P8-27 |  |
| NODE_1_CURR | Analog | In | AIN0 | P9-39 |  |
| NODE_2_OFF | Digital | Out | GPIO2_9 | P8-44 | Low Pulse |
| NODE_2_ON | Digital | Out | GPIO2_8 | P8-43 | Low Pulse |
| NODE_2_STATE | Digital | In | GPIO2_23 | P8-29 |  |
| NODE_2_CURR | Analog | In | AIN1 | P9-40 |  |
| NODE_3_OFF | Digital | Out | GPIO2_11 | P8-42 | High Pulse |
| NODE_3_ON | Digital | Out | GPIO2_10 | P8-41 | High Pulse |
| NODE_3_STATE | Digital | In | GPIO2_24 | P8-28 |  |
| NODE_3_CURR | Analog | In | AIN2 | P9-37 |  |
| NODE_4_OFF | Digital | Out | GPIO2_13 | P8-40 | High Pulse |
| NODE_4_ON | Digital | Out | GPIO2_12 | P8-39 | High Pulse |
| NODE_4_STATE | Digital | In | GPIO2_25 | P8-30 |  |
| NODE_4_CURR | Analog | In | AIN3 | P9-38 |  |
| UART_1_TTL_TXD | Digital | Out | UART1_TXD | P9-24 |  |
| UART_1_TTL_RXD | Digital | In | UART1_RXD | P9-26 |  |
| UART_2_TTL_TXD | Digital | Out | UART2_TXD | P9-21 |  |
| UART_2_TTL_RXD | Digital | In | UART2_RXD | P9-22 |  |
| UART_4_TTL_TXD | Digital | Out | UART4_TXD | P9-13 |  |
| UART_4_TTL_RXD | Digital | In | UART4_RXD | P9-11 |  |
| UART_5_TTL_TXD | Digital | Out | UART5_TXD | P8-37 |  |
| UART_5_TTL_RXD | Digital | In | UART5_RXD | P8-38 |  |
| 48V_RET_GF | Digital | Out | GPIO1_4 | P8-23 | High SS |
| 48V_GF | Digital | Out | GPIO1_0 | P8-25 | High SS |
| TEST_RET_GF | Digital | Out | GPIO1_1 | P8-24 | High SS |
| TEST_GF | Digital | Out | GPIO1_29 | P8-26 | High SS |
| GF_CURR | Analog | In | AIN4 | P9-33 |  |
| H2O_ALARM | Digital | In | GPIO1_5 | P8-22 |  |
| 48V_MON | Analog | In | AIN5 | P9-36 |  |
| ENVR_ALARM | Digital | In | GPIO1_31 | P8-20 |  |
| EEPROM_WP | Digital | Out | GPIO1_30 | P8-21 | High SS |

Digital Signals

NODE_1_ON

NODE_1_OFF

NODE_1_STATE

NODE_2_ON

NODE_2_OFF

NODE_2_STATE

NODE_3_ON

NODE_3_OFF

NODE_3_STATE

NODE_4_ON

NODE_4_OFF

NODE_4_STATE

48V_RET_GF

48V_GF

TEST_RET_GF

TEST_GF

H2O_ALARM

ENVR_ALARM

EEPROM_WP

Analog Signals

NODE_1_CURR

NODE_2_CURR

NODE_3_CURR

NODE_4_CURR

GF_CURR

48V_MON