Instrument Simulator Software Requirements/Functionality

02-05-2014

Chad Kecy

- Setup / Initialization

1. DC-DC Converter Information

- Query user for voltage/wattage information. Must have this info prior to any operation.

2. Port #3

- Query user how port #3 will be configured (serial or 4-20ma) (default to serial?)

3. Serial Port Configurations

- Query user on baud rate, parity, stop bits, flow control (default to 9600,n,8,1,none?)

B. Power + Load Monitoring

- This functionality should be running in the background at all times checking the following:

- Each port is not drawing more than  Pmax

- Each port (1-3) is not drawing more than 2 amps

- Port #4 is not drawing more than 6 amps

- The combined output power of all four ports does not exceed Pmax.

C. Operating Modes

1. Instrument Mode (available on ports 1-4)

- Activate the requested port’s electronic dummy load at the programmed level

- Stream stored data (either serial or 4-20ma) (ports 1-3 only)

- Read instrument load voltage and current

- Calculate instrument power

2. Manual Load (available on ports 1-4)

- Accept user data and activate requested port’s electronic dummy load

- Verify user data (current/power) to be within port’s limit

- Verify user data (current/power) to within overall power limitations

3. Serial Stream (available on ports 1-3)

- Requested port will output either a requested string or a stored string

4. 4-20ma Stream (available on port #3)

- Will output either a requested current or a stored series of current values out of port #3.