Shallow Water FOCE Sensor Node Board

Hardware Interface Document, v0.2

Chad Kecy

03-28-2014

System Level Description

The role of the Sensor Node (SN) as part of the swFOCE experiment is to interface a wide range of underwater instruments. The SN provides power and serial communication to three instruments, and also has one high-current power only port. Up to four SNs connect to the Gateway Node (GN) and communicate via RS-232 serial communication. The GN provides the SN with 48V (75W).The SN can provide 12Vdc to 48Vdc with a maximum of 50 watts available per SN. The SN utilizes RS-232 serial communication with a maximum baud rate of 1Mbps. One of the ports can be configured to accept data from 4-20ma instruments. Data from the individual instruments is stored locally at the SN until requested from the GN. The GN handles all decision making; the SN is simply a data aggregator.

Figure One: Possible Sensor Node configuration on Shallow Water FOCE

Function Description

DC-DC Conversion

The SN contains a single DC-DC converter which provides a common DC voltage for the attached instruments. The board has been designed to accept two families of converters which allow the user to determine the output voltage and the maximum output power (for greater efficiency). The input and output filters were designed to accommodate all of the available converters. A jumper selectable trim up/down potentiometer is also included. The current and voltage level of the instrument bus voltage (BUS_VOLT) is monitored and recorded. A digital signal from the PIC can disable the DC-DC converter in the event power to all of the instruments needs to be removed quickly. The logic power is provided by a separate voltage regulator which is supplied directly by the input 48Vdc.

Load Paths

The instrument bus voltage is split into four paths: three standard instrument ports and one high-current power port. The standard instrument ports are limited to 2A of current, regardless of bus voltage. The high-current port can provide up to the maximum of 6A. The current usages of all four ports are monitored individually. Load switching to each port is accomplished via dual coil latching relays. Separate on and off pulses are used to change the relay states. Digital signals indicating the states of the individual relays are sent back to the PIC for independent confirmation. The high-current port has a different family of relay that allows for higher current usage; it can handle the maximum current at 50W for any of the chosen DC-DC converters.

The total power consumption of the four output ports should not exceed the rated power level of the installed DC-DC converter. There is no hardware monitoring of power usage; this will need to be monitored in software.

PIC and Instrument Communication

The SN board contains a PIC24F microcontroller with four serial UARTs. Three are used for communication to the instruments, and the fourth is for communication back to the GN. Separate one-channel transceivers are used in the event of a failure only affecting one channel. The three instrument serial transceivers may be placed into a sleep mode to reduce overhead power consumption. The third port may be configured to accept 4-20ma instruments. In this case, serial port #3 is inactive and the 4-20ma data is read through the ADC. The PIC handles all of the digital input and output signals, the four UARTs, and the spi bus. An onboard boot loader allows the operating program to be loaded remotely.

Configuration Bits

Due to all of the possible hardware configurations of the Sensor Node board, and the fact that the board will be in a potted housing, there are eight bits which may be read back to confirm the hardware configuration. These are set by eight dip switches prior to potting.

Analog Signals

A SPI bus based 16-bit Analog-to-Digital Converter (ADC) is used to measure the various analog signals on board. All of the voltages are scaled to 3.0 volts. The signals measured are the bus voltage and current, the individual instrument currents, and the 4-20ma signal, if used.

Board Population Options

The Sensor Node can be configured in a number of ways to suit the needs of the end user.

DC-DC Converter Selection

There are two families of CUI Inc. DC-DC converters, VYB and VHB, which can be used on the SN board. External filtering and control components have been selected to accommodate any voltage/power combination.

|  |  |  | Power | Input | Output | Current |
| --- | --- | --- | --- | --- | --- | --- |
| Manuf | Series | Manufacturer PN# | (Watts) | (Volts) | (Volts) | (Amps) |
| CUI Inc | VYB10W | VYB10W-Q48-S12 | 10 | 18-72 | 12 | 0.833 |
| CUI Inc | VYB10W | VYB10W-Q48-S15 | 10 | 18-72 | 15 | 0.666 |
| CUI Inc | VYB15W | VYB15W-Q48-S12 | 15 | 18-75 | 12 | 1.25 |
| CUI Inc | VYB15W | VYB15W-Q48-S15 | 15 | 18-75 | 15 | 1 |
| CUI Inc | VYB20W | VYB20W-Q48-S12 | 20 | 18-75 | 12 | 1.667 |
| CUI Inc | VYB20W | VYB20W-Q48-S15 | 20 | 18-75 | 15 | 1.333 |
| CUI Inc | VYB20W | VYB20W-Q48-S24 | 20 | 18-75 | 24 | 0.834 |
| CUI Inc | VHB50W | VHB50W-Q48-S12 | 50 | 18-75 | 12 | 4.16 |
| CUI Inc | VHB50W | VHB50W-Q48-S15 | 50 | 18-75 | 15 | 3.33 |
| CUI Inc | VHB50W | VHB50W-Q48-S24 | 50 | 18-75 | 24 | 2.08 |
| CUI Inc | VHB50W | VHB50W-Q48-S28 | 50 | 18-75 | 28 | 1.78 |
| CUI Inc | VHB50W | VHB50W-Q48-S48 | 50 | 18-75 | 48 | 1.04 |
| CUI Inc | VHB75W | VHB75W-Q48-S12 | 75 | 18-75 | 12 | 6.25 |
| CUI Inc | VHB75W | VHB75W-Q48-S15 | 75 | 18-75 | 15 | 5 |
| CUI Inc | VHB75W | VHB75W-Q48-S24 | 75 | 18-75 | 24 | 3.12 |
| CUI Inc | VHB75W | VHB75W-Q48-S48 | 75 | 18-75 | 48 | 1.56 |

Table One: List of available DC-DC converters

4-20ma Circuit

Instrument port #3 can be configured to accept 4-20ma devices. Resistors R44-R47, and R51 are used to select between serial communication and 4-20ma signals.

|  | R44 | R45 | R46 | R47 | R51 |
| --- | --- | --- | --- | --- | --- |
| Serial | OPEN | 0 OHMS | 0 OHMS | OPEN | 0 OHMS |
| 4-20ma | 0 OHMS | OPEN | OPEN | 0 OHMS | OPEN |

Table Two: Jumper selection for Instrument Port #3.

In addition, the user has the option of not populating certain components if one or the other port #3 options will not be used.

If the user plans on using serial communication only and never 4-20ma, the following components can be considered No Stuff: F1, R44, R47, R48, R49, R50, R52, R53, C58, C59, C60, C62, U24, and U25.

If the user plans on using 4-20ma on port #3 and never serial communication, the following components can be considered No Stuff:  R45, R46, C53, C54, C55, C56, C57, and U23.

Note on R51: Special attention must be made with respect to R51. If using serial communication, it needs to be populated with a zero ohm resistor. This prevents electrical noise from interfering with the ADC (U20). If using 4-20ma, R51 needs to be No Stuff in order to present the 4-20ma signal to the ADC.

Logic Power LED

During the bench top test and debug phase it may be useful to have a power on indication. When R12 is populated with a zero ohm resistor, LED D2 indicates the presence of +3.3Vdc. Removing R12 reduces the power consumption by approximately 12mW.

Port #4 Feedback Resistor

Due to the switching relay on port #4 not having a second set of poles, an opto-isolator is used to read the relay state feedback. In order to keep the forward diode current in the nominal operating range, the current limiting resistor has to match the bus voltage.

| BUS_VOLT | R33 | Watts |
| --- | --- | --- |
| 12 Vdc | 12k | 0.250 W |
| 15 Vdc | 12k | 0.250 W |
| 24 Vdc | 24k | 0.250 W |
| 28Vdc | 24k | 0.250 W |
| 48 Vdc | 47k | 0.250 W |

Table Three: Current limiting resistors for U17

Connector Population

The power and signal connectors for the SN board should be populated during initial testing; this facilitates easier bench testing. If ultimately potting the SN board in a housing, the connectors can be removed and the whips can be directly soldered to the board. However, depending on the user’s final deployment method, the power and signal connectors can remain populated.

The power connectors (J1, J2, J3, J4, and J5) were designed to accommodate 18AWG wire. The communication connectors (J7, J8, J9, and J10) were designed to accommodate 24AWG wire.

In addition, the programming connector, J6, may be removed prior to deployment, if desired.

Board Configuration Feedback

The user configuration can be set into memory by using the 8 position dip switch, SW1, as indicated in table four. [Also include a picture when available].

The 8-bit word can be read back using CONFIG_BIT_0 through CONFIG_BIT_7; these correspond to digital port B, signals 8 through 15. See table five.

| Watts | Voltage | Port #3 | SW1-1 | SW1-2 | SW1-3 | SW1-4 | SW1-5 | SW1-6 | SW1-7 | SW1-8 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 75 | 48 | Serial | OFF | OFF | OFF | OFF | OFF | OFF | ON | ON |
| 75 | 24 | Serial | ON | OFF | OFF | OFF | OFF | OFF | ON | ON |
| 75 | 15 | Serial | OFF | ON | OFF | OFF | OFF | OFF | ON | ON |
| 75 | 12 | Serial | ON | ON | OFF | OFF | OFF | OFF | ON | ON |
| 50 | 48 | Serial | OFF | OFF | ON | OFF | OFF | OFF | ON | ON |
| 50 | 28 | Serial | ON | OFF | ON | OFF | OFF | OFF | ON | ON |
| 50 | 24 | Serial | OFF | ON | ON | OFF | OFF | OFF | ON | ON |
| 50 | 15 | Serial | ON | ON | ON | OFF | OFF | OFF | ON | ON |
| 50 | 12 | Serial | OFF | OFF | OFF | ON | OFF | OFF | ON | ON |
| 20 | 24 | Serial | ON | OFF | OFF | ON | OFF | OFF | ON | ON |
| 20 | 15 | Serial | OFF | ON | OFF | ON | OFF | OFF | ON | ON |
| 20 | 12 | Serial | ON | ON | OFF | ON | OFF | OFF | ON | ON |
| 15 | 15 | Serial | OFF | OFF | ON | ON | OFF | OFF | ON | ON |
| 15 | 12 | Serial | ON | OFF | ON | ON | OFF | OFF | ON | ON |
| 10 | 15 | Serial | OFF | ON | ON | ON | OFF | OFF | ON | ON |
| 10 | 12 | Serial | ON | ON | ON | ON | OFF | OFF | ON | ON |
| 75 | 48 | 4-20ma | OFF | OFF | OFF | OFF | ON | OFF | ON | ON |
| 75 | 24 | 4-20ma | ON | OFF | OFF | OFF | ON | OFF | ON | ON |
| 75 | 15 | 4-20ma | OFF | ON | OFF | OFF | ON | OFF | ON | ON |
| 75 | 12 | 4-20ma | ON | ON | OFF | OFF | ON | OFF | ON | ON |
| 50 | 48 | 4-20ma | OFF | OFF | ON | OFF | ON | OFF | ON | ON |
| 50 | 28 | 4-20ma | ON | OFF | ON | OFF | ON | OFF | ON | ON |
| 50 | 24 | 4-20ma | OFF | ON | ON | OFF | ON | OFF | ON | ON |
| 50 | 15 | 4-20ma | ON | ON | ON | OFF | ON | OFF | ON | ON |
| 50 | 12 | 4-20ma | OFF | OFF | OFF | ON | ON | OFF | ON | ON |
| 20 | 24 | 4-20ma | ON | OFF | OFF | ON | ON | OFF | ON | ON |
| 20 | 15 | 4-20ma | OFF | ON | OFF | ON | ON | OFF | ON | ON |
| 20 | 12 | 4-20ma | ON | ON | OFF | ON | ON | OFF | ON | ON |
| 15 | 15 | 4-20ma | OFF | OFF | ON | ON | ON | OFF | ON | ON |
| 15 | 12 | 4-20ma | ON | OFF | ON | ON | ON | OFF | ON | ON |
| 10 | 15 | 4-20ma | OFF | ON | ON | ON | ON | OFF | ON | ON |
| 10 | 12 | 4-20ma | ON | ON | ON | ON | ON | OFF | ON | ON |

Table Four: Configuration Settings

|  |  |  | B8 | B9 | B10 | B11 | B12 | B13 | B14 | B15 | Dec |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Watts | Voltage | Port #3 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 |  |
| 75 | 48 | Serial | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 63 |
| 75 | 24 | Serial | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 62 |
| 75 | 15 | Serial | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 61 |
| 75 | 12 | Serial | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 60 |
| 50 | 48 | Serial | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 59 |
| 50 | 28 | Serial | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 58 |
| 50 | 24 | Serial | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 57 |
| 50 | 15 | Serial | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 56 |
| 50 | 12 | Serial | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 55 |
| 20 | 24 | Serial | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 54 |
| 20 | 15 | Serial | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 53 |
| 20 | 12 | Serial | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 52 |
| 15 | 15 | Serial | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 51 |
| 15 | 12 | Serial | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 50 |
| 10 | 15 | Serial | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 49 |
| 10 | 12 | Serial | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 48 |
| 75 | 48 | 4-20ma | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 47 |
| 75 | 24 | 4-20ma | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 46 |
| 75 | 15 | 4-20ma | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 45 |
| 75 | 12 | 4-20ma | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 44 |
| 50 | 48 | 4-20ma | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 43 |
| 50 | 28 | 4-20ma | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 42 |
| 50 | 24 | 4-20ma | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 41 |
| 50 | 15 | 4-20ma | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 40 |
| 50 | 12 | 4-20ma | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 39 |
| 20 | 24 | 4-20ma | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 38 |
| 20 | 15 | 4-20ma | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 37 |
| 20 | 12 | 4-20ma | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 36 |
| 15 | 15 | 4-20ma | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 35 |
| 15 | 12 | 4-20ma | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 34 |
| 10 | 15 | 4-20ma | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 33 |
| 10 | 12 | 4-20ma | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 32 |

Table Five: Reading back the configuration settings.

*

*

*

*

*

*.....Space reserved for a picture of the dip switch

*

*

*

*

*

*

*

*

*

Test and Debug Features

Programming and Serial Debugging

During the initial test phase of the board, serial port #4 can be used as a debug port in conjunction with a Hyperterminal window. The PIC is programmed through J6. The signals are compatible with Microchip’s ICD programmers. Jumper JP1 allows the user to manually reset the PIC in the event of a programmer lockup or malfunction.

Heartbeat LED

D19 is a through-hole LED on extended legs. This allows the user to mount the LED in such a way that it is visible through any potting compound. D19 can be used during initial test, final deployment, or both.

ADC Calibration

Channel 7 of the ADC (U20) is unused and by default is tied to ground via a zero ohm resistor. The user has the option of creating a voltage divider via R54 and R43 to calibrate the analog-to-digital hardware and software. If Vmeas is the voltage to be measured by the ADC, then:

R54 = R43 * (3.0V – Vmeas) / Vmeas))

Vmeas = 3.0V * (R43 / (R43 + R54))

For optimal use of the voltage reference, U22, the current draw should be between 100uA and 1mA. For example, R43 = 10k and R54 = 20k results in Vmeas = 1.0V and the current drawn is 100uA.

If no calibration is desired, install a zero ohm resistor for R43 and leave R54 open.

Board Level Signal Descriptions

Analog Signals

INSTR_CURR

This is a 0-3V analog voltage signal which corresponds to the total current usage of the DC-DC converter. The conversion factor is Y(amps) = 2.0 x INSTR_CURR.

INSTR_VOLT

This is a 0-3V analog voltage signal which corresponds to the output voltage of the DC-DC converter. The conversion factor is Y(volts) = 16.0 x INSTR_VOLT.

INSTR_1_CURR

This is a 0-3V analog voltage signal which corresponds to the current usage of the instrument being monitored on port #1. The conversion factor is Y(amps) = 0.667 x INST_1_CURR.

INSTR_2_CURR

This is a 0-3V analog voltage signal which corresponds to the current usage of the instrument being monitored on port #2. The conversion factor is Y(amps) = 0.667 x INST_2_CURR.

INSTR_3_CURR

This is a 0-3V analog voltage signal which corresponds to the current usage of the instrument being monitored on port #3. The conversion factor is Y(amps) = 0.667 x INST_3_CURR.

INSTR_4_CURR

This is a 0-3V analog voltage signal which corresponds to the current usage of the instrument being monitored on port #4. The conversion factor is Y(amps) = 2.0 x INST_4_CURR.

4_20MA_VOLT

This is a 0-3V analog voltage signal which corresponds to the 4-20ma circuit, if used. The conversion factor is Y(milliamps) = 6.1 x 4_20MA_VOLT + 4.0. The scaling for 4 to 20mA ends at 2.62V. A reading of 2.97 volts (22.1ma) indicates an error in the measurement.

Digital Output Signals

ADC_*SHDN

This is an active low input to the ADS8344 ADC. Pulling this line low puts the ADS8344 in a low power shutdown mode.

UART_1_*SHDN

This is a logic low digital signal which puts the serial transceiver for port #1 into a low power shutdown mode.

UART_2_*SHDN

This is a logic low digital signal which puts the serial transceiver for port #2 into a low power shutdown mode.

UART_3_*SHDN

This is a logic low digital signal which puts the serial transceiver for port #3 into a low power shutdown mode.

INSTR_1_ON

This is a logic high level positive going pulse which turns on the power relay to the instrument on port #1. The minimum pulse width is 10msec and should not exceed 50msec.

INSTR_1_OFF

This is a logic high level positive going pulse which turns off the power relay to the instrument on port #1. The minimum pulse width is 10msec and should not exceed 50msec.

INSTR_2_ON

This is a logic high level positive going pulse which turns on the power relay to the instrument on port #2. The minimum pulse width is 10msec and should not exceed 50msec.

INSTR_2_OFF

This is a logic high level positive going pulse which turns off the power relay to the instrument on port #2. The minimum pulse width is 10msec and should not exceed 50msec.

INSTR_3_ON

This is a logic high level positive going pulse which turns on the power relay to the instrument on port #3. The minimum pulse width is 10msec and should not exceed 50msec.

INSTR_3_OFF

This is a logic high level positive going pulse which turns off the power relay to the instrument on port #3. The minimum pulse width is 10msec and should not exceed 50msec.

INSTR_4_ON

This is a logic low level positive going pulse which turns on the power relay to the instrument on port #4. The minimum pulse width is 10msec and should not exceed 50msec.

INSTR_4_OFF

This is a logic low level positive going pulse which turns off the power relay to the instrument on port #4. The minimum pulse width is 10msec and should not exceed 50msec.

DC-DC_CNTL

This is an active low control signal which disables the onboard DC-DC converter. This removes power to all of the attached instruments simultaneously.

Digital Inputs

INSTR_1_STATE

This is a logic level feedback signal which corresponds to the state of the relay on port #1. Logic level high indicates the relay is on.

INSTR_2_STATE

This is a logic level feedback signal which corresponds to the state of the relay on port #2. Logic level high indicates the relay is on.

INSTR_3_STATE

This is a logic level feedback signal which corresponds to the state of the relay on port #3. Logic level high indicates the relay is on.

INSTR_4_STATE

This is a logic level feedback signal which corresponds to the state of the relay on port #4. Logic level high indicates the relay is on.

ADC_BUSY

This is an output of the ADS8344 that indicates the operating status of the ADC. Its operation is different depending on whether using internal or external clocking. It is in a high impedance state when *CS is high.

CONFIG_BIT_x (0-7)

These are 8 digital bits which can be read back by the user to verify the hardware configuration of the SN module. CONFIG_BIT_7 is the MSB. See tables four and five for the exact codes.

SPI Bus Signals

SPI_SCLK

The clock input for the SPI bus; it can also be used to control the ADC conversion steps.

SPI_MISO

The SPI data output of the ADS8344.

SPI_MOSI

The SPI data input to the ADS8344.

SPI_*CS

This is the active low SPI chip select signal.

| PIC24FJ256GA106 Pinout for swFOCE Sensor Node Board |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pin # | Func1 | Func2 | Func3 | Func4 | Func5 | Func6 | Func7 | Functional Description |
| 1 | CN63 | PMD5 | RE5 |  |  |  |  | INSTR_3_OFF |
| 2 | SCL3 | CN64 | PMD6 | RE6 |  |  |  | INSTR_4_ON |
| 3 | SDA3 | CN65 | PMD7 | RE7 |  |  |  | INSTR_4_OFF |
| 4 | PMA5 | RP21 | C1IND | CN8 | RG6 |  |  | SPI_*CS |
| 5 | C1INC | RP26 | CN9 | PMA4 | RG7 |  |  | SPI_SCLK |
| 6 | C2IND | RP19 | CN10 | PMA3 | RG8 |  |  | SPI_MISO |
| 7 | *MCLR |  |  |  |  |  |  | *MCLR |
| 8 | C2INC | RP27 | CN11 | PMA2 | RG9 |  |  | SPI_MOSI |
| 9 | Vss |  |  |  |  |  |  | Vss |
| 10 | Vdd |  |  |  |  |  |  | Vdd |
| 11 | PGEC3 | AN5 | C1INA | RP18 | CN7 | RB5 |  | OPEN |
| 12 | PGED3 | AN4 | C1IN | RP28 | CN6 | RB4 |  | OPEN |
| 13 | AN3 | C2INA | CN5 | RB3 |  |  |  | OPEN |
| 14 | AN2 | C2INB | RP13 | CN4 | RB2 |  |  | OPEN |
| 15 | PGEC1 | AN1 | Vref- | RP1 | CN3 | RB1 |  | OPEN |
| 16 | PGED1 | AN0 | Vref+ | RP0 | CN2 | PMA6 | RB0 | OPEN |
| 17 | PGEC2 | AN6 | RP6 | CN24 | RB6 |  |  | PROG_CLK |
| 18 | PGED2 | AN7 | RP7 | CN25 | RB7 |  |  | PROG_DATA |
| 19 | AVdd |  |  |  |  |  |  | AVdd |
| 20 | AVss |  |  |  |  |  |  | AVss |
| 21 | AN8 | RP8 | CN26 | RB8 |  |  |  | CONFIG_BIT_0 |
| 22 | AN9 | RP9 | CN27 | PMA7 | RB9 |  |  | CONFIG_BIT_1 |
| 23 | TMS | AN10 | CVref | CN28 | PMA13 | RB10 |  | CONFIG_BIT_2 |
| 24 | TDO | AN11 | CN29 | PMA12 | RB11 |  |  | CONFIG_BIT_3 |
| 25 | Vss |  |  |  |  |  |  | Vss |
| 26 | Vdd |  |  |  |  |  |  | Vdd |
| 27 | TCK | AN12 | CTED2 | CN30 | PMA11 | RB12 |  | CONFIG_BIT_4 |
| 28 | TDI | PMA10 | AN13 | CTED1 | CN31 | RB13 |  | CONFIG_BIT_5 |
| 29 | AN14 | CTPLS | RP14 | CN32 | PMA1 | RB14 |  | CONFIG_BIT_6 |
| 30 | AN15 | REFO | RP29 | CN12 | PMA0 | RB15 |  | CONFIG_BIT_7 |
| 31 | SDA2 | RP10 | CN17 | PMA9 | RF4 |  |  | OPEN |
| 32 | SCL2 | RP17 | CN18 | PMA8 | RF5 |  |  | "Test LED" |
| 33 | RP16 | CN71 | RF3 |  |  |  |  | INSTR_4_STATE |
| 34 | RP30 | CN70 | RF2 |  |  |  |  | INSTR_3_STATE |
| 35 | ASCK1 | RPI45 | INT0 | CN72 | RF6 |  |  | DC-DC_CNTL |
| 36 | SDA1 | CN84 | RG3 |  |  |  |  | ADC_BUSY |
| 37 | SCL1 | CN83 | RG2 |  |  |  |  | ADC_*SHDN |
| 38 | Vdd |  |  |  |  |  |  | Vdd |
| 39 | OSCI | CLKI | CN32 | RC12 |  |  |  | OSCI |
| 40 | OSCO | CLKO | CN22 | RC15 |  |  |  | OSCO |
| 41 | Vss |  |  |  |  |  |  | Vss |
| 42 | RTCC | RP2 | CN53 | RD8 |  |  |  | UART_3_*SHDN |
| 43 | RP4 | CN54 | RD9 |  |  |  |  | UART4_TTL_TXD |
| 44 | RP3 | CN55 | PMCS2 | RD10 |  |  |  | UART4_TTL_RXD |
| 45 | RP12 | CN56 | PMCS1 | RD11 |  |  |  | OPEN |
| 46 | RP11 | CN49 | RD0 |  |  |  |  | UART1_TTL_TXD |
| 47 | SOSCI | C3IND | CN1 | RC13 |  |  |  | OPEN |
| 48 | SOSCO | C3INC | RPI37 | CN0 | RC14 |  |  | OPEN |
| 49 | RP24 | CN50 | RD1 |  |  |  |  | UART1_TTL_RXD |
| 50 | RP23 | CN51 | RD2 |  |  |  |  | UART2_TTL_TXD |
| 51 | RP22 | CN52 | PMBE | RD3 |  |  |  | UART2_TTL_RXD |
| 52 | RP25 | CN13 | PMWR | RD4 |  |  |  | UART3_TTL_TXD |
| 53 | RP20 | CN14 | PMRD | RD5 |  |  |  | UART3_TTL_RXD |
| 54 | C3INB | CN15 | RD6 |  |  |  |  | UART_1_*SHDN |
| 55 | C3INA | CN16 | RD7 |  |  |  |  | UART_2_*SHDN |
| 56 | Vcap | Vddcore |  |  |  |  |  | Vcap |
| 57 | ENVREG |  |  |  |  |  |  | ENVREG |
| 58 | CN68 | RF0 |  |  |  |  |  | INSTR_1_STATE |
| 59 | CN69 | RF1 |  |  |  |  |  | INSTR_2_STATE |
| 60 | CN58 | PMD0 | RE0 |  |  |  |  | INSTR_1_ON |
| 61 | CN59 | PMD1 | RE1 |  |  |  |  | INSTR_1_OFF |
| 62 | CN60 | PMD2 | RE2 |  |  |  |  | INSTR_2_ON |
| 63 | CN61 | PMD3 | RE3 |  |  |  |  | INSTR_2_OFF |
| 64 | CN62 | PMD4 | RE4 |  |  |  |  | INSTR_3_ON |

Table Six: Pin assignments of the SN board PIC

| PIC24FJ256GA106 Pin and Coding Assignments |  |  |  |  |
| --- | --- | --- | --- | --- |
| Pin # | Func1 | Signal Type | Default State | Schematic Name |
| 1 | RE5 | Digital Output | LOW | INSTR_3_OFF |
| 2 | RE6 | Digital Output | HIGH | INSTR_4_ON |
| 3 | RE7 | Digital Output | HIGH | INSTR_4_OFF |
| 4 | RG6 | SPI | HIGH | SPI_*CS |
| 5 | RP26 | SPI |  | SPI_SCLK |
| 6 | RP19 | SPI |  | SPI_MISO |
| 7 | *MCLR | Programming |  | *MCLR |
| 8 | RP27 | SPI |  | SPI_MOSI |
| 9 | Vss |  |  | Vss |
| 10 | Vdd |  |  | Vdd |
| 11 | RB5 | Digital Output | HIGH | OPEN |
| 12 | RB4 | Digital Output | HIGH | OPEN |
| 13 | RB3 | Digital Output | HIGH | OPEN |
| 14 | RB2 | Digital Output | HIGH | OPEN |
| 15 | RB1 | Digital Output | HIGH | OPEN |
| 16 | RB0 | Digital Output | HIGH | OPEN |
| 17 | PGEC2 | Programming |  | PROG_CLOCK |
| 18 | PGED2 | Programming |  | PROG_DATA |
| 19 | AVdd |  |  | AVdd |
| 20 | AVss |  |  | AVss |
| 21 | RB8 | Digital Input |  | CONFIG_BIT_0 |
| 22 | RB9 | Digital Input |  | CONFIG_BIT_1 |
| 23 | RB10 | Digital Input |  | CONFIG_BIT_2 |
| 24 | RB11 | Digital Input |  | CONFIG_BIT_3 |
| 25 | Vss |  |  | Vss |
| 26 | Vdd |  |  | Vdd |
| 27 | RB12 | Digital Input |  | CONFIG_BIT_4 |
| 28 | RB13 | Digital Input |  | CONFIG_BIT_5 |
| 29 | RB14 | Digital Input |  | CONFIG_BIT_6 |
| 30 | RB15 | Digital Input |  | CONFIG_BIT_7 |
| 31 | RF4 | Digital Output | HIGH | OPEN |
| 32 | RF5 | Digital Output | LOW | "Test LED" |
| 33 | RF3 | Digital Input |  | INSTR_4_STATE |
| 34 | RF2 | Digital Input |  | INSTR_3_STATE |
| 35 | RF6 | Digital Output | HIGH | DC-DC_CNTL |
| 36 | RG3 | Digital Input |  | ADC_BUSY |
| 37 | RG2 | Digital Output | HIGH | ADC_*SHDN |
| 38 | Vdd |  |  | Vdd |
| 39 | OSCI | Clock |  | OSCI |
| 40 | OSCO | Clock |  | OSCO |
| 41 | Vss |  |  | Vss |
| 42 | RD8 | Digital Output | HIGH | UART_3_*SHDN |
| 43 | RP4 | Serial |  | UART4_TTL_TXD |
| 44 | RP3 | Serial |  | UART4_TTL_RXD |
| 45 | RD11 | Serial | HIGH | OPEN |
| 46 | RP11 | Serial |  | UART1_TTL_TXD |
| 47 | RC13 | Digital Output | HIGH | OPEN |
| 48 | RC14 | Digital Output | HIGH | OPEN |
| 49 | RP24 | Serial |  | UART1_TTL_RXD |
| 50 | RP23 | Serial |  | UART2_TTL_TXD |
| 51 | RP22 | Serial |  | UART2_TTL_RXD |
| 52 | RP25 | Serial |  | UART3_TTL_TXD |
| 53 | RP20 | Serial |  | UART3_TTL_RXD |
| 54 | RD6 | Digital Output | HIGH | UART_1_*SHDN |
| 55 | RD7 | Digital Output | HIGH | UART_2_*SHDN |
| 56 | Vcap |  |  | Vcap |
| 57 | ENVREG |  |  | ENVREG |
| 58 | RF0 | Digital Input |  | INSTR_1_STATE |
| 59 | RF1 | Digital Input |  | INSTR_2_STATE |
| 60 | RE0 | Digital Output | LOW | INSTR_1_ON |
| 61 | RE1 | Digital Output | LOW | INSTR_1_OFF |
| 62 | RE2 | Digital Output | LOW | INSTR_2_ON |
| 63 | RE3 | Digital Output | LOW | INSTR_2_OFF |
| 64 | RE4 | Digital Output | LOW | INSTR_3_ON |

Table Seven: Pin assignments and software default states.