Possible test plan for determining the source of the MVC rebooting issue.

I propose we turn on one chamber at a time in an effort to isolate the source of the problem. It may take less time if we begin by having all of the instruments powered up on the chamber under observation (problem may occur quicker). Once the problem arises, re-initialize the chamber with one less Sensor Node. If the problem is narrowed down to a particular Sensor Node, then we can begin turning off the actual instruments on the Sensor Node one at a time.

One question is how long do we wait before determining that the rebooting problem will not happen? We have seen instances of it happening within 5 days or up to 3 weeks.

I propose the following order. Begin with chamber #1, everything on. If the problem does not surface, move onto chamber #2.

Chamber 1 (Camera, no thruster)

All Sensor Nodes and Instruments on

Turn off Sensor Node SN# 9

Turn off Sensor Node SN# 8

Turn off Sensor Node SN# 7

Turn off Sensor Node SN# 6

Chamber 2 (No camera, thruster)

All Sensor Nodes and Instruments on

Turn off Sensor Node SN# 11

Turn off Sensor Node SN# 18

Turn off Sensor Node SN# 17

Turn off Sensor Node SN# 16

| GN SN#/Port | SN SN# | Port | Instrument | Instr Manuf | Model # | Serial # | SSDS ID# | Sensor Unit |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| #1 / 1 | 6 | 1 | pH Probe | SeaBird | SBE18S | 181067 | 1813 | 2 |
| #1 / 1 | 6 | 2 | pH Probe | SeaBird | SBE18S | 181069 | 1812 | 2 |
| #1 / 1 | 6 | 3 | CTD-O2 | SeaBird | SBE52-43 | 624878-0047 | 1642 | 2 |
| #1 / 1 | 6 | 4 | Pump | Seabird | SBE5T | 057936 | 1850 | 2 |
| #1 / 2 | 7 | 1 | pH Probe | SeaBird | SBE18S | 181066 | 1818 | 1 |
| #1 / 2 | 7 | 2 | pH Probe | SeaBird | SBE18S | 181065 | 1817 | 1 |
| #1 / 2 | 7 | 3 | CTD-O2 | SeaBird | SBE52-43 | 5248473-0044 | 1733 | 1 |
| #1 / 2 | 7 | 4 | Pump | SeaBird | SBE5T | 054258 | 1837 | 1 |
| #1 / 3 | 8 | 1 | PAR Sensor | Satlantic |  | 401 | 1824 |  |
| #1 / 3 | 8 | 2 | PAR Sensor | Satlantic |  | 400 | 1851 |  |
| #1 / 3 | 8 | 3 | [open] |  |  |  |  |  |
| #1 / 3 | 8 | 4 | [open] |  |  |  |  |  |
| #1 / 4 | 9 | 1 | Valve Box | MBARI | N/A | 1 | 1846 |  |
| #1 / 4 | 9 | 2 | Motor |  |  |  |  |  |
| #1 / 4 | 9 | 3 | Flowmeter | Omega | FMG-554 |  |  |  |
| #1 / 4 | 9 | 4 | Motor |  |  |  |  |  |

| #2 / 1 | 16 | 1 | pH Probe | SeaBird | SBE18S | 180943 | 1848 | 3 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| #2 / 1 | 16 | 2 | pH Probe | SeaBird | SBE18S | 181070 | 1833 | 3 |
| #2 / 1 | 16 | 3 | CTD-O2 | SeaBird | SBE52-43 | 5248473-0046 | 1752 | 3 |
| #2 / 1 | 16 | 4 | Pump | Seabird | SBE5T | 054310 | 1836 | 3 |
| #2 / 2 | 17 | 1 | [open] |  |  |  |  |  |
| #2 / 2 | 17 | 2 | [open] |  |  |  |  |  |
| #2 / 2 | 17 | 3 | [open] |  |  |  |  |  |
| #2 / 2 | 17 | 4 | [open] |  |  |  |  |  |
| #2 / 3 | 18 | 1 | PAR Sensor | Satlantic |  | 402 | 1820 |  |
| #2 / 3 | 18 | 2 | [open] |  |  |  |  |  |
| #2 / 3 | 18 | 3 | [open] |  |  |  |  |  |
| #2 / 3 | 18 | 4 | [open] |  |  |  |  |  |
| #2 / 4 | 11 | 1 | Valve Box | MBARI | N/A | 3 | 1847 |  |
| #2 / 4 | 11 | 2 | Motor |  |  |  |  |  |
| #2 / 4 | 11 | 3 | Flowmeter | Omega | FMG-554 | 2 |  |  |
| #2 / 4 | 11 | 4 | Motor |  |  |  |  |  |