FOCE Chassis Power Supply Engineering Change Orders

A - 3.3V_SYS regulator wired incorrectly

Pin 3 of U24 should be tied to the common node of R61/C75.

As built:

Reworked to:

B - Configure rs-232 Rx/Tx lines

Remove R63 and R65.

C - Modify Beta-Dyne on/off circuitry

Remove R46

D - Oscillations on 3.3V Beta-Dyne

Remove L3 and replace with short (18AWG bare wire) [this is a temporary fix]

E - Replace input capacitors

Replace C42, C43.

Old part: 82uF PXH series

New part: 100uF MVE series. PN# EMVE630ADA101MJA0G

Replace C14, C15, C29, C30.

Old part: 180uF PXH series

New part: 220uF MVE Series. PN# EMVE500ADA221MJA0G

F - Connect input and output return lines

Connect U10-3 to U10-8 with 20AWG wire

Connect U13-3 to U13-8 with 20AWG wire

Connect U16-3 to U16-8 with 20AWG wire

G - Connect SPI bus #2 clock signal to MAX6675's

Connect U1-26 to U18-5.

H - Tie ADC shutdown* signal high

Solder a 30AWG wire from V3 to U3-20 (or any 3.3V location).

J - Add a chassis ground connection

Add a connection between C13/C17 node, C27/C31 node, and C40/C44 node to chassis ground connection (screwhole added to rev B board).