# Installing QuickFOCE on a Linux Server

## Background

QuickFOCE consists of an electronics can containing power conversion hardware and a Digi Portserver TS4 MEI, a 4 port Ethernet to serial converter.  The 4 serial ports service:

- A Seabird SBE-52 CTD mounted to measure the ambient (background) water

- Two Seabird SBE-18S pH sensors – one for in the chamber, one for ambient conditions

- A motor controller

These instruments are sampled and controlled by a Linux workstation on shore.  In the case of SW-FOCE at Hopkins, this server is kelpbed.swfoce.mbari.org.  The workstation for the MBARI test tank setup is fangtooth.shore.mbari.org.

## Linux

The preferred, tested Linux setup is to install CentOS 6.5.  Once CentOS  is installed, you need to do the following as root to install the required packages (some of these are required for the Digi driver, some for the xFOCE software):

- yum update

- yum install gcc

- yum install libtool

- yum install gcc-c++.x86_64

- yum install kernel-devel

- yum install emacs  (my preferred editor, YMMV)

- yum install ncurses-devel ncurses

- yum install tcl

- yum install tk

- yum install openssl-devel

- yum install libuuid-devel

- yum install mercurial

## Digi Driver

The correct version of the Digi driver for CentOS is dgrp-1.9.  The gzipped tar file is in ProjectLibrary  under “901219 SW-FOCE\Software\QuickFOCE\Digi\ 40002086_S.tgz.

- Extract this tarball somewhere convenient

- ./configure

- make all

- make install

- make postinstall

You can now run the Digi gui, dgrp_gui.  Mapping the correct Portserver should be obvious.  We recommend that you use static IP addressing for the Portserver.

## xFOCE Software

To build the xFOCE software, you’ll need to find and download the following:

- The xFOCE source.  This can be obtained from Mercurial via the command line
hg clone https://bitbucket.org/klheadley/xfoce

- libmodbus – as of this writing, the appropriate version is libmodbus_3.0.6.tar.gz.  You may get it from libmodbus.org

- zeromq and czmq, both available from zeromq.org.   Due to a build incompatibility (which hopefully will be resolved soon), as of this writing you’ll need slightly old version of the above.  Use zeromq-4.0.1.tar.gz and czmq-2.0.2.tar.gz

### Support software

For each of the support packages, libmodbus, zeromq, and czmq, you need to do the following:

- untar the .gz file

- cd into the directory just created

- ./configure

- make

- (as root) make install

### xFOCE software

- edit configure.ac to change to AC_PREREQ([2.60]), and to allow a $host other than those configured.  The default $host clause should do an AC_MSG_WARN() instead of AC_MSG_ERROR(), and use the macros from i686-pc-linux-gnu

- libtoolize –f

- aclocal

- autoconf

- automake –add-missing

- autoreconf

- ./configure

- make

You do not need to ‘make install’, though you may do so if you wish.  But it’s easiest just to run it from the build directory.  To do so, you need to:

- Set the XF_HOME environment variable to the xFOCE install directory

- . $XF_HOME/src/scripts/xfhome $XF_HOME

- gwrun [-f config_file]

An example config file is included in $XF_HOME/conf/simple.conf.  So you can simply run

gwrun –f simple.conf