## Contents
Exportable FOCE Requirements
Version | Description | Date | Topics
0 | Early discussions by SW/HW teams | 2011-12-06 | Engineering discussion of essential features and specs
1 | Discussion (Barry, Peltzer, Walz, FOCE engr) | 2012-02-01 | pH offset range experimental design for biology instrument suite priorities MBARI permitting limits maintenance and telemetry autonomy/reliability Instrument Suite Worksheet
2 | SE/EE sidebar discussions | 2012-03-15 | metadata tracking velocity actuators
3 | Pre-Concept DR | 2012-04-13
4

## requirements
Design | Telemetry
Design | Device I/O Support
Design | Core Software
Design | Software Language Support
Design | Input Voltage Section
Design | LV DC Conversion
Design | Environmental
Design | Environmental Sensors
Functional | Performance
Functional | Environmental
Functional | Core Instrument Suite
Functional | Operation & Maintenance
Functional | Data Management |  |  | unspecified
Functional | User Interface |  |  | required
Functional | Power Management |  |  | optional
Functional | Reliability |  |  | preferred
Type | Category | Reference | Topic | Level | Spec | Comments | Add/Chg/Del since previous version
Functional | Performance | FP.1 | number of experimental units | required | 3 control units (w/ CO2) 3 experimental units (w/o CO2) 400L CO2/year (MBARI permit) CO2 generator and deliver systems must be appropriately scalable to meet experimental requirements
Functional | Performance | FP.2 | maintain a pH offset relative to a reference value | required | min: -0.2 typ: -0.4 to -0.8 max: -1.5 pH range: resolution: 0.05 pH units, 0.1 unit may be OK higher resolution (0.003-0.01) required for chemistry experiments Vint:= 5 cm/sec | Natural variation of ~ 0.8 to 1 pH unit can occur over a day
Functional | Performance | FP.3 | maintain a specified constant pH value | required | 7.0 < pH < 8.2 independent of ambient pH internal <= 5 cm/sec
Functional | Performance | FP.4 | maintain a specified Vint | required | 2-5 cm/sec resolution?
Functional | Environmental | FE.1 | Max depth | required | 40m (130 ft) Note diver-servicable requirement Not suitable for active surf, river mouth
Functional | Environmental | FE.2 | Water temperature range | required | -2 to 30 degrees C
Functional | Environmental | FE.3 | Bio-fouling protection/prevention | required | Hard to spec, Can document recommendations Environmental requirements trades (e.g. toxicity) Ties to reliability and maintenance (is a maintenance schedule issue) e.g. passive, Cu…
Functional | Environmental | FE.4 | Min external VH20 | required | 0 cm/sec
Functional | Environmental | FE.5 | Max external VH20 | required | 50 cm/sec
Functional | Environmental | FE.6 | Min internal VH20 | required | 2 cm/sec
Functional | Environmental | FE.7 | Max internal VH20 | required | 0 cm/sec 50 cm/sec 2 cm/sec
Functional | Environmental | FE.8 | Min/max surge | required | Must maintain pH in some surge conditions. Surge is mostly moving water back and forth in chamber – should not affect control (?)
Functional | Environmental | FE.9 | surge replication | preferred | To the extent possible, the system should replicate natural surge as needed. The mechanism for this may not be needed for all experiments, but is critical in some (e.g. for filter feeders); this should be a modular feature
Functional | Environmental | FE.10 | Light transmission | required | chamber should enable maximum natural light to be transmitted into the chamber for greatest replication of ambient conditions
Functional | Environmental | FE.11 | Water flow | required | chamber should enable maximum natural water flow to be transmitted into the chamber for greatest replication of ambient conditions | for grasses, e.g., know order of magnitude diff between inside and outside
Functional | Core Instrument Suite | FIS.1 | Instrument suite summary | unspecified | see worksheet
Functional | Core Instrument Suite | FIS.2 | pH | required | 1xn experiment, 1 control, 1 environment | Electrical and communications interfaces? Is reference measured in ambient? Control? valid T required for pH measurement - may be internal or external to pH sensor number required is for basic science; additional may be needed for control purposes
Functional | Core Instrument Suite | FIS.3 | CTD | required | T required for pH measurement 1xn experiment, 1 control, 1 environment | high on the nice to have list, but not required for all science number required is for basic science additional may be needed for control purposes
Functional | Core Instrument Suite | FIS.4 | DO2 | required | not required for pH control required for biology experiment not required for deep/chemistry experiment 1xn experiment, 1 control, 1 environment | number required is for basic science; additional may be needed for control purposes
Functional | Core Instrument Suite | FIS.5 | PAR | preferred | 1xn experiment, 1 control, 1 environment | high on the nice to have list A PAT tag pin-diode may be OK, if it can be adequately calibrated Must provide good comparative inside v outside) measurement of spectrum may be possible to use reduced number of sensors (share among chambers)
Functional | Core Instrument Suite | FIS.6 | H2O flow (velocity) | required | 1xn experiment, 1 control, 1 environment Nice to have as sci context data, required for control Range, precision, rate? | 2D OK Either volume or velocity Not both Sci would like to know how much (low pH) water has flowed through over time period, and in which direction. Could use inexpensive flow meters instead of velocimeter. number required is for basic science; additional may be needed for control purposes
Functional | Core Instrument Suite | FIS.7 | H2O flow (volume) | required | Either volume or velocity Not both
Functional | Core Instrument Suite | FIS.8 | Camera (time lapse video) | required | xFOCE must support camera expansion Required for swFOCE How will video be used? Frame rate 1/min max 1/30-60min to 1/day typ previously estimated at 10 frames/min) Resolution: publication quality (e.g. HD, Coolpix…5-12 MP) Supplement to measurement data | access to video early on is important to know if data quality of OK to get final result
Functional | Core Instrument Suite | FIS.9 | Camera (video) | unspecified | TBD | For swFOCE observation - spec TBD (e.g. w/ pan and tilt, be able to view antennae movement)
Functional | Core Instrument Suite | FIS.10 | Light panel | required | Required for camera Output, angle, power…
Functional | Core Instrument Suite | FIS.11 | Flow effectorr(s) | required | Quantity: 1 min, app-specific, number and type may change Spec Depends: VH20 range pH offset range
Functional | Core Instrument Suite | FIS.12 | ESW effector(s) | required | Quantity 1 min; depends number of thrusters, use of valves Spec Depends: pH offset range deployment duration
Functional | Core Instrument Suite | FIS.13 | ESW valve(s) | required | Depends on injection design
Functional | Core Instrument Suite | FIS.14 | ESW pH | preferred | Useful when at end of ESW supply to modify feed algorithm
Functional | Core Instrument Suite | FIS.15 | ESW CTD | optional | Should, but not core Number, type?
Functional | Core Instrument Suite | FIS.16 | Instrument expansion capability | preferred
Functional | Operation & Maintenance | FOM.1 | accessibility | required | easy diver access to chamber
Functional | Operation & Maintenance | FOM.2 | accessibility | required | sensors positioned to enable access to experiment
Functional | Operation & Maintenance | FOM.3 | accessibility | preferred | it must be possible to reposition cameras for maintenance and observation
Functional | Operation & Maintenance | FOM.4 | Small-boat deployable | required | Zodiac/whaler
Functional | Operation & Maintenance | FOM.5 | Diver serviceable | required | Installation, cleaning, instrument replacement, recovery
Functional | Operation & Maintenance | FOM.6 | In-situ replaceable instruments | preferred | Must for pH (cal) Others should? Maintenance schedule
Functional | Operation & Maintenance | FOM.7 | sealable chamber | required | Remotely operated is requirement?
Functional | Operation & Maintenance | FOM.8 | In-situ pH calibration | optional | Depends on cost trade Nice to have
Functional | Operation & Maintenance | FOM.9 | Deployment/experiment duration | required | Design for 6 month deployment | Objective is to run realistic long- term, in-situ community /multi species experiments months to years variety of environments (lakes, estuaries, oceans) Something that could work from a Lobo-class mooring would be great. LTER sites (yrs-decade records) would benefit from FOCE. Typical experiments would run < 6 months' a couple car batteries for a few weeks at a time would be OK
Functional | Operation & Maintenance | FOM.10 | Operate continuously | required | Runs overnight w/o monitoring min
Functional | Operation & Maintenance | FOM.11 | Operate autonomously | preferred | Set it up and run it, but person in loop to correct anomalies …requirements for event detection and/or data visualization
Functional | Operation & Maintenance | FOM.12 | Maintenance schedule | unspecified | Will depend on application, ranging from every day or two (e.g. at start of shallow experiment) to every few days, with periodic system checks. | Checks may be infrequent, need to be able to tell whether system is working (need more information on appropriate granularity of diagnostics and repair - red/green light, console with report(s) and utils, etc. ) What faults is the user capable of addressing? What components should be field-replaceable?
Functional | Data Management | FDM.1 | Data logging (remote side) | required | Requirement is that all data must be recorded (independent of telemetry function) and available at end of experiment. Full-resolution science and system data (for forensic reconstruction of events/anomalies) Must be independent of telemetry
Functional | Data Management | FDM.2 | Adjustable device sample rates | preferred | Must be able to set sample rates to provide best resolution for diagnostics, and lower sample rates for logging and performance. See data logging rates below
Functional | Data Management | FDM.3 | Data logging rates (core instruments) | required | TBD… CTD, pH, velocity/flow Rates up to 2 Hz (10 Hz?) Must support system and science data stream rates Depends on environmental conditions, velocity sampling strategies
Functional | Data Management | FDM.4 | Data archive (shore side) | unspecified | Split opinions – must vs should Single point of failure/enabler Vs Not strictly required to do FOCE Part of system to get data to shore Maybe not automatic Maybe very low tech (scp) Cost vs value
Functional | Data Management | FDM.5 | Data log format | optional | Timestamped data Any format OK as long as complete, robust. If not human readable, need tools to access records. Segmented logs? Must be able to request specific records using specified time range(s) or number of records | Simple timestamped comma delimited flat files are OK Need mechanism to decimate/filter data (too much data is hard to work with)
Functional | Data Management | FDM.6 | Data log capacity | required | Must record all science and system data for 6 month deployment. Depends: data rates, number of streams, sample rates, data log format
Functional | Data Management | FDM.7 | Near real-time telemetry | required | Full data resolution if possible, but all data must be recorded even w/o telemetry link. Must provide data required to assess health and status of experiment. Rate: best effort to reflect full-resolution logged data.
Functional | Data Management | FDM.8 | Telemetry data summarization | preferred | Data telemetry must support system health and status reporting. Should make best effort to reflect full-resolution logged telemetry, but this is not required. Alarm is summary Some form probably needed
Functional | Data Management | FDM.9 | Near real-time data visualization | preferred | Must be able to see from shore whether the system is working. What does it mean to be working? How do we tell? What streams? How often updated? Should, but not must? What is cost?
Functional | Data Management | FDM.10 | Event detection and notification | preferred | Should do if telemetry Data telemetry must support system health and status reporting. Other event detection and reporting would be great, but not required. What events? Sci or Sys-only?
Functional | Data Management | FDM.11 | Timestamp data | required | Need to correlate data streams in memory Accuracy? Resolution?
Functional | Data Management | FDM.12 | Time distribution | unspecified | Timestamp data Time distribution Timing determinism
Functional | Data Management | FDM.13 | Timing determinism | required | ~0.5 max sample rate – Nyquist freq.
Functional | Data Management | FDM.14 | Metadata Management | unspecified | The instrument data stream shall (automatically) include metadata describing the identity and state of the instrument and its driver. New metadata shall be (automatically) added to the stream when the state of the instrument or driver changes in a way that affects data interpretation. User shall be able to associate any instrument data record with the instrument and driver metadata in effect at the time the data was acquired. | klh not sure this a strict requirement - it is possible to control pH in the box w/o doing this. I see the value, but the users need to weigh in, and we need to understand how it drives other requirements (e.g. choice of processor, language, etc.). The costs of meeting this proposed requirement need to be determined; alternative implemtations need to be proposed and traded.
Functional | User Interface | FUI.1 | System configuration and control | required | Supports: Power management Data stream read (instantaneous) Instrument config?/mgmt Control tuning Experiment control Retrieve data logs pH cal?
Functional | User Interface | FUI.2 | Data visualization | required | Supports: Power management Data stream read (instantaneous) Instrument config?/mgmt Control tuning Experiment control Retrieve data logs pH cal?
Functional | Power Management | FPM.1 | Total system power | unspecified | Minimize… Depends: power system implementation choice Amount of water moved What does current FOCE actually use? TBD Watts
Functional | Power Management | FPM.2 | Max inrush current | unspecified | Need spec (like sys power)
Functional | Power Management | FPM.3 | Input voltage range | unspecified | “
Functional | Power Management | FPM.4 | Input current range | unspecified | “
Functional | Power Management | FPM.5 | Monitor individual load voltages and currents | unspecified | Range? Rate/resolution/accuracy? In hardware must Operator monitoring should
Functional | Power Management | FPM.6 | Monitor bus voltages and currents | unspecified | Range? Rate/resolution/accuracy? In hardware must Operator monitoring should
Functional | Power Management | FPM.7 | Switch individual load power | required | Must be maintainable
Functional | Power Management | FPM.8 | Switch bus power | unspecified
Functional | Reliability | FR.1
Design | Telemetry | DT.1 | Support wired communication to shore | optional | Requirement is for some type of communication to shore; this is an implementation option
Design | Telemetry | DT.2 | Support wireless communication to shore | optional | Requirement is for some type of communication to shore; this is an implementation option
Design | Telemetry | DT.3 | Support TCP/IP protocols | optional | Requirement is for some type of communication to shore; this is an implementation option
Design | Telemetry | DT.4 | Support 10BaseT Ethernet or better | optional | Requirement is for some type of communication to shore; this is an implementation option
Design | Telemetry | DT.5 | Support remote console | required | e.g., serial, ssh, telnet
Design | Device I/O Support | DDI.1 | RS-232 device I/O | optional | Requirement is to support the core set of instruments. This is an implementation option. up to 115200 bps Number of channels? Expandable?
Design | Device I/O Support | DDI.2 | Analog device I/O | optional | Requirement is to support the core set of instruments. This is an implementation option. up to 115200 bps Number of channels? Expandable?
Design | Device I/O Support | DDI.3 | Ethernet device I/O | optional | Requirement is to support the core set of instruments. This is an implementation option. Number of channels? Input range? Expandable?
Design | Device I/O Support | DDI.4 | HTTP device I/O | optional | Requirement is to support the core set of instruments. This is an implementation option.e.g. cameras
Design | Device I/O Support | DDI.5 | Memory mapped device I/O | optional | Requirement is to support the core set of instruments. This is an implementation option.Number of channels? Expandable? e.g. enclosure sensors
Design | Device I/O Support | DDI.6 | USB device I/O | optional | Requirement is to support the core set of instruments. This is an implementation option.
Design | Core Software | DCS.1 | Use open source/free software | optional
Design | Core Software | DCS.2 | Support IPC (inter-process communication) | optional
Design | Core Software | DCS.3 | Thread support | optional
Design | Core Software | DCS.4 | Support for chosen I/O methods | required | RS232, RS422, RS485 (or other bus), Ethernet, Analog…
Design | Core Software | DCS.5 | Support in-situ software updates | optional | Nice to have
Design | Core Software | DCS.6 | Support in-situ firmware updates | optional | Nice to have
Design | Core Software | DCS.7 | Support dynamic device driver loading |  | Start, stop, replace individual instrument drivers
Design | Core Software | DCS.8 | Support integration of new sensors | required | Requirement is to lower barriers of initial adoption/participation by new users. Make it easy to try FOCE technology, w/ less emphasis on extending or customizing. Not strictly required to provide generic mechanism here, but do want to show others an expansion path. e.g. instrument service framework Support for serial, analog, Ethernet, multidrop bus Support for polled, streaming
Design | Core Software | DCS.9 | Support data pub/sub | optional | Implementation option streaming control: Per channel/per device?
Design | Core Software | DCS.10 | Support system verification and test Pre- and post-deployment | required | Must be able to verify that the system is working before (and after) it is deployed
Design | Core Software | DCS.11 | Enable modification of control algorithms/strategies | required | Must make it easy to tune and adjust control strategy and/or algorithms for various applications Configurable at minimum
Design | Core Software | DCS.12 | System data monitoring | required | Update rates? Logging? Telemetry? Bus voltage(s)/current(s) Device voltage/current Over/under voltage conditions Ground faults Enclosure Temp Enclosure Humidity Enclosure Pressure Enclosure Water intrusion Heat sink Temp(s) Disk, CPU, memory usage Must be recorded locally Actual list may depend…
Design | Core Software | DCS.13 | Support hardware power management | required | Requirement is to minimize power consumption where possible. Support whatever modes are available and chosen for selected hardware sleep modes, clock throttling, etc.
Design | Software Language Support | DSL.1 | Support C language | optional | Implementation option
Design | Software Language Support | DSL.2 | Support Java language | optional | Implementation option
Design | Software Language Support | DSL.3 | Support BASIC language | optional | Implementation option
Design | Software Language Support | DSL.4 | Support LabView language | optional | Implementation option
Design | Software Language Support | DSL.5 | Support a visual prog language | optional | For what?
Design | Software Language Support | DSL.6 | Support MatLab language | optional | For what?
Design | Software Language Support | DSL.7 | Support scripting languages | optional
Design | Input Voltage Section | DIV.1 | Accept input DC volt up to 400Vdc | required | What drives 400V requirement? Most choose a (common) input voltage … somewhat changeable at low cost. Requirement: Accept inputs that may be able to down-convert to 48V
Design | Input Voltage Section | DIV.2 | Accept input DC volt up to 50Vdc | unspecified | see above 400V reqmt
Design | Input Voltage Section | DIV.3 | Accept input DC current to x amps | unspecified
Design | Input Voltage Section | DIV.4 | Provide galvanic isolation to source | unspecified
Design | Input Voltage Section | DIV.5 | Provide inrush limiting | unspecified
Design | Input Voltage Section | DIV.6 | Be fused / cicuit breaker | unspecified
Design | LV DC Conversion | DLV.1 | Provide multiple low voltages | unspecified | 5, 12, 24, 48
Design | LV DC Conversion | DLV.2 | Provide voltage monitoring | unspecified | Hardware level
Design | LV DC Conversion | DLV.3 | Provide current monitoring | unspecified
Design | LV DC Conversion | DLV.4 | Monitor for over voltage to instr | unspecified
Design | LV DC Conversion | DLV.5 | Monitor for under voltage to instr | unspecified
Design | LV DC Conversion | DLV.6 | Monitor for over current to instr | unspecified
Design | LV DC Conversion | DLV.7 | Provide a resettable circuit breaker | unspecified
Design | LV DC Conversion | DLV.8 | Monitor for ground faults | unspecified
Design | LV DC Conversion | DLV.9 | Provide analog reference voltage | unspecified
Design | LV DC Conversion | DLV.10 | Automatic over-voltage protection | unspecified | +/- 5, 12 V
Design | LV DC Conversion | DLV.11 | Automatic over-current protection | required
Design | LV DC Conversion | DLV.12 | Ground fault detection | unspecified
Design | LV DC Conversion | DLV.13 | Galvanic isolation (per load) | unspecified
Design | Environmental Sensors | DIS.1 | Provide ambient temp monitoring | required | 0.016 Hz (1 min) nominal
Design | Environmental Sensors | DIS.2 | Provide ambient humidity monitoring | required | 0.016 Hz (1 min) nominal
Design | Environmental Sensors | DIS.3 | Provide ambient pressure monitoring | required | 0.016 Hz (1 min) nominal
Design | Environmental Sensors | DIS.4 | Provide water presence detection | required | 0.016 Hz (1 min) nominal
Design | Environmental Sensors | DIS.5 | Provide temp monitoring of various subsystems | unspecified | More specific
Design | Environmental | DE.1 | Non-toxic | unspecified | May be required in some applications; no general requirement, but less hazardous is better :o)
Design | Environmental | DE.2 | Non-hazardous to animals | unspecified | May be required in some applications; no general requirement, but less hazardous is better :o)
Design | Environmental | DE.3 | light transmission | required | materials should be selected that enable the best fidelity of natural light transmission (to replicate local environment)
Design | Environmental | DE.4 | out gassing | required | materials must minimize out-gassing spec?

## worksheet
FOCE Requirements Worksheet | Concept for gathering requirements - maybe filled out by end-users - researchers (science instruments) and engineers (system devices) or by engineers via interview w/ researchers
geographic location
depth (approx)
distance from shore (approx)
objectives
experiment duration
number of experiments
number of controls
environment (waves, currents, temp, proximity to public access, etc…)
power source (Local utility, generator, solar, wind, batteries, etc.)
telemetry connection (Copper, Fiber, Wireless, none, etc.)
CO2 source (bottled gas, generator, etc.)
Sensor/Actuator | device type | pH | T | CTD | DO2 | PAR | flow | pCO2 | velocity | timelapse camera | video camera | light panel
Requirement | Degree to which device is required to attain science objectives R: required P: preferred O: optional D: depends (use depends on inclusion another device) | R | D | P | P | P | P | P | O | P | O | D
Priority | Level of importance in preferred devices on a scale of 1 to 5 (e.g. what would you choose to keep running if power were limited) 1(less important)-5 (most important) | P: 5; P: 5; P: 4; P: 3; P: 3; P: 4
Criticality | Degree to which experimental outcomes depend on device operation and/or data . 0: diagnostic context (possibly not needed) 1: other measurements depend on operation, but data can be backed out if lost 2: context related to outcome (valid outcomes possible if lost) 3: critical context or measurement (can not succeed if completely lost) | 3 | 1 | 2 | 2 | 2 | 2 | 2 | 0 | 2 | 0 | 0
Experimental Units | total number of (e.g. 1) or number per (e.g. 1xn) experiment | R: 1xn; D: 1xn; P: 1xn; P: 1xn; P: 1xn
Control Units | total number of (e.g. 1) or number per (e.g. 1xn) control | bar: R to P
Environment Units | total number of (e.g. 1) or number per (e.g. 1xn) environment | bar: R to P
Resolution | measurement resolution | R: 0.05 units; D: 0.001 deg C; P: 0.001 deg C 0.1 m
Min Sample Rate | miniumum rate (max period) at which it is useful to acquire and store data | R: 10 min
Preferred Sample Rate | preferred rate (max period) at which it is useful to acquire and store data | R: 5 min
Data Shelf Life | Period after which (contextual) data is not useful and may be discarded (blank: never)
Operation period | blank: 24x7xduration (e.g., only during daylight hours, 2 h/day, etc.)
Min Telemetry Period | how often is telemetry needed during the experiment U: unsure N: never R: upon request or specify period in hours, days, etc. | R: R, 60 m
Telemetry Service | A: all L: latest D: decimated/sub-sampled S: stats (min,max,avg) O: other (specify) blank: unsure | bar: R
Maintenance | will need to change/replace this instrument N - never F - on failure A - as needed S -on a schedule <specify> O: other <specify> blank - unsure | bar: R
Specifications | relevant specifications (manufacturer, model, resolution, accuracy, precision, technical specs, etc.)
Comments |  | D: pH depends on T, and may have integrated T measurement T measurement should take place phycially close to pH measurement; D: only needed for cameras wavelengths affect experiment (manage on-time)

## eswCalcs
flow meter range
ESW concentration (mmol/L) | 40
LCO2Density (kg/m^3) | 770
LCO2Density (kg/L) | 0.77
p (kg/m^3) | 1031.4 | 1022.8
w (m) | 0.5
h (m) | 0.5
l (m) | 3
v (m/s) | 0.02 | 0.02 | 0.04 | 0.04 | 0.1 | 0.1 | 0.2 | 0.2 | 1 | 1
pH shift | 0.1372 | 0.2872 | 0.1372 | 0.2872 | 0.1372 | 0.2872 | 0.1372 | 0.2872 | 0.1372 | 0.2872
CO2 (umol/kg) | 78 | 150.5 | 78 | 150.5 | 78 | 150.5 | 78 | 150.5 | 78 | 150.5
V (m^3/s) | 0.005 | 0.005 | 0.01 | 0.01 | 0.025 | 0.025 | 0.05 | 0.05 | 0.25 | 0.25
massH2O (kg/s) | 5.157 | 5.157 | 10.314 | 10.314 | 25.785 | 25.785 | 51.57 | 51.57 | 257.85 | 257.85
massH2O (kg/min) | 309.42 | 309.42 | 618.84 | 618.84 | 1547.1 | 1547.1 | 3094.2 | 3094.2 | 15471 | 15471
CO2 (mmol/min) | 24.1348 | 46.5677 | 48.2695 | 93.1354 | 120.6738 | 232.8386 | 241.3476 | 465.6771 | 1206.738 | 2328.3855
ESW (L/min) | 0.60337 | 1.1642 | 1.2067 | 2.3284 | 3.0168 | 5.821 | 6.0337 | 11.6419 | 30.1685 | 58.2096
ESW (L/hr) | 36.2021 | 69.8516 | 72.4043 | 139.7031 | 181.0107 | 349.2578 | 362.0214 | 698.5157 | 1810.107 | 3492.5783
phSensitivity (deltaPH/mmol/kg) | 2.55
deltaPH | 0.8
v (m/s) | 0.02 | 0.04 | 0.1 | 0.2 | 0.5 | 1
V (m^3/s) | 0.005 | 0.01 | 0.025 | 0.05 | 0.125 | 0.25
massH2O (kg/s) | 5.157 | 10.314 | 25.785 | 51.57 | 128.925 | 257.85
massH2O (kg/min) | 309.42 | 618.84 | 1547.1 | 3094.2 | 7735.5 | 15471
ESWUseRate (L/min) | 2.4268 | 4.8536 | 12.1341 | 24.2682 | 60.6706 | 121.3412
ESWUseRate (L/hr) | 145.6094 | 291.2188 | 728.0471 | 1456.0941 | 3640.2353 | 7280.4706
ESWUseRate (L/day) | 3494.6259 | 6989.2518 | 17473.1294 | 34946.2588 | 87365.6471 | 174731.2941
LCO2 (kg/day) | 6.1519 | 12.3037 | 30.7593 | 61.5187 | 153.7967 | 307.5935
LCO2 (lb/day) | 13.5624 | 27.1248 | 67.8121 | 135.6241 | 339.0603 | 678.1206