## Slide 1: Software side of xFOCE
(Computing system trades)
FIXIT

## Slide 2: Computing Architectures User Concerns
Performance: proven science approach, data quality
Flexibility: wide range of environments, scenarios
Ease:
  Build
  Test
  Use, maintain
  Extend
Cost

## Slide 3: A few assumptions
Implementer profile | Researchers plus 2-4 post docs, grad students 0-1 engineer
Technical | Access to modest manufacturing facilities Software is harder than hardware, mechanical Strong preference for familiar COTS solutions
Implementation schedule | A few months to build
Experiments | Campaign-style deployments 12 month service life Remote and/or restricted sites
Lifecycle | Prefer to extend capability, rather than re-spin from scratch for successive funding rounds

## Slide 4
Controller
Enriched
Seawater
Power
Data
CO2
H2O
FOCE Essentials
Flow
Flow
CTD
pH
CTD
pH
DO
DO

## Slide 5: Computing Architectures Elements
Instruments
Software
Housings
Electronics
Computers
Cabling
For each, many options
Many ways to arrange them
In a FOCE experiment

## Slide 6: Computing Architectures Shrinking the Universe of Trades
Look at decisions that have greatest impact on user concerns
  Build vs Buy
  Location matters: Shore vs Surface vs Seafloor
  Topology: how the pieces are connected

## Slide 7: Build vs Buy General Principles
Buy what you can’t/shouldn’t build
  Processors, instruments, cabling
  Software
Build (or borrow) as needed
  Chambers, housings
  Use COTS components if possible
  Get help on tough key builds
    contract, open source
Beware Frankensystem
  Patchwork integration of COTS parts
  Power consumption
  Rube Goldberg-ness of mechanical/electrical interfaces

## Slide 8: Location Matters General Principles
Connected vs Disconnected
  Generally easier to generate and store power, ESW on shore
  Telemetry bandwidth
Trades for off-shore elements
  Accessibility (to you, the elements, others)
  Reliability
Keep what’s offshore/sub-surface as simple and robust as possible
  Basics: Enriched seawater, power
  HW: Cables, connectors, housings
  SW: Applications, protocols, processes
Play your application, tech strengths and budget
  Permits, environment
  Experiment duration, depth
  Data access requirements
  Maintenance schedule, spares
  Available design/fabrication resources

## Slide 9
Controller
Enriched
Seawater
Power
Data
What pieces inside?
Build or Buy?
Where to put the pieces?
How are pieces connected?
CO2
H2O
Flow
Flow
CTD
pH
CTD
pH
DO
DO
What the software developer sees…
…What the user needs to consider

## Slide 10
Some things must be on the seafloor
  Chamber, instruments
  Actuators (pumps, fans, etc.)
But location of everything else is flexible…
  Power, ESW
  Controller
  Data storage, user interfaces
And, there are many ways to implement all of it…

## Slide 11
Our goal is to narrow the choices
Choices depend on experiment and constraints
  Budget, schedule
  Environment
  Technical resources
xFOCE community may draw from different approaches

## Slide 12
A study of some approaches
  Grand Centralized Workstation
  Industrial Computers
  Sensor Node
Zoom in, evaluate key features and trades
  Build/buy what and why: hardware, software
  Location: shore, surface, seafloor
  Connections: topology and infrastructure
Next stop: Grand Centralized Workstation…

## Slide 13
Controller
Power
Data
Grand Centralized Workstation
Location View
Data
Flow
Flow
CTD
pH
CTD
pH
DO
DO
Component | Location | Why
Controller HW | Seafloor | Experiment continues if comms fails
Controller Software | Seafloor | ‘’
UI Software | Shore | Easier maintenance
Power | Seafloor/Shore | Need fine-grained power management at site
Data | Seafloor/Shore | Experiment continues if comms fails
ESW | Seafloor | Depth, distance from shore

## Slide 14
Controller
Power
Data
Flow
Flow
CTD
pH
CTD
pH
DO
DO
Grand Centralized Workstation
Data
Component
Build/Buy View
Component | Build/Buy | Implementation
Controller HW | Buy | PC104 16 serial ports
Controller Software | Build Open Source | Linux/Java SIAM Open Source Dataturbine
Archive, UI | Build | PC Workstation MacOS, Windows, Linux, Java SSDS Custom FOCE GUI
Power | Build | Custom chassis
Data | Build | SIAM SSDS Open Source
ESW | Build | Custom electronics, Software, chassis COTS pumps, controllers

## Slide 15
Flow
Flow
CTD
pH
CTD
pH
DO
DO
Controller
Power
Data
Connections View
Data
1000 MBps
Ethernet
RS232
10 kW Cable
Grand Centralized Workstation
Ethernet-Serial
Ethernet

## Slide 16: Summary Grand Centralized Workstation
Benefits
  Long term, deep operation required
  High bandwith comms supports video
  Large experimental chamber
  One software platform
  Software reuse
  COTS computers, IO solutions
Costs
  Expensive to build, maintain
  Power consumption
  Specialized for cabled observatory
    Ship, ROV service
  Custom HW/SW
  Custom Ti housings
  Large physical size
Bottom Line for xFOCE
  Central computing w/ workstation computing environment
  Much custom hardware and software
  Expensive and technical to build and use

## Slide 17
Next up: Industrial Computers…

## Slide 18: Summary Industrial Computer Approach
Benefits
  Mostly COTS
  Easier to build
  Potentially long cable runs
Costs
  High cost
  High power consumption
  Cable may be very technical, especially from surface expression in active environment
The Bottom Line
  Expensive, more straightforward design pattern
  min custom HW/SW
  Best suited to cabled system

## Slide 19: Summary Sensor Node Approach
Why consider this
  Reduce power*
  Avoid custom serial IO expansion
  Modularizes data acquisition
  Many variations possible
    COTS processors, languages
    Topologies
    *Connection backbone (cost: power)
    Component upgrades
    Drivers on Gateway
The Rub
  Two software development platforms
  More custom hardware to build, configure, maintain
  More cables, connectors, housings
  Video expansion path unclear
  Multiple system clocks
The Bottom Line
  It’s a good design pattern, with balance of build/buy
  flexible, extensible, moderate cost
  Variants trade complexity, power, modularity

## Slide 20: Approach Scorecard
Approach | Performance | Flexibility | Ease |  |  | Cost
 |  |  | Build | Use | Extend | 
Deep Cabled |  |  |  |  |  | 
WAGO Lego |  |  |  |  |  | 
Sensor Node |  |  |  | … |  | 

I want that!
Not so much

## Slide 21: Other technologies and approaches
Options | Why not here
CAN bus | 
Inductive cables | 
OASIS, BRS | 
Pachube | 
MatLab | 
Acoustic comms | 

## Slide 23: Computing Architectures Considered Approaches
Location Matters  Build vs Buy  | Connected | Semi-Connected | Stand-alone
Custom |  |  | 
Consumer COTS |  |  | 
High grade COTS |  |  | 
xFOCE
CP FOCE
SW FOCE
eFOCE
Deep FOCE
Sensor Node
IPC
SIAM+Digi
SensorPod
?
Slide order? later?
What are Sensor Node etc?

## Slide 24: The Options
What to look for:
mixed
custom
cots/os
Build vs Buy:
by subsystem, color
Location Matters:
More stuff @...
Shore 
Surface 
Seafloor 
Remember to breathe :o)
10 kW
1000 MBps
Eth, Serial
Star
Topology, Resources:
Power, Bandwidth
Backbone
Topology

## Slide 25: Computing Architectures Deep Cabled Approach
Description
  1 year experiment
  1000m Operation
  High Power, Bandwidth cable
Features
  Central computing w/ many serial ports
  Custom electronics & housings
  Custom/open software
  Ethernet, serial
  In situ ESW gen
  ROV CO2 handling
  Video
Examples
  Deep FOCE
ESW
mixer
CO2
FOCE Control
Flow
ESW dist
Power
Dist
video
instruments
mixed
custom
cots/os
SIAM
PC104
Sci/Control
SSDS
RDV
OSDT
Browser
(video)
SIAM
10 kW
1000 MBps
Eth, Serial
Star
terminal
Mac/Win/Lin
DFOCE GUI
Power
Shore

## Slide 26: Computing Architectures WAGO Lego Approach
Description
  1 year experiment
  40-? m Operation
  Med Power, Bandwidth cable
  Not stand-alone (power)
Features
  Custom electronics & housings
  COTS/custom software
  Ethernet, serial, modbus
  On shore ESW gen
  Video
Examples
  Deep FOCE
ESW
mixer
CO2
FOCE Control
flow
ESW dist
Power
Dist
video
instruments
mixed
custom
cots/os
WAGO
Serial MUX
Sci/Control
Browser
(video)
LabView
0.2 kW
2-10 MBps
Eth, RS422
Star/Bus
terminal
National
Inst
Mac/Win/Lin
Serial MUX
WAGO
LV GUI
GF ISO
video
MUX
Shore
Power

## Slide 27: Computing Architectures Sensor Node Approach
Description
  0.5-1 year experiment
  40 m Operation
  stand-alone (shown) or cable
Features
  Custom electronics & housings
  COTS/custom software
  Mixed topology
  Ethernet, serial, modbus
  Off shore ESW gen
Examples
  eFOCe
ESW
mixer
CO2
FOCE Control
Flow
ESW dist
Power
instruments
mixed
custom
cots/os
Beagle e.g. Gateway
Power Dist
Sci/Control
Browser
2 kW
0.1-2 MBps
Eth/DSL?
RS422
Star
terminal
Mac/Win/Lin
Serial MUX
GUI
GF ISO
Diesel Gen
Electronics
FOCE Control
Beagle
Android
Telem
Custom SW
Custom SW
Cell
RF
WiFi
FOCE Control
Sensor Node
Custom SW
Arduino, e.g.

## Slide 28: What was tough for us (but may be easier for you)
Understanding range or users’ resources and capabilities
Understanding user requirements
Likeliest application configurations, constraints
  Stand- alone vs connected
What will users re-use, re-invent (and why)
Trading processors, programming languages
Exporting custom hardware