| Topic | Note | Presentation | Source |
| --- | --- | --- | --- |
| pH flushing pH offsets vs currents | Consider reservoir for equilibration? Yes, but doesn't equilibrate until mixed w/ ambient water. Even with reservoir, needs time to equilibrate else must flush entire system if using exactly desired pH. | kirkwood | paytan |
| xFOCE max flow requirement | Flow speed fixed? It's variable. We chose 10 cm/s 10 is low for shallow sys. big recirculator needed for 50 (50 would be better), mostly just an issue of CO2 use. | kirkwood |  |
| Flow meter tech choice | Are flow meters mechanical? Elec/Mag - no moving parts, no flow obstruction |  |  |
| pCO2 measurement | How difficult to add pCO2? They are slow, in situ ones SAMI 1 sample/15 min - better than discrete sample diving Looked at Quantros - new things coming out trying to min diving 1/2 wks At heron, dove daily, now working on automated sampling |  |  |
| pCO2 measurement (labor) | diurnal changes (to pCO2? TCO2?) if there's a lot of activity 70% work over 8 mo was processesing discrete samples. |  |  |
| Experiment Design | Number of chambers: swFOCE 3 experiment /3 control xFOCE 1 experiment/1 control |  |  |
| Discrete water sampling | want to characterize carbonate chemistry in control, inside treatment and ambient hard prob to deal w/ in terms of people required |  |  |
| Discrete water sampling | how to manage backflow - must sample reference far enough away from chambers |  |  |
| Discrete water sampling | could use osmo sample that collects slowly, and then sample daily, weekly, etc. could sample background conveniently from intakes to ESW, but can't get chamber, control samples that way |  |  |
|  | night sampling is important - mylar bags w/ fixative, solenoids. Can add to FOCE, but not core |  |  |
| Chamber maintenance | cleaning needed daily (chamber and instruments) - At Heron, 20 min if done daily, longer if you let it go. Also, scrubbing chamber makes biofouling worse. |  |  |
| xFOCE community resources | would be good to have place to go to have someone build one for groups w/ minimal engr experience can do meetings/workshops with other xFOCE users (engrs, sci users w. experience) prior to your own project - take advantage of user community | headley | paytan kline |
| Software architecture | to mitigate mult software platforms, make SN software more like firmware - parameterized - so that it doesn't need to be touched often. |  | roman |
| GF detection (cost/benefit) | adding GF detection/isolation adds a lot of function and robustness w/o adding much complexity. |  | mellinger |
| Electronics form factor | are there packaging standards for small processors? Lots of bolted-together pieces may be sort of kludgy. |  | mellinger |
| Custom hardware (user role) | what custom hardware does the user have to implement - what is target for how much the user would implement - to what degree do they need design custom HW? (idea is to make designs and doc that users can use to make basic foce w/o going under the hood) |  | au |
| GF detection (electrical safety) | how do you get around not useing GF on isolated (moored system) for safety? May mean you have to add custom HW/SW to make safe and reliable... cpFOCE used 12V system, so no human danger |  | au |
| Chamber maintenance | - Acid bath easier for cleaning - Don't scrub the chamber - creates more nucleation sites for biofouling | shane | kline |
| Flow effectors | impeller bilge pumps (aquarium type) clogged w/ sand, didn't work well paytan - longer RV types may be better than other cartridge types... |  | kline paytan |
| Diesel as CO2 source | CO2 scrubbing - filters remove particulates, leaving CO2 and Nitrogen(?). how much is it necessary to clean it - how pure the CO2 needed? what about trace organics messing up experiment? |  |  |
| CO2 use | 1xD size (CO2) tank lasted 1 month on Heron e.g. antarctic has diesel, but shipping CO2 expensive 20 cylinders/day, depending on scale, size of chamber, etc |  | kline |
| xFOCE community resources | It would be good to have in the xFOCE doc package a decision tree relating chamber size, CO2, pump size and other parameters |  | kline |
| System architecture (reliability) | fundamental requirement is that failure of one device shouldn't bring down whole experiment. with xFOCE, there are minimal instruments, so losing one thing means you have to stop experiment and fix it. e.g. if GF in one instrument, turn off all three | kecy | chaffey kirkwood ... |
| System architecture (reliability) | if you've been running for a long time and know relation between T and pH, would want to keep going, even w/o pH e.g., then assume variability post-experiment. but you can't close the loop and control pH it is valuable to know which sensor is bad... that's why we have spares - replace it w/ good unit. What is field replaceable unit size? can diver go out and do the swap? paytan - but would not want to stop CO2 flow and keep running. For long term study, things will fail... can you put it in steady state, and keep going. don't eliminate option to have high reliability possibility even if you know GF to a sensor, you still can't tell whether its the sensor, connector, or what Consider implementing fallback/failsafe/limp mode... |  | paytan kirkwood ... |
| System architecture (reliability, granularity) | argues for having expansion room at Gateway level to Different instruments have diff cal periods - so may be disruptive to re cal MTBF, service intervals are different, if nodes are low cost, then can increase numbers - e.g. use one node for one instrument...but GW has to be able to expand failure is different from scheduled maint in failure, must bring back both sides. |  | chaffey ... |
| Power consumption | what is power for wago? ~50-70W |  |  |
| Power consumption | power for (cartridge bilge) pumps overwhelms computer (NI C-Rio) power at heron |  | kline |
| Communications backbone (DSL on Copper) | did we consider using DSL for swFOCE? runs over copper, long distances, big savings in connectors fiber is less user friendly. high speed and robust. more info in appendices. |  | chaffey? |
| Video | how do you leave in option for video? have bandwidth - video on each chamber? no reqmt |  |  |
| ESW Delivery | power budget for pumps, circulators? most power was to accelerate what pumps for ESW inject? Peristaltic 25W, gravity does flow change w/ tide if using gravity? can modulate valve opening time on shore. need 2 atm head CO2 on ESW gen |  | ... kirkwood paytan |
| Engineering data streams | is it valuable to monitor bus voltages, gfs, etc. we add lots of alarms, that at then ignored...is this helpful? don't usually monitor these, but maybe would be useful(?) | Salamy | ...? paytan |
| swFOCE use model | what is model for science use? at first Jim, up to scientists to arbitrate use between COS, HMS, etc. how long will it run, maintenance handoff, primary sponsors HMS, MBARI - may be only users for a while. once it's up and running (a lot of effort), often good to get parallel data from many users. MBARI and HMS primary users, would be great to involve collaborators to extent possible. |  |  |
| swFOCE use model | could, e.g. pass water to another chamber for another experiment. |  |  |
| Data management | need a data integration system w/ single timeline to know how observations link together - to be able to see measurements on a common timeline - including water samples, failures, cals, etc. e.g. on GUI to be able to put them together, CANON/DSS style it would streamline data analysis. the more integrated it is, the faster it is to get useful information out of the system. |  |  |
| Data management | record keeping, daily reports (automated): system may want to log data faster than you have summary reports for. How fast is fast enough ondiff parameters - community recommendations for pH? CTD?... |  |  |
| Discrete water sampling | should sample (water) with divers first to see if there is a signal? how fast are things changing? sample hourly at first to establish diurnal patterns? background signal could change worth doing manual water samples initially, in chamber, things may change more slowly. how easy is it to take phys sample - mesh bags, carbiner 8x50ml took about 10 min. 5 syrnge....can put port in side Carbonate chem analysis (e.g. DIC) takes a lot of time (very labor intensive, even with automation; around the clock at Heron) |  |  |

[Post-Review Inputs]

<ed p about=pCO2, measurement resolution>

You [Adina] are, of course, correct that two CO2 parameters are needed to fully

qualify the CO2 system.  We will add that to the x-FOCE requirements.

After you left I spoke with a few folks at MBARI about this and a couple

of ideas came to mind:

1. The Contros pCO2 sensor is a good candidate for real-time measurements

especially of the background water in situ.  And there is the new Aanderaa

sensor coming-out soon; someone should have a look at that and see if it

is adequate.  We have had bad luck with the SAMI sensors in terms of

producing usable data; plus their need for reagents limits both the length

of deployment and frequency of measurement.  One of these sensors

(whichever folks choose) per site should suffice.

2. For sw-FOCE at HMS, I think we should consider adding a small CO2

equilibrator to the inlet line so that we can measure pCO2 in real-time.

I built one of these with Catherine Goyet many years ago and ran it at

sea.  By cycling between different sources, one can measure pCO2 of the

surface seawater, xCO2 of the air (both inside and outside the pump

house), and reference gases to keep things calibrated.  Combine this with

an inline pH sensor and we can have the CO2 system on the ESW supply line

pretty well nailed down.

3. Adding a Contros (or other sensor) to the experimental chambers will

get expensive if we do everyone, but having it in one might be useful.  We

will need to talk about this more with the team.

On another topic, I've attached some screen grabs from our CO2 release

trials to show how fast conditions can change inside the flume and from

this you can infer how often we need to sample.  In each figure, the top

trace is the six pH sensors, the bottom is the CTD temp outside (blue) and

inside (red) the flume.  The 174...bmp figure shows temperature drifting

along slowly and then a very abrupt 0.1 deg C shift in about 5 mins.  The

pH_w_Temp.bmp shows the temperature cycle over a 5 day period and one can

clearly see the tidal influence on temperature cycling and a daily range

of 0.5 deg C -- which is quite a lot at 890m depth.  In both figures, the

CTD sampling is once per minute (0.01666... Hz).  For our SeaBird CTDs,

this is the fastest rate they go in the non-profiling mode.  WE could go

slower but I think then we will lose some of the finer detail.  I think it

is better to log all this data and then box-car average in post processing

to suit the individual investigators needs.

</ed p>

<tom o about=instrument swapping, maintenance cycles, power>

I think a few reviewer comments highlight the importance of instrument installation and swapping procedures:

- Users want to continue acquiring data even if an instrument fails; i.e. don't want to shut down the entire experiment if a single instrument fails or must otherwise be replaced. Among other implications, this suggests that instrument "hot swap" capability is useful, either at the sensor pod or instrument level.

- The maintenance frequency differs from one instrument to another. If the sensor pod is the basic replacement unit, this means that instruments must be grouped on pods such that all instruments on a pod should be replaced at the same time. In this scenario, some pods might have fewer than 3 instruments - how many serial ports are then required on the gateway node?

I think we should sketch out the instrument/pod replacement process in some detail, and how frequently replacement occurs.

Regarding the need for separate tool chains for gateway and sensor pod software development, Brent Roman suggests that the pod firmware be kept generic with no hard-coded instrument-specific logic. This suggests either running the drivers on the gateway, downloading drivers to pods at runtime, or running "universal" parameterized drivers on the pods.

Davey reiterated that power consumed by the pump overwhelms the other power draws in the system, so it seems that designing an ultra-low power sensor pod may be overkill.

</tom o>

<mark c about=USB for serial expansion>

Something I thought of after I left was that the Gateway to node serial architecture seems a little old school. Rather than build custom hardware, can one use COTS USB to serial converters and do the fan out with USB? That way if you had low power loads only, like sensors, they may actually run from USB 2.0 voltage sources. For high power loads you of course would have your power distribution system. The Gateway would have a multitude of USB ports.

All small electronic devices have USB ports and you could possibly leverage from the power negotiation and device discovery capabilities of the standard. You could also plug the nodes into any portable device for testing or?

You might be able to get much hardware COTS and if not, you would be developing and sourcing something that is progessing forward and would be of broader interest (USB instrument interface node).

Just a thought. Ed (Mellinger) thought it would be too power hungry.

http://www.smxrtos.com/articles/usb_art/waysusb.htm

</mark c>

<tom o about=meeting notes>

Adina: could be useful to continue measurements with surviving instruments even if one instrument dies. E.g. use historical correlation data to continue CO2 injection in "stand by" mode if pH sensor fails. This has implications for electrical design; individual instrument isolation or not?

Chaffey: would be very useful to know exactly which sensor failed.

Davey: if sensor node is basic replacement unit, have to carefully think about how to group instruments based on their calibration/maintenance cycles

If some nodes are partially populated for the above reason, have to ensure that gateway has sufficient serial capability for many nodes.

Davey: pump power overwhelms everything else.

Chaffey: Coletti, McGill have been using DSL very effectively

Roman: Can mitigate multiple-toolchain issue (gateway + sensorpod) by keeping sensorpod firmware generic, so that all sensorpods have the same firmware. Put instrument-specific logic (drivers) on gateway, or make a "generic" driver  for sensorpod that can be parameterized at runtime

</tom o>