- OTG IOIO Board

  - 16 inputs 0-3.3V  3mV resolution

  - Output 3.3V, 5V

  - Digital I/O 3.3V logic 46 pins

- Control of Light source

Mother board:

- Input 12-15 V = > Vin power to OTG IOIO board

- GPIO indicators: LED surface mount, output to scope

Tom tried: analog input, UART

8/5/2015

- Power control of mother board/ other components

  - Headphone signal out – frequency recognition. Low power high reliability control, can have several channels.

  - Microphone?

  - Camera flash LED control – higher power control option

- Removable battery for power control of phone/ reset

  - Keep the battery inside

  - Get rid of battery all together – connect 4V regulated power to the battery compartment

  - Is phone going to try charging non existing battery?

- Camera control

  - Focus

    - Auto control – need focusing feature

    - Pre-focus and disable

    - Software control

  - Image capture

Optimization

- Operation mode(s)

  - Automated, unattended operation

  - Scan on external trigger (start < 30 sec required, < 10 sec desirable)

  - Detection, alarm, warning

  - Identification – desirable

  - Throughput – desirable

- Operation in the filed

  - Surface, low pressure with option to optimize for high pressure

  - Terrestrial

  - Independent to orientation

- Cost – low cost system

  - Initial parts cost <3k

  - Maintenance once per 6 month or better

- Longevity

  - Biofouling

  - Maintenance

- Operating mode

  - External trigger

Design

- Flow path

  - Pressure compensated, insulated flow path (preferred)

  - Insulated flow path

  - Prime fast / slow scan

    - Pump: 2 modes, slow and fast

  - Freeze flow. Is leakage important?

    - Closure valve

  - Combined pump and valve, as in case of gear pump?

- Optical cell

  - Optical performance

  - Cleaning

- Optical path

  - Split screen with time delay

    - Off-centerline issue for optics

- Illumination

  - Several angles – for visibility

    - Forward illumination

    - Back illumination

  - Fluorescence channel

- Detector

- System command-and-control

- Maintenance

  - Assess to flow cell => cleaning

  - Assess to flow cell => repair

- Housing – split design

  - Option 1 - fixed split with “cuvette insert”

    - Issue – fixed geometry

    - How to channel light

    - How to clean – not good assess.

  - Option 2 – separate housings

    - How to connect housings

      - Flex cable

      - Some sort of “sealed pipe” connection

- Detector housing

  - 2 cell phones

  - Optical path – fold

  - Electric board – 1 or 2

  - At least one phone – screen accessible. Can 2 be accessible?

  - Cell phone mounting – replaceable for other cell phones

  - Sealed optical path