## Slide 1: Ultra-fast fluorescence imaging of phytoplankton for species identification and monitoring
Robert Chan
HKBU
SKLMP

## Slide 2: Team membership
Phytoplankton Species Identification by fluorescence imaging (Psi- Fi)
  HKBU members: JX Fu and R.K.Y. Chan
  SKLMP members: Paul Lam, Leo Chan, Maggie Mak, Vicky Wu, Jianglai Wu and Yin Hang
  Woods Hole members: Don Anderson, Michael Brosnahan, Mindy Richien, Heidi Sosik
  OUC member: Xuzhu Wang
  Nanjing member: Yu Hongxia, Zhang Xiaowei
  Industrial members:


## Slide 3
Micro phytoplankton
Imaging phytoplankton in the field
  Sizes ranges from submicron to 100’s micron
Length 64-115 m
www.marbot.gu.se
www.hkredtide.org

## Slide 4
Lake Tai

## Slide 5
Phytoplankton detection
Government Agency (HKAFCD)
  Water samples
  Bulk fluorescence measurement
  Lab tests
    Microscopic inspection and counting
    Spectral measurements
    Up to days to get results
Flow cytometry
  Particles are squeezed in a narrow streamline flow
  Counting is usually done with a laser crossing the flow

## Slide 6
Flow cytometer
Sample nozzle
Laser
Filter
Photodetector
Use Laser-induced fluorescence (LIF) to identify particles

## Slide 7
Current technologies
Light scattering
  High throughput (1000’s particles per sec)
  Simple optical design
  Inaccurate size/mass measurements
  Inexpensive set-up
  OUC and HKBU are developing low cost systems
Imaging (IFCB Woods Hole)
  Complex optical microscopy and cytometry
  Good size/mass discrimination
  Best method for species identifications
  Very slow throughput
  Expensive equipment

## Slide 8
Imaging FlowCytobot
Heidi Sosik with Imaging FlowCytobot ready for testing. credit Woods Hole
Throughput 0.25 ml/min

## Slide 9: Ultra fast flow 2D fluorescence imaging with light sheet microscope
(a) Side view
(b) Top view
Objective10X/0.2
Cylindrical lens
Flow capillary
To camera
Objective
W 40X/0.75
Major breakthrough in imaging cytometry - Image quality not dependent on speed of flow ( no motion blur)
High-speed flow of ~1 m/s with volume flow rate 1 ml/min
Particle counts 10 particles / ml detected in 1 min.
Wu J.L and Chan R.K.Y.; “A fast fluorescence imaging flow cytometer for phytoplankton analysis”; October 2013, Vol. 21, No. 20, Optics Express

## Slide 10: Field Psi-Fi instrumentation
Use high power laser 100 mW CW laser
  Increase flow rate by 5 time to 5 ml/min
Use 300 µm square tubes
  Increase flow rate to 10 ml/min
WHOI needs to sample 100 ml
  Acquisition time is 10 min. for Psi-Fi
  Acquisition time is over 6 hours for IFCB
Lake Tai phytoplankton density is much larger and particles are also larger in size

## Slide 11: 2D imaging
Image acquisition speed 1 s/frame
Large size range discrimination
1 – 200 µm

## Slide 12: PSF deconvolution image processing (Tikhonov-Miller)

## Slide 13: Image deconvolution

## Slide 24: A small selection of species in one water sample

## Slide 25: Results of experiment
Clear images of whole internal structure
  No motion blur
  No depth of focus problems
Good dark background
  High S/N
Distinct patterns of chlorophyll organelle
  Species specific for particle identification?
  Taxonomy
  Life cycle studies
  Parasitic activities

## Slide 26: 1: Vertical water column phytoplankton identification by light sheet flow tube microscopy
Light sheet imaging
Up to 10 suction heads at 1 m apart
Water holding tanks
On boat or raft
Schematic diagram showing the water sampling system and fast light sheet imaging flow cytometer

## Slide 27
Long Term Buoy stations
WiFi communication to base station
Standalone imaging system

## Slide 28: Technical problems/challenges
Design of flow-tube
  Enclosed chamber (HKBU, IndP)
Design of watertight chamber (WHOI, SKLMP, IndP)
  Air conditioning and humidity control (SKLMP, IndP)
  Water inlet and outlet (WHOI, HKBU)
  Electronic interface (WHOI, SKLMP, IndP)
  Power cables (IndP)
Design of fluidics for water sampling and flow tube (SKLMP, OUC, IndP)
Biofouling issues
  Cleansing agent sodium azide (WHOI,SKLMP)
  Maintenance of water sampler
Computing
  Feedback control (OUC, IndP)
  Data compression and transmission (OUC, IndP)
    IFCB has 30 MB/ml, Psi-Fi >30 MB/ml or >5 MB/s
  Image recognition (SKLMP, OUC)

## Slide 29
2: 3D-sectioning by light sheet flow tube microscopy using microfluidic particle trapping technology
High-speed flow of 1 mm/s for 3D sectioning
Volume flow rate 0.5 µl/min
Lateral point spread function (PSF) 0.81±0.07 µm
Axial PSF 1.42±0.15 µm
5.39±0.13µm
Flow sheath
Wu et al; “A light sheet based high throughput 3D-imaging flow cytometer for phytoplankton analysis”; Optics Express, Vol. 21, p. 14474-80, June 2013

## Slide 30
High speed imaging for 3D sectioning

## Slide 31: High speed imaging for 3D sectioning

## Slide 32
3D sectioning of Procentrum sp. and Gambierdiscus sp.

## Slide 33
SEM images
Gambierdiscus

## Slide 34
Experimental results
High speed camera at 1333 frames per sec
Flow rate 1 mm/s
Sectioning width is 0.75 µm
Not depth of field (DOF) problems
Speed of image acquisition typically around 30 ms per particle
  Compared to confocal microscopes which can take up to minutes
2D frame can capture more than one particle at a time
  At least 115X faster than the latest spinning disk confocal microscope
  Ideal for dense chained particles such as Lake Taihu
Axial resolution 2 µm
}

## Slide 35
2: 3D-sectioning by light sheet flow tube microscopy using microfluidic particle trapping technology
Schematic diagram depicting the microfluidic system for phytoplankton particle capture
Photodetector
Detritus
Phytoplankton
Laser sheet
To 3D light sheet imaging flow tube
Non-uniform Electric field
DEP force
Dielectrophoresis

## Slide 36
Conveyors belt particle analysis
Look at the biology of the particles
  Particle trapping
  3D particle analysis using 1
  Microfluidic plays the vital part in the light sheet flow cytometry
  Conveyor belt particle analysis
    Spectral imaging (Raman HKBU)
    DNA Lab-in-a-chip (Nanjing University)
    Mass spectroscopy (Xiamen University)

## Slide 37: Thank you