Advantages of autonomous sampling in situ with mobile robots

Long duration untethered autonomous vehicles overcome many limitations of research vessels and stationary platforms. These vehicles can sample broad areas and volumes and have already revolutionized ocean science [Bellingham, 2011]. We expect that cytometers integrated with these robotic vehicles would greatly advance marine microbiology. Unmanned Surface Vehicles (USVs) or autonomous surface vehicles (ASV) are robot boats that move along the sea surface, and may be equipped with profilers for acquiring data at depth. Autonomous underwater vehicles (AUVs) move in three dimensions and may be propeller driven, or buoyancy driven in the case of gliders. These vehicles provide number of benefits compare to manned vessels.

Automated / advanced surveillance: autonomous vehicles can be dispatched to a specific location, and with underwater vehicles, to specific depths, enabling 2D and 3D surveys. Search can be commanded by operator, another vehicle or mooring, to regions of interest identified by other methods; or vehicle itself may alter search patterns in response to in situ sensor data.

Cost/scalability: operating costs vary, but as a point of comparison,

USV “Wave Glider” (Liquid Robotics of Sunnyvale, CA) costs MBARI about $900/day to run; while 41 meter RV Rachel Carson costs about $8,000/day. Due to lower cost of equipment and operation, they are relatively easy scalable assets: multiple vehicles can be deployed easier than multiple research vessels.

Endurance/ persistent presence:

some AUVs/ USVs are capable of weeks deployments and even months for gliders, spanning thousands of kilometers, they enable continuous presence in regions of interest, provide opportunities to observe microbial community changes over periods of many days and weeks, and enable quick response to events such as blooms.

Combination of lower cost, scalability, and endurance provides persistent presence of a platform with data gathering sensors and improves special and temporary resolution.

Access to extreme environments. Rapid environmental and microbial change is occurring underneath polar ice, yet direct access by crewed vehicles or divers is difficult. The recent discovery of major phytoplankton blooms beneath Arctic ice in the Chukchi Sea [REF] points to our current limited knowledge of these important ecosystems.  AUVs have already performed CTD and ice-thickness surveys underneath Arctic and Antarctic sea ice, and a cytometer payload could enable under-ice microbial surveys [REF].

Cunningham et al (2003) describe a flow cytometer integrated with the AutoSub AUV, which measured changes in microbial populations across a water mass front in the North Atlantic in 2001. AutoSub was equipped with a CTD, spectrophotometer, and water sampler in addition to the cytometer. The sampling area was determined before AUV deployment by SeaWIFS satellite ocean color images that indicated a coccolithophore bloom south of the Isles of Scilly. The AUV was programmed to run a 10 km transect from south to north at 9 meters depth across the bloom front. Particle size and fluorescent particle number data from the cytometer were largely consistent with other onboard instruments as well as post-cruise microscopy of the triggered water samples and analysis of ship-collected data at the north and south ends of the transect. While these results illustrate the potential for AUV cytometry, we know of no other examples of AUV-cytometer integration in 12 years since the AutoSub experiment. There are known integrations with stationary buoys or cable-to-shore systems (IFCB references), but we must consider the technical challenges of integration and operation of cytometers autonomous vehicles.

Power: autonomous vehicles are often designed to minimize non-propulsive power ("hotel load") in order to maximize endurance. Many of today’s cytometers use on order of 50 watts when sampling, and some (e.g. CytoBuoy, IFCB) must sample for several minutes at least. These power levels are higher than what is provided by most small, long-duration AUVs (Table 1).

Size: Even “small” marine cytometers such as SeaFlow are roughly the size of a large microwave oven, and exceed the payload capacity of many AUVs (Table 1)

Telemetry: While cytometer data can be stored onboard for later recovery, in many cases it’s beneficial to transmit at least decimated data in (near) real time, e.g. so that other assets may respond to a phytoplankton bloom, or a manned vessel can sample an area of interest in more detail. AUVs often use 3G/4G cellular modem (0.44 Mbps) for telemetry in coastal areas, but must resort to Iridium satellite (2400 bps) in remote or offshore areas.

Environmental conditions. The instrument should be operable within depth and temperature range of the target organisms. For a phytoplankton, this depth range from 0 to about 200 meters, and temperatures from -1.8C in the photic zone to about 38C in the tropicss. The instrument should be rugged enough to withstand mechanical shock and vibration during typical deployment and recovery cycle, as well as being operable to varying tilt and pitch during deployment.

Risk of vehicle loss: today’s cytometers generally cost on order of $100K and more; if the vehicle is lost, then so is an expensive instrument. High cytometer cost also discourages procurement of multiple instruments, e.g. for deployment on a fleet of vehicles.

While these constraints are daunting, we note that increasingly capable non-cytometer instruments are being integrated with autonomous vehicles. MBARI's Environmental Sample Processor (ESP) is an autonomous ecogenomic “laboratory in a can”, which processes water samples in situ with heat and reagents to extract and analyze microbial genetic material, and transmits results to shore (Scholin, 2013). The third generation instrument is under development  with key requirement to be for compatible with low power mobile platforms in general, and MBARI’s Long Range Autonomous Underwater Vehicle (LRAUV) in particular. The device is contained in a cylinder 28 cm in diameter by 61 cm length, draws between 4 and 20 watts while operating, and is rated to fifty meters depth. The prototype was recently integrated with an LRAUV and underwent initial sea trials in Monterey Bay. The ESP relies on small low-power electronics, microfluidic technologies, and stringent power-management strategies.  Based on the ESP and other examples we believe that development of small low-power cytometers for autonomous vehicles is likely in the near future. The ESP's microfluidic technology and sophisticated chemical and physical sample handling procedures could even point the way to onboard cell staining and sorting required for many cytometry applications.