Cytometers set sail with sea-going mobile robots

AUTHORS

Peter Lopez

NYU School of Medicine

Thomas C. O'Reilly

Denis Klimov

Monterey Bay Aquarium Research Institute

Marine microbial communities play a critical role in Earth's biosphere. Phytoplankton produce approximately 50% of breathable oxygen, regulate atmospheric CO2, and form the base of the marine food web. Toxic algal blooms have major effects on coastal ecosystems and fisheries as well as on human health. Instruments known as flow cytometers (Fig. 1) produce quantitative multi-parametric measurements of individual particles at rates of thousands per second or more (Shapiro 2003), providing knowledge of phytoplankton distribution, diversity and abundance in the ocean. Measurements of cytometric parameters such as light scattering and fluorescence can reveal cell size and intrinsic pigments, which are useful in characterizing plankton taxonomy, abundance and viability (Fig. 2). Most cytometric measurements are made from ships, buoys, and cable-to-shore observatories. Integration of cytometers into autonomous surface and underwater vehicles could greatly advance marine microbiology by enabling better spatial and temporal sampling, facilitating faster "event response", providing access to climate-sensitive polar under-ice regions, and reducing costs relative to ship operations. In this article we discuss instrument and vehicle development trends that point the way to these mobile robot cytometers.

FIGURE 1

Generic diagram of a flow cytometer. Particles for analysis are introduced into a nozzle or cuvette for analysis, where hydrodynamic focusing produces a single-file flow of particles past a focused laser. Light scattering and fluorescence measurements are made by photomultiplier tubes (PMTs) detecting selected wavelengths as determined by optical bandpass filters. PMT output is converted by analog to digital converters (ADCs) for computer display of particle intensity levels. Image used with permission, Semrock.com

FIGURE 2

Flow cytometric analysis of plankton.  Levels of 692nm chlorophyll fluorescence plotted against forward light scattering intensity allow discrimination of various groups, Prochlorococcus (Pro), Synechococcus (Syn), and eukaryotes (Euk). Sample is spiked with 0.75micrometer yellow green “YG” beads (Polysciences, Inc.) as a size reference. Image courtesy of Amy Zimmerman, MBARI

Long-duration untethered autonomous vehicles complement other data gathering platforms such as research vessels and moorings since vehicles can sample broad areas and volumes, and have proven themselves revolutionary in the ocean sciences (Bellingham, 2014). The inclusion of cytometers on these vehicles would greatly advance marine microbiology, but this integration presents challenges. Existing vehicles have limited size and power budget that must be considered. Traditional flow cytometer designs offer unique challenges for autonomous vehicle integration. The typical hydrodynamic focusing utilized for optimal positioning of particles in the flow stream requires a source of particle-free “sheath” fluid. Laser light sources, focusing optics and detection systems must maintain precise alignment. In addition many desirable flow cytometric assays such as total DNA or viability assessment require processing of material before entering the flow cytometer for analysis (Veldhuis and Kraay, 2000, Lomas et al, 2011). Data processing and communication capabilities add another constraint to cytometer integration. Clearly integration of flow cytometers as AUV payloads would benefit from optimized cytometer design.

A consideration of the minimum number of flow cytometric parameters that can provide reliable characterization of plankton is important for producing a focused instrument design. The fact that many species of plankton can be identified by cytometric measurement of the intrinsic qualities of the organism, such as light scattering (approximate size) or fluorescence from native pigments, helps to reduce the parameters needed (Sosik et al, 2010). In addition, novel approaches to the traditional flow cytometer design can further minimize the size/power footprint of an instrument destined for a vehicle. Instruments that simplify sample handling by measuring intrinsic phytoplankton fluorescence without the need for staining cells beforehand or sorting them in the cytometer are the best candidates for early integration.

Parameters that have proven beneficial for marine application of cytometers include Coulter volume, light scattering and fluorescence, time-of-flight pulse shape characterization, and imaging techniques. Some instruments have been developed and deployed that combine some of these techniques; the CytoBuoy (CytoBuoy b.v., Woerden, the Netherlands) for pulse shape and image capture, and the Imaging FlowCytobot (WHOI, Woods Hole, MA) adds image analysis. The wide range of sizes (0.2-200 micrometers) and shapes (round, rod-like, chains) of marine phytoplankton adds a daunting challenge in designing one cytometric system that can accurately identify all types of plankton. Traditional laboratory based cytometers have not been able to analyze such a large dynamic range of particle sizes in one instrument without adding complexity to the system (Beckman Coulter Astrios EQ), or designing systems specific to a portion of the dynamic range (Apogee A50).

<<H1>>STATE OF THE ART

<<H2>>Autonomous vehicle benefits and constraints

Autonomous surface vehicles (ASVs) are robot boats that move along the sea surface, and may be equipped with profilers for acquiring data at depth. Autonomous underwater vehicles (AUVs) move in three dimensions and may be propeller driven, or buoyancy driven in the case of gliders. Autonomous vehicles provide a number of benefits compare to manned vessels:

Automated surveys: Autonomous vehicles can be dispatched to a specific location, and underwater vehicles, to specific depths, enabling 2D and 3D surveys. The survey may be commanded by a human operator or another machine, or the vehicle itself may alter survey patterns in response to in situ sensor data.

Operational cost/scalability: As a point of comparison, a Wave Glider ASV (Liquid Robotics, Sunnyvale, CA) costs MBARI about $900/day to operate; while MBARI's 41-meter RV Rachel Carson costs about $8,000/day. Due to lower cost of equipment and operation, multiple autonomous vehicles running multiple missions can be used at lower cost than one research vessel.

Endurance: Some autonomous vehicles are capable of long deployments lasting weeks or even months, spanning thousands of kilometers; for example the AutoSub LR (Furlong et al, 2012), the LRAUV (Hobson et al, 2012), and the Spray glider (Davis et al, 2002). These vehicles enable continuous presence in a region of interest, provide opportunities to observe microbial community changes over time, and enable quick response to events such as blooms.

Persistent presence: The combination of lower cost, scalability, and endurance can greatly enhance data gathering in spatial and temporal domains.

Access to extreme environments: Rapid environmental and microbial change is occurring underneath polar ice, yet direct access by manned vehicles or divers is difficult. Recent discovery of major phytoplankton blooms beneath Arctic ice in the Chukchi Sea (Arrigo et al, 2012) points to our current limited knowledge of these important ecosystems. AUVs have already performed CTD and ice-thickness surveys underneath Arctic and Antarctic sea ice (Wadhams et al, 2006, Williams et al, 2014), and a cytometer payload could enable under-ice microbial surveys.

These benefits could be utilized for many mobile cytometer applications, including the following projects evaluated at the Monterey Bay Aquarium Research Institute (MBARI).

Controlled, Agile, and Novel Observing Network (CANON): This MBARI project aims to monitor microbial populations and their effects on ecosystems in and around Monterey Bay, emphasizing long-term measurements with autonomous platforms, vehicles and instruments (mbari.org, 2015). CANON includes studies of small phytoplankton (Prochlorococcus, Synechococcus, and eukaryotes) in offshore mesoscale gyres. Other studies emphasize the development lifecycle of larger cells and algal blooms in the inner bay, including potentially toxic Pseudo-nitzschia species. A mobile cytometer would complement physical, chemical sensors, samplers, and ecogenomic instruments such as the Environmental Sample Processor (ESP - see below for more details).

Microbial population changes in ice-covered Arctic waters. Li et al (2009) documented microbial population changes in the Arctic Ocean in concert with climate-driven changes, with picoplankton (less than 2 micrometers) replacing larger cells such as diatoms (2 - 20 micrometers) over a period of several years. These changes may be associated with thinning ice thickness and areal extent as well as freshening and stratification of surface water. Potential shifts in phytoplankton and bacteria community structure will have major repercussions on carbon sequestration at depth in the ocean (Forest et al. 2011). Thus far most sampling had been performed in open water, and much more would be learned from under-ice sampling with a mobile cytometer.

But before autonomous mobile cytometers become reality, we must also consider the technical challenges of operating instruments aboard autonomous vehicles:

Power: Energy is a precious commodity on most autonomous vehicles. Some of today’s marine cytometers consume more than 35W when sampling, and must sample for several minutes at least. These power levels could seriously decrease the endurance of small vehicles.

Size: Even compact marine cytometers such as SeaFlow (Swalwell et al, 2011) in its present configuration, are roughly the size of a large microwave oven, and exceed the payload capacity of many vehicles.

Telemetry: Some autonomous vehicles use cell phone networks close to shore but must use low bandwidth, expensive satellite links when far from shore or in remote regions. Limited telemetry links pose challenges for applications that require near real-time data telemetry to shore.

Environment: The vehicle and its cytometer should be operable within the depth and temperature range of the target phytoplankton, i.e. the photic zone; between 0 and about 200 meters, and from -1.8C to about 38C in the tropics. The instrument should be rugged enough to withstand mechanical shock and vibration during a deployment and recovery cycle, and should be operable within expected tilt/pitch angles during a deployment.

Risk of vehicle loss: Today’s cytometers generally cost on order of $100K and more; if the vehicle is lost, then so is an expensive instrument. High cytometer cost also discourages procurement of multiple instruments, e.g. for deployment on a fleet of vehicles.

Table 1 shows characteristic payload volume and power characteristics for just a few autonomous vehicles with endurance of a week or longer; images of the vehicles themselves are shown in Figure 3.

TABLE 1

Some key characteristics of the Liquid Robotics Wave Glider SV3, Scripps' Spray glider, and MBARI's Long Range AUV, as described in publications, marketing materials and personal communication with developers. Actual endurance will depend on energy needs of integrated instruments, including a cytometer.

|  | Wave Glider SV3 | Spray glider | Long Range AUV |
| --- | --- | --- | --- |
| Type | wave-propelled autonomous surface vehicle | buoyancy-driven glider | propeller-buoyancy hybrid |
| Approximate payload dimensions (cm) | 2 x deck boxes: 380x380x127 1 x deck box10x37x17 towed body: 15 dia x 100 len | 11 dia x 15 len | 30 dia x 33 len |
| Max payload weight (kg) | 45 | 3.5 | 10 |
| Hotel load (watts) | 5 - 20 | 1 | 50+ |
| Payload energy | 150W solar panels, 16 MJ storage batteries | 13 MJ (Li primaries) | 18 MJ (Li secondaries) 40 MJ (Li primaries) |
| Typical endurance | months | 88 days, 354 depth cycles | 11 days, 400 km 27 days, 1000 km |
| Reference | liquidr.com, 2015; Brian Kieft (MBARI) | ucsd.edu, 2015; Brent Jones (MBARI) | Brian Kieft (MBARI) |

TABLE 2

Characteristics of some available automated marine cytometers, as described in publications, marketing materials and personal communication with developer (*).

|  | Imaging FlowCytobot | SeaFlow | CytoSub/ CytoBuoy |
| --- | --- | --- | --- |
| Type | Imaging flow cytometer | Benchtop flow cytometer | Imaging line-scan flow cytometer |
| Application | Submersible to 40 m | Shipboard | Submersible to 200 m |
| Detectable cell size | 1-100 m | 0.5-20 m | 0.5-1500 m |
| Power | 35W | 250W (includes PC/monitor) | 50-60W |
| Dimensions | 26D x102L cm cylinder | 68 x 49 x 33 cm (PC not included) | 32D x 50L cm cylinder |
| Unattended endurance | up to 11 months * | 5 days * | TBD |

FIGURE 3

Some long-duration autonomous vehicles; (a) Wave Glider (b) Spray glider (c) Long Range AUV. Photographs used with permission (a) Tom O'Reilly (MBARI, 2014); Paul Coenen (MBARI); (b) Kim Fulton-Bennett (MBARI, 2012), Christopher Wahl, (c) Todd Walsh (MBARI, 2013); Brett Hobson, Brian Kieft (MBARI), Thomas Hoover (Ocean Aero)

(a)

(b)

(c)

While these conditions are daunting, we note that increasingly sophisticated instruments are being integrated with autonomous vehicles. MBARI's Environmental Sample Processor (ESP) is an eco-genomic “laboratory in a can”, processing samples with heat and reagents to extract and analyze microbial genetic material (Scholin, 2013). The instrument is made possible by low-power electronics, microfluidic technologies, and stringent power-management strategies. The 3rd generation 3G-ESP is under development for low power mobile platforms, MBARI's Long Range AUV (LRAUV) in particular. 3G-ESP is contained in a nose section 28 cm diameter by 61 cm long, draws between 4 and 20 watts, and is rated to 50 meters depth. The prototype underwent successful initial sea trials in Monterey Bay in early 2015. Based on the 3G-ESP and other examples we believe that development of small low-power cytometers for autonomous vehicles is possible in the near future.

<<H2>>Marine Cytometers

At this point only one flow cytometer integrated with an AUV has been reported (Cunningham et al, 2003). They describe a flow cytometer integrated with the AutoSub AUV, which measured microbial populations across a bloom front in the North Atlantic. AutoSub was equipped with a basic science package as well as a water sampler and was deployed into area of interest defined by SeaWIFS satellite ocean color imaging. Cytometric data recorded during this brief hours-long deployment were largely consistent with onboard instruments as well as lab analysis of water samples.

Since 2003, several other automated cytometers have been developed for applications aboard ships or stationary platforms such as piers, buoys, and cabled observatories. We describe a few of these below, and compare some of their characteristics in Table 2.

The SeaFlow cytometer is designed to automatically acquire data over several days without human interaction; its primary deployment mode is aboard ships of opportunity where the instrument continually samples water from the ship's intake for weeks at a time (Swalwell et al, 2011). The SeaFlow "virtual core" design eliminates the need for sheath fluid by adding two additional detectors to determine when a cell is optimally positioned for optical measurement, thereby simplifying the fluidics system. The instrument analyzes cells in the size range of 0.5 - 20 micrometers (pico- and nanoplankton), including Prochlorococcus and Synechococcus, and enables extended two-dimensional surveys of near surface water (Figure 4). Palevsky et al (2013) used continuous underway SeaFlow data to help determine phytoplankton's contribution to the carbon dioxide air-sea flux in Alaskan waters.

FIGURE 4

Distribution of (a) Prochlorococcus, (b) Synechococcus and (c) nanoplankton in Hawaiian waters, as measured by the SeaFlow cytometer. Large spatial scale views are on the left, finer scale on the right (used with permission, from Swalwell et al, 2011).

FIGURE 5

Monterey Bay phytoplankton images captured by Imaging FlowCytobot (instrument loaned to MBARI courtesy of Dr. Heidi Sosik, WHOI)

Imaging FlowCytobot (IFCB) is designed to automatically image, quantify and identify cells larger than 5 micrometers (Olson and Sosik, 2007). IFCB typically draws in a 5 ml water sample every twenty minutes, and triggers a flash lamp and camera when scattering or fluorescence photomultiplier tube (PMT) signal exceeds a preset threshold. Image scale is typically about 0.3 micrometers per pixel, allowing many cells to be analyzed to the genus or even species level (Figure 5). IFCB has been deployed for more than six months on piers and cabled observatories. Campbell et al (2010) describe detection of toxic algal blooms (Dinophysis) with IFCB in a Texas estuary; public health officials quickly responded to the detection by suspending shellfish harvesting until the bloom subsided. IFCB draws considerable electrical power (35W continuously) and generates several gigabytes of data per day; thus the instrument is usually deployed at locations with wired connections to shore providing power and data. IFCB was recently deployed on a moored raft in Salt Pond (Nauset, MA), with no wired shore connection. The raft was equipped with solar panels, a bank of batteries and a gasoline generator that automatically recharged the batteries as required. Communication to shore was via radio link to a shore-based Internet connection (Olson, personal communication).

The CytoSense cytometer uses a one-dimensional scanning technique to capture particle morphological as well as optical information.  Forward scattering, side scattering, and up to three fluorescent channels of each particle are sampled at a frequency such that multiple data points are acquired for each channel. Analysis of the resulting channel pulse shapes yields details of particle morphology and internal structure that in turn aids taxonomic classification, sometimes to the species level (Dubelaar and Gerritzen, 2000).  An optional camera can provide 2-dimensional pictures to further aid analysis. The manufacturers claim an exceptionally wide dynamic range of measured particle sizes, from 0.5 m through 800 m. The instrument is commercially available packaged for buoy as well as submerged applications to 200 meters depth. Thyssen et al (2008) describe a two-month phytoplankton time-series acquired every 30 minutes in the Bay of Marseilles by the submersible "CytoSub" version of the instrument. The cytometer was deployed at 2 meters depth at the end of a quay with a cable to provide power and data connection to the instrument. Nitrate, nitrite, phosphate, salinity, temperature and wind conditions were also measured at the site. From the data they observed fluorescence and abundance changes of seven distinct resolved cell populations coincident with the diel cycle in light, temperature and salinity, as well as in response to strong wind events and resulting hydrologic and nutrient level changes.

<<H1>>THE FUTURE

Electronics, lasers, software and fluidics have evolved in the years since the first AutoSub cytometer deployments, resulting in smaller, lower-power instruments that will likely be compatible with mobile marine robots in the near future.

Sosik and Olson are currently modifying their Imaging FlowCytobot for integration with autonomous surface vehicles such as Wave Glider, autonomous JetYak (whoi.edu, 2014), and AutoNaut (AutonautUsv.com, 2015). The instrument will be towed behind the vehicle at a depth of several meters. They anticipate the vehicle will provide sufficient power to allow sampling by the instrument once every 20 minutes. All raw data and images will be stored onboard for post-recovery analysis. Selected images and potentially processed data will be returned to shore over the available telemetry channels in near real-time; 3G/4G cellular modem (0.5-4 Mbps) in coastal areas, Iridium satellite (2400 bps) in more remote regions (Sosik, personal communication).

While the Imaging FlowCytobot will collect invaluable data from near-surface waters, the instrument is currently too large and consumes too much power for a small, long duration AUVs such as LRAUV, capable of surveys throughout the photic zone. New marine cytometers now under development begin to address these AUV constraints. J. Swalwell of the University of Washington is developing an "immersion flow cytometer" that increases signal-to-noise, simplifies flow cell design, and potentially lowers cost (US Provisional Patent Application Serial No. 62/101,036). The instrument utilizes a low-cost primary ball lens with high index of refraction that is partly immersed in the seawater sample. Analogous to oil or water immersion microscopy, the immersion results in an estimated 60% higher numerical aperture than found in SeaFlow and Influx (BD Biosciences, San Jose CA) cytometers. Chromatic aberration in the ball lens causes laser and fluorescent colors to be dispersed behind the lens, similar to the effects of a prism or grating. Filters and detectors for the wavelengths of interest can be optimally positioned behind the lens, resulting in lower contamination by scattered light at each wavelength and a corresponding signal-to-noise increase. Since the primary lens is immersed directly in the sample, the need for small orifices and sample-delivery tubing may be reduced or eliminated. The instrument utilizes the same "virtual core" technology as the SeaFlow cytometer (Swalwell et al, 2011), eliminating the need for sheath fluid. The cheaper lens, improved signal-to-noise, and simplified fluidic design is expected to result in a smaller (shoebox sized), lower power (roughly 15W) and lower cost instrument (Swalwell, personal communication), with a high potential for deployment on an autonomous vehicle.

Wu et al (2013) describe an imaging approach where a plane of 450 nm laser light (“light sheet”) is used to excite cross-sectional fluorescent images, acquired sequentially, as phytoplankton flow through the light sheet. The instrument generates an integrated cross sectional image of each cell, and is capable of flow rates of at least 1 ml/minute, which is four times the flow rate of Imaging FlowCytobot. The image quality is such that cells larger than about 5 micrometers can often be taxonomically identified; Orientation of the cells can vary as they pass through the light sheet, and the authors discuss taxonomic identification approaches that take this into account. They intend to miniaturize their prototype such that it is compatible with small autonomous platforms such as moorings and AUVs (R. Chan, personal communication).

Though it does not qualify as a "traditional" cytometer, holographic microscopy is another technique that automatically measures individual cell optical characteristics. Conventional microscope optical elements are replaced by a laser, pinhole, and CCD detector. The device acquires holograms containing three-dimensional information from particles within the analyzed volume, eliminating the need to physically align the cells in single-file. Cell images can be reconstructed into multiple image planes. Graham and Smith (2010) describe a submersible holographic microscope with 7.4 micrometer spatial resolution; a commercial version of the instrument has been integrated with MBARI's Dorado AUV and is being evaluated for analysis of large phytoplankton and zooplankton (sequoiasci.com, 2015).  Bochdansky et al (2013) describe a submersible instrument that acquires several holograms per second; a commercial version of this instrument claims sub-micron spatial resolution (4deep.com, 2015); its physical dimensions of 30 cm length and 10 cm diameter make it a potential candidate for integration with a small autonomous vehicle.

Small low-powered cytometers may someday be capable of the many applications that require cell pre-staining and sorting.  Recent advances in microfluidics and manipulation techniques embodied by small ecogenomic sensors such as the AUV-integrated Environmental Sample Processor may point the way in this regard (Scholin, 2013).

Consumer markets of electronics such as mobile phones and computers are rapidly driving the miniaturization and cost-reduction of batteries, processors, cameras, and other sensors. Small, simple low-cost imaging cytometers have been implemented on mobile phones for blood analysis and other medical applications in the developing world (Mudanyali et al, 2012). Perhaps some of these innovations could be explored for marine applications as well, that might dramatically lower instrument cost.

Once a cytometer is integrated with a vehicle, "adaptive sampling" could be used to conserve power and telemetry bandwidth, turning on the instrument only when there is likely something interesting to sample. Zhang et al (2010) describe water sampling of thin phytoplankton layers by MBARI’s Dorado AUV. The AUV flies a vertical saw-tooth profile through regions that may contain layers, and their algorithm triggers a water sampling bottle when fluorescence and backscattering sensors indicate that the vehicle is moving through the center of a layer. Godin et al (2012) describe an algorithm to track moving patches of phytoplankton. The algorithm adjusts vehicle speed, heading and depth based on the chlorophyll signal from an onboard fluorometer, enabling the vehicle to remain near the centroid of the phytoplankton patch as it drifts with the current. The algorithm could be adapted to power on a cytometer only when the vehicle is in a patch, e.g. to allow the instrument to capture the growth and decline of microbes with time.

Some applications require real-time telemetry of cytometer data to shore, e.g. to issue health warnings in case of a harmful algal bloom. But many cytometers generate gigabytes of raw data each day, and telemetry of this data volume is problematic. In coastal regions the vehicle may use a 3G/4G cell modem, but beyond 100 km from shore an Iridium satellite link (2400 bps, compressed) may be the only option (Kocak and Clark, 2013), although faster satellite links may be available in the near future. These slow links can result in very long transmission times, high modem energy consumption, high airtime costs, and long surface times for autonomous vehicles.  Instead of telemetering the entire dataset, onboard processing could return a brief data summary. In the case of cytometer fluorescence and scattering data, onboard cluster analysis could produce a simple list of detected taxa and their abundance(Aghaeepour et al, 2013). Alternatively Li (1997) describes a “cytometric diversity index" in the form of a few floating point numbers that is calculated based on light scattering and chlorophyll fluorescence, and can be used to estimate the structure of phytoplankton communities in real-time, e.g. to track blooms or biological fronts. In the case of image data, Sosik and Olsen (2007) describe image processing "machine learning" software that classifies individual phytoplankton cells to the genus or species level, based on size, shape, symmetry, texture and other characteristics. Their software currently runs on shore-side machines, but they hope to adapt it to the smaller processors on autonomous vehicles.

<<H1>>Acknowledgements

The authors wish to thank Brent Jones, Brian Kieft, Yanwu Zhang, Amy Zimmerman, and Thom Maughan (MBARI); Jim Bellingham, Heidi Sosik, and Rob Olson (WHOI), Jarred Swalwell (University of Washington); George Dubelaar and Dominic Stuart (CytoBuoy b.v.); Alex Cunningham (University of Strathclyde),  Henry Ruhl (National Oceanography Centre), Jeffrey T Sherman (UCSD), Harry Nelson (Fluid Imaging Technologies), Stephen Jones and John Samson (4Deep) for personal communication regarding autonomous vehicles, instrument  specifications, and latest research.

Corresponding Authors

Peter Lopez

NYU School of Medicine

540 First Ave, SK2 Admin

New York, NY 10016 USA
Email: Peter.Lopez@med.nyu.edu

Thomas C. O'Reilly

Monterey Bay Aquarium Research Institute

7700 Sandholdt Road

Moss Landing, CA 95039 USA

Email: oreilly@mbari.org

Denis Klimov
EE, R&D Engineering
Monterey Bay Aquarium Research Institute (MBARI)
7700 Sandholdt Rd
Moss Landing, CA 95039
Email: klimov@mbari.org

References

4Deep.com. 2015. Submersible microscope. Available at: http://4-deep.com/submersible-microscope/ (accessed 24 February 2015).

Aghaeepour N., Finak G., The FlowCAP Consortium, The DREAM Consortium, Hoos H., Mosmann T.R., Brinkman R., Gottardo R. & Scheuermann R.H. 2013. Critical assessment of automated flow cytometry data analysis techniques. Nature Methods 10, 228–238

Arrigo, K. R., Perovich, D. K., Pickart, R. S., Brown, Z. W., van Dijken, G. L., Lowry, K. E., ... & Swift, J. H. 2012. Massive phytoplankton blooms under Arctic sea ice. Science, 336(6087), 1408-1408.

AutoNautUsv.com. 2015. Available at: http://www.autonautusv.com (accessed 24 February 2015).

Bellingham, J. G. 2014. Have robot, will travel. Methods in Oceanography, 10, 5-20.

Bochdansky, A. B., Jericho, M. H., & Herndl, G. J. 2013. Development and deployment of a point source digital inline holographic microscope for the study of plankton and particles to a depth of 6000 m. Limnology and Oceanography: Methods, 11(1), 28-40.

Campbell, L., R.J. Olson, H.M. Sosik, A. Abraham, D.W. Henrichs, C.J. Hyatt, and E.J. Buskey. 2010 First harmful Dinophysis (Dinophyceae, Dinophysiales) bloom in the U.S. is revealed by automated imaging flow cytometry. J. Phycol.,46, 66 – 75.

Cunningham, A., McKee, D., Craig, S., Tarran, G., Widdicombe, C. 2003. Fine-scale variability in phytoplankton community structure and inherent optical properties measured from an autonomous underwater vehicle. Journal of Marine Systems, 43 (1–2) (2003), pp. 51–59

Davis, R. E., Eriksen, C. C., & Jones, C. P. 2002. Autonomous buoyancy-driven underwater gliders. The Technology and Applications of Autonomous Underwater Vehicles, 37-58.

Dubelaar, G. B. J., & Gerritzen, P. L. 2000. CytoBuoy: a step forward towards using flow cytometry in operational oceanography. Scientia Marina, 64(2), 255-265.

Forest, A., Tremblay, J.-É., Gratton, Y., Martin, J., Gagnon, J., Darnis, G., Sampei, M., Fortier, L., Ardyna, M., Gosselin, M., Hattori, H., Nguyen, D., Maranger, R., Vaqué, R., Marrasé, C., Pedrós-Alió, C., Sallon, A., Michel, C., Kellogg, C., Deming, J., Shadwick, E., Thomas, H., Link, H., Archambault, P., Piepenburg, D. (2010) Biogenic carbon flows through the planktonic food web of the Amundsen Gulf (Arctic Ocean): A synthesis of field measurements and inverse modeling analyses. Progress in Oceanography. 91(4):410-436

Furlong, M. E., Paxton, D., Stevenson, P., Pebody, M., McPhail, S. D., & Perrett, J. 2012, September. Autosub Long Range: A long range deep diving AUV for ocean monitoring. In Autonomous Underwater Vehicles (AUV), 2012 IEEE/OES (pp. 1-7). IEEE.

Godin, M. A., Ryan, J. P., Zhang, Y., & Bellingham, J. G. 2012, December. Localization and Tracking of Submerged Phytoplankton Bloom Patches by an Autonomous Underwater Vehicle. In AGU Fall Meeting Abstracts (Vol. 1, p. 1869).

Graham, G. W., & Nimmo Smith, W. A. M. 2010. The application of holography to the analysis of size and settling velocity of suspended cohesive sediments. Limnology and Oceanography: Methods, 8(1), 1-15.

Hobson, B. W., Bellingham, J. G., Kieft, B., McEwen, R., Godin, M., & Zhang, Y. 2012, September. Tethys-class long range AUVs-extending the endurance of propeller-driven cruising AUVs from days to weeks. In Autonomous Underwater Vehicles (AUV), 2012 IEEE/OES (pp. 1-8). IEEE.

Kocak, D. M., & Clark, A. M. 2013. Motivations and Methods for Modern Maritime Communications: A Survey of Recent and Emerging Trends. Marine Technology Society Journal, 47(5), 7-18.

Li W.K. 1997. Cytometric diversity in marine ultraphytoplankton. Limnol Oceanogr 42:874–80.

Li, W.K., McLaughlin, F. A., Lovejoy, C., & Carmack, E. C. 2009. Smallest algae thrive as the Arctic Ocean freshens. Science, 326(5952), 539-539.

liquidr.com. 2015. Specification sheets. Available at: http://liquidr.com/resdown/downloads/spec-sheets.html (accessed 26 February 2015)

Lomas, M. W., Bronk, D. A., & van den Engh, G. 2011. Use of flow cytometry to measure biogeochemical rates and processes in the ocean. Annual review of marine science, 3, 537-566.

MBARI.org. 2015. CANON: Controlled, agile, and novel observing network. Available at http://www.mbari.org/canon/ (accessed 25 February 2015)

Mudanyali, O., Dimitrov S., Sikora U,Padmanabhan S.,Navruza I., and Ozcan A. 2012. Integrated rapid-diagnostic-test reader platform on a cellphone Lab Chip, 2012,12, 2678–2686

Olson, R. J., & Sosik, H. M. 2007. A submersible imaging‐in‐flow instrument to analyze nano‐and microplankton: Imaging FlowCytobot. Limnology and Oceanography: Methods, 5(6), 195-203.

Palevsky, H. I., Ribalet, F., Swalwell, J. E., Cosca, C. E., Cokelet, E. D., Feely, R. A., & Quay, P. D. 2013. The influence of net community production and phytoplankton community structure on CO2 uptake in the Gulf of Alaska. Global Biogeochemical Cycles, 27(3), 664-676.

Scholin CA. 2013. Ecogenomic sensors. In: Levin SA, (ed.) Encyclopedia of Biodiversity 2nd edn Volume 2. Academic Press: Waltham, MA, 690–700.

Semrock.com, 2015. http://www.semrock.com/flow-cytometry.aspx

sequoiasci.com. 2015. Testing Sequoia's LISST-Holo on an autonomous underwater vehicle. Available at: http://www.sequoiasci.com/article/testing-sequoias-lisst-%C2%ADholo-on-an-autonomous-underwater-vehicle/ (accessed 24 February 2015).

Shapiro, H. M. 2003. Practical flow cytometry. John Wiley & Sons. 736 pp.

Sosik, H. M., & Olson, R. J. 2007. Automated taxonomic classification of phytoplankton sampled with imaging‐in‐flow cytometry. Limnology and Oceanography: Methods, 5(6), 204-216.

Sosik, H.M., Olson, R.J., & Armbrust, E.V. 2010. Flow cytometry in phytoplankton research. Chlorophyll a Fluorescence in Aquatic Sciences: Methods and Applications. pp.171–185

Swalwell, J. E., Ribalet, F., & Armbrust, E. 2011. SeaFlow: A novel underway flow-cytometer for continuous observations of phytoplankton in the ocean. Limnology and Oceanography: Methods, 9(10), 466-477.

Thyssen, M., Mathieu, D., Garcia, N., & Denis, M. 2008. Short-term variation of phytoplankton assemblages in Mediterranean coastal waters recorded with an automated submerged flow cytometer. Journal of Plankton Research, 30(9), 1027-1040.

UCSD.edu. 2015. Underwater gliders. Available at: http://spray.ucsd.edu/pub/rel/info/spray_description.php  (accessed 26 February 2015).

Veldhuis, M.J.W., & Kraay. G.W. 2000. Application of flow cytometry in marine phytoplankton

research: current applications and future perspectives. Scientia Marina 64(2): 121-134

Wadhams P, Wilkinson JP, McPhail SD. 2006. A new view of the underside of Arctic sea ice. Geophysical Research Letters. 33:L04501. doi: 10.1029/2005GL025131

Whoi.edu. 2014. The JetYak. Available at https://vimeo.com/113751116 (accessed 25 February 2015).

Williams, G.D., Wilkinson, J., Maksym, T., Kunz, C. & Singh, H. 2014. Very thick and heavily deformed Antarctic sea ice captured in 3-D by autonomous underwater vehicle, International Symposium on Sea Ice in a Changing Environment - Proceedings of the Hobart Symposium, 10-14 March 2014, Hobart, Tasmania, Australia, pp. 69A834

Wu, J, Chan, RKY. 2013. A fast fluorescence imaging flow cytometer for phytoplankton analysis. Opt Express Oct 7;21(20):23921-6. doi: 10.1364/OE.21.023921

Zhang, Y., McEwen, R.S., Ryan,  J.P., Bellingham,  J.G., Harvey, J. & Vrijenhoek, J. 2010. Acquiring peak samples from phytoplankton thin layers and intermediate nepheloid layers by an autonomous underwater vehicle with adaptive triggering. AGU 2010 Fall Meeting, Abstract OS51C-1334, San Francisco, CA, U.S.A., December 2010.