## Slide 1: Desktop testbed: goals
Gain familiarity with object and methods
  Fluidics
  Optics
Find out what is possible
  Scale performance to deployable prototype
  Design optimization
Test some ideas
  Pressure gradients
  Bulk vs. hydro focusing

## Slide 2: Key criteria
Adaptability to autonomous platform
Science “performance”
Maintenance stability
Actionable data
Size, power
Shock
Tilt
Population diversity
Throughput
Maintenance
Alignment
Biofouling
Trigger sampling
Telemetry
Screening/sorting
Limiting use on many platforms
Population statistics in real time is challenging
Length of deployment
Unattended operation
Configurability
Temporal, spatial resolution

## Slide 3: Benchmark: fluorescent imaging
Filtering 25 ml on 0.2um black polycarbonate filter
Sampling, 125 ml
Fixing with Glutaraldehyde
Visual identification, counting
Translating stage
Change objective: 10x, 40x, 100x
Entering spreadsheet
Notes: Multiple steps, manual, lab equipment. Can we automate?

## Slide 4: Benchmark: fluorescent imaging
Ciliate with sequestered chloroplasts at 100x
Zeiss Axioplan fluorescence microscope
Courtesy of Lisa Ziccarelli
Red: chlorophyll
Different physiological state?

## Slide 5: Benchmark: fluorescent imaging
Pseudo-nitzschia, at 100x
Courtesy of Lisa Ziccarelli
Zeiss Axioplan fluorescence microscope
Chaetoceros, at 100x

## Slide 6: Pigment fluorescence: how you do it?
(1) http://www.oceanopticsbook.info/view/optical_constituents_of_the_ocean/_phytoplankton
(2) http://mhl.nsw.gov.au/projects/berowra/chloro.php
Chlorophyll absorption (1)
Pigments absorption (2)
(3) http://www.photobiology.com/photobiology99/contrib/danuta/
Chlorophyll (auto) fluorescence (3)
Light source + excitation filter
Detector + emission filter
bandpass
long wave pass
Notes: Useful technique – how you do it? Auto fluorescence of native pigment

## Slide 7: Cell phone cytometer?
Courtesy of Ken Johnson
Fluorescent Imaging cytometry on a cell phone
Inexpensive components
Biomedical applications

## Slide 8: Cell phone test stand
X-Y-Z stage
Optical base plate
Cell phone
Sample on a slide
Optical filter film (for fluorescence)
Micro lens

## Slide 9: Cell phone microscope
With red film filter
Cross-talk!!!
Kyocera C5155 2048x1536 1.6um/pixel
HTC Nexus One 2592x1944 1.3um/pixel
1 mm grid
85 um line
6-10 um texture

## Slide 10: Optics: chasing cross-talk
LED Luxeon Royal Blue 455 nm
+
Red filter (Roscolux #19)
Blue is still visible
Red filter OD not enough
LED Luxeon Royal Blue 455 nm
+
Blue filter (Intor 450-40)
+
Red filter (Roscolux #19)
Pink light is visible:
Long tail of blue LED

## Slide 11: Optics: filters
LED Luxeon Royal Blue 455 nm
Roscolux #19 red film
Cross-talk regions
Band-pass filter
Better off-band suppression

## Slide 12: Optics: Filters
Solid plastic
Color film
Interference filters

## Slide 13: Optics: Filters
LED Luxeon Royal Blue 455 nm
+
Intor 450-40 filter
Delvie 2422 plastic filter
Isolation region
Cross-talk removed But filter is thick!

## Slide 14: Testbed
X-Y-Z stage
Base plate
Imaging Source cameras:
BW and Color 1600x1200
Blue LED
White LED
Microscope objective
Lens tube

## Slide 15: Testbed: imaging scale
4x objective
1600x1200 pixel
1600 um wide = > 1 um / pixel
10x objective
1600x1200 pixel
650um wide = > 0.4 um / pixel
100 um optical glass target

## Slide 16: Testbed: BW imaging, slide
4x objective
1600x1200 pixel, 1 um / pixel
RG630 filter
BW camera: all pigments – one color
Fluorescence
“White” light
Samples courtesy of Holly Bowers
160 x 160 um region
160 x 160 um region

## Slide 17: Testbed: color, fixed sample
10x objective
1600x1200 pixel, 0.4 um / pixel
Delvie 2422 filter
Color camera: ID pigment group
Fluorescence
“White” light
Samples courtesy of Lisa Ziccarelly
65 x 65 um region
65 x 65 um region

## Slide 18: Testbed: color, net tow sample
10x objective
1600x1200 pixel, 0.4 um / pixel
Delvie 2422 filter
Color camera: ID pigment group
Fluorescence
White light
Samples courtesy of Julio Harvey
Multiple cells per slide, each in its own bubble!

## Slide 19: Comparison
Includes processing platform
Limited control of camera
Better camera control
Access to image trigger signals
Multi-camera architecture
Use of microscope objectives
Compact, low power
Need processor board
Cell phone obsolescence
Single camera?

## Slide 20: Learned
2 channels:
Synchronous acquisition seems to be valuable
Fluorescence signal is weak
  Longer exposure
  Higher excitation power
  Fluorescence of materials
Depth of field is quite short at high resolution
Cell phone microscope: not bad!
Images of same object
organic vs. inorganic
Flow cell configuration
High dynamic range of detection

## Slide 21: Next steps
Energy: sensitivity, power, exposure
Software: image processing
Flow cell design
Optimization
Architecture
Processing platform

## Slide 22: Software
Simplified imaging processing algorithm
  Particle length and width
  Fluorescent vs non-fluorescent
  Area, integral signal => proxy to biomass
Triggered image logging, other actions
Prototype runs on laptop, should be embedded in the future
Explore hardware acceleration, e.g. with FPGA or GPU

## Slide 23: Example processing chain, implemented in GPU/FPGA
From “A hardware accelerated approach for imaging flow cytometry”, Lee, Meng et al (2013)