## Slide 1: Cytometer technology for autonomous platforms: a feasibility study
Presented for Jupiter Research Foundation
Tom O’Reilly
Denis Klimov
Francisco Chavez
Andy Hamilton
3/10/2015

## Slide 2: Contributors and collaborators
Methods and applications
Instruments and training
Samples and analysis
Heidi Sosik (WHOI)
Rob Johnson (WHOI)
Jarred Swalwell (UW)
Harry Nelson, Kevin Stewart (Fluid Imaging)
Stephen Jones, John Samson (Resolution Optics)
Jim Bellingham
Francisco Chavez
Alex Worden
John Ryan
Magda Gutowska
Steve Haddock
Ken Johnson
Peter Lopez (NYU)
Holly Bowers
Reiko Michisaki
Camille Poirier
Lisa Ziccarelli
Tim Pennington
Julio Harvey
April Woods
Jason Smith
Gernot Friederich

## Slide 3: Key CANON questions about single-celled organisms
What organisms are present?
How are the organisms distributed spatially and temporally?
How do organisms and ecosystems respond to climatic/oceanographic changes?
Population dynamics, harmful algal blooms...

## Slide 4: Project goals and activities
Look at available and emerging cytometers, potentially compatible with AUV, WG
Evaluate instruments in lab, at sea
Develop bench-top test bed
Formulate implementation path
  Science use cases
  Buy-vs-build

## Slide 5: Imaging cytometers
Imaging Flow Cytobot
Add camera to detectors
Image analysis for cell size, shape, internal structure, taxonomy...
Handles morphologically complex cells
Lower size-limit ~2 microns (diffraction limit)

## Slide 6: Imaging Flow Cytobot
Fluorescence or scatter PMT signal triggers camera and flashlamp
1380 x 1034 b/w CCD camera
Max image rate ~10 Hz
Autonomous operation, 6+ month deployments

## Slide 7: Imaging Flow Cytobot: Auto taxonomic classifier
Feature extraction + machine-learning classifier
Extracts size, shape, symmetry, texture...
Open source software

## Slide 8: Light sheet imaging cytometer
Laser
Wu and Chan, 2013
Cells flow through blue laser light sheet, generate chl-a images
Light sheet thickness ~5 microns; ~100 ms exposure per frame
No motion blur
Image deconvolution to remove out-of-focus light

## Slide 9: Light sheet imaging cytometer
1 ml/min flow
0.75 um resolution, identify species > 5 um
Can generate Z-stacks, 3D images
Wu and Chan, 2013
Notes: Light sheet 5.4 um thick
Image acq speed: ~30 ms/particle

## Slide 10: Holographic microscopy
No objective lens, object info captured from large volume, post-acquisition focusing
LISST Holo (courtesy J. Ryan)
Submersible Microscope
Notes: Holo: 4.5 microns/pixel, ~1 minute to acq image.
SM: 16-50 fps, lateral resolution ~1 um, 5W

## Slide 11: Target organism size
In general, more small cells than large cells
10 um
1 mm
10 cm
Abundance (A)
Size (D)
0.1 um
10 m
Courtesy J Bellingham
VN: volumetric rate required to detect N cells per unit time
VN
A∝Db
(-1.3 < b < -0.6)

## Slide 12: Application scale
Different methods have different optimal size range
‘Hybrid’ methods help overcome limitations, expand application scale
1 um
10 um
100 um
1000 um
particle size
0.1 um
Flow cytometry
Imaging
Worden
Scholin, Ryan
Vrijenhoek, Haddock
Chavez

## Slide 13: Some throughput limits
Assume ρ = 10,000 cells/ml, N = 1000 cells (CV = 3%): “Typical” flow cytometer f = 1 ml/hr; must measure for 359 sec
Core flow rate – small cells not aligned/resolvable at high flow rates
‘Dead time’ increases with event rate, can miss rare/large cells
  Reduce detector gain to decrease event rate
Optimize setup for target cells
slow flow
fast flow
sheath
core
laser

## Slide 14: Mobile autonomous cytometer applications
Three-dimensional surveys
Bloom patch tracking
Thin-layer detection/sampling
Identify regions of interest for intensive ship sampling
Draw on 3G ESP, LRAUV-Dorado experience

## Slide 15
Aug
Collaborative surveys
Notes: This is the present MBARI CANON paradigm – integrated use of diverse robotic platforms equipped with chemical and biological sensors, shore based assets, satellite systems, and modeling. For the near term (and perhaps mid term) we anticipate requiring a ship which would carry researchers conducting laboratory analysis on samples collected by the ship and the robotic systems.
CANON uses multiple mobile platforms to track individual aggregations of organisms to better observe how they respond to and interact with their surrounding environment.
New platforms and sensors required. Sensors such as the ISUS nitrate sensor, dissolved oxygen, various types of fluorometers, particle size, and ability to identify individual microbial species (e.g. using ESP) are necessary, along with water sampling for later laboratory analysis. Deployment time requirements are set by the duration of biological events - for example a coastal bloom might last about a month. This drives a need for a platform that can last about a month (much longer than Dorado) and carry a more complex sensor suite than a glider, motivating the Tethys Long Range AUV (which is now in testing).

## Slide 16: Autonomous vehicle challenges
Power, size
Telemetry
Thermal, orientation, shock
Sampling – flow rate, moving vs stationary, ‘gulp’ vs continuous

## Slide 17: Existing autonomous mobile cytometer platforms
Wave Glider (in progress)
  Jupiter Research Foundation mobile microscope
  WG Imaging Flow Cytobot (WHOI, NSF)
AUV
  CytoSub prototype on AutoSub (2002)
JRF microscope payload

## Slide 18: Instrument comparisons
Imaging Flow Cytobot - flow cytometer with camera (O’Reilly)
FlowCam – laminar flow cell with chl PMT and camera (Hamilton)
LISST Holo – holographic microscope (Ryan)
Submersible Microscope – holographic microscope (Klimov)
LISST-100X – bulk scattering (Ryan)
Lab fluorescent microscope - Zeiss Axioplan2 (Ziccarelli, Michisaki)
BD Influx – “Gold standard” lab cytometer (Poirier)
SeaFlow cytometer (Swalwell)
Notes: Goals: gain experience with flow cytometry, compare instrument capabilities and nuances.

## Slide 19: Instrument comparisons
Particle size distributions
Particle concentrations
Images
Sampling methods
Tests
  Culture samples in the lab
  At sea on Rachel Carson, pumped rosette sampling

## Slide 20: Image comparisons
LISST Holo: 4.4 micron/pixel
IFCB: 0.32 micron/pixel
Flowcam: 0.55 micron/pixel

## Slide 21: IFCB and FlowCam
IFCB
FlowCam
FlowCam
IFCB

## Slide 22: IFCB and FlowCam
IFCB
FlowCam
FlowCam
IFCB
WG microscope

## Slide 23: Particle size distributions
IFCB, FlowCam measure equivalent circular diameter (ECD)
LISST-100X derives particle dimensions from ‘bulk’ forward scatter
Reinecker et al, 2008

## Slide 24: Particle size distributions
Also measured size with Zeiss Axioplan-2

## Slide 25: Particle size distribution
Isochrysis
LISST100X
Zeiss

## Slide 26: Particle size differences
IFCB Nannochloropsis
3.1 pixel/micron
FlowCam, IFCB: particle boundary, size (ECD) determined by image processing

## Slide 27: Particle size differences: calibration
LISST-100X – no calibration method provided
IFCB and FlowCam – calibrate with beads
IFCB: 9 micron bead, 3.5 pixels/micron
Carefully calibrated IFCB
  Note fuzzy bead edge
FlowCam – no reliable bead data collected

## Slide 28: Paragon sample
IFCB

## Slide 29: Particle size distribution
Paragon
IFCB PMT gain set for larger cells
LISST100X peaks for each shape axis
IFCB 150 um Nitex mesh
FlowCam 100 um flow cell

## Slide 30: Evaluation: more to do...
Evaluate particle concentration results
Compare IFCB, FlowCam, SeaFlow pumped rosette profiles to InFlux lab cytometer (with J. Smith, Swalwell, Poirier)
Revisit Submersible Microscope -Try latest software release

## Slide 31: Lessons from evaluations
Calibration against common standard
Instrument settings influence which cells are analyzed
Software settings influence results (e.g. where is the cell “edge”?)
Sampling: continuous vs discrete
Notes: We note ‘quirks’, pros and cons of each instrument – but this slide shows some general observations.
Statistics and sample draw are especially relevant for moving platform. IFCB not designed for fine scale spatial resolution (e.g. on profiler)

## Slide 32: Klimov: Bench top test-bed

## Slide 33: Desktop testbed: goals
Gain familiarity with object and methods
  Fluidics
  Optics
Find out what is possible
  Scale performance to deployable prototype
  Design optimization
Test some ideas
  Pressure gradients
  Bulk vs. hydro focusing

## Slide 34: Key criteria
Adaptability to autonomous platform
Science “performance”
Maintenance stability
Actionable data
Size, power
Shock
Tilt
Population diversity
Throughput
Maintenance
Alignment
Biofouling
Trigger sampling
Telemetry
Screening/sorting
Limiting use on many platforms
Population statistics in real time is challenging
Length of deployment
Unattended operation
Configurability
Temporal, spatial resolution

## Slide 35: Benchmark: fluorescent imaging
Ciliate with sequestered chloroplasts at 100x
Zeiss Axioplan fluorescence microscope
Courtesy of Lisa Ziccarelli
Red: chlorophyll
Different physiological state?

## Slide 36: Benchmark: fluorescent imaging
Filtering 25 ml on 0.2um black polycarbonate filter
Sampling, 125 ml
Fixing with Glutaraldehyde
Visual identification, counting
Translating stage
Change objective: 10x, 40x, 100x
Entering spreadsheet
Courtesy of Lisa Ziccarelli
Notes: Multiple steps, manual, lab equipment. Can we automate?

## Slide 37: Benchmark: fluorescent imaging
Pseudo-nitzschia, at 100x
Courtesy of Lisa Ziccarelli
Zeiss Axioplan fluorescence microscope
Chaetoceros, at 100x

## Slide 38: Benchmark: fluorescent imaging
Traditional method, time series
Laboratory equipment, manual
Hours per slide
Can we emulate?
  Simplified version
  Autonomous, unattended operation
Testing with slides
How you do it?
Notes: Multiple steps, manual, lab equipment. Can we automate?

## Slide 39: Pigment fluorescence: how you do it?
(1) http://www.oceanopticsbook.info/view/optical_constituents_of_the_ocean/_phytoplankton
(2) http://mhl.nsw.gov.au/projects/berowra/chloro.php
Chlorophyll absorption (1)
Pigments absorption (2)
(3) http://www.photobiology.com/photobiology99/contrib/danuta/
Chlorophyll (auto) fluorescence (3)
Light source + excitation filter
Detector + emission filter
bandpass
long wave pass
Notes: Useful technique – how you do it? Auto fluorescence of native pigment

## Slide 40: Cell phone cytometer?
Courtesy of Ken Johnson
Fluorescent Imaging cytometry on a cell phone
Inexpensive components
Biomedical applications

## Slide 41: Cell phone test stand
X-Y-Z stage
Optical base plate
Cell phone
Sample on a slide
Optical filter film (for fluorescence)
Micro lens

## Slide 42: Cell phone microscope
With red film filter Cross-talk!!!
Kyocera C5155 2048x1536 1.6um/pixel
HTC Nexus One 2592x1944 1.3um/pixel
1 mm grid
85 um line
6-10 um texture
Stained pseudo -nitzschia

## Slide 43: Optics: chasing cross-talk
LED Luxeon Royal Blue 455 nm
+
Red filter (Roscolux #19)
Blue is still visible
Red filter OD not enough
LED Luxeon Royal Blue 455 nm
+
Blue filter (Intor 450-40)
+
Red filter (Roscolux #19)
Pink light is visible:
Long tail of blue LED

## Slide 44: Optics: Filters
Solid plastic
Color film
Interference filters

## Slide 45: Optics: Filters
LED Luxeon Royal Blue 455 nm
+
Intor 450-40 filter
Delvie 2422 plastic filter
Isolation region Cross-talk removed But filter is thick!
Intensity, counts
Transmission, %

## Slide 46: Optics: chasing cross-talk
LED Luxeon Royal Blue 455 nm
+
Blue filter (Intor 450-40)
+
Red filter (Roscolux #19)
Yellow light slightly visible – may upgrade blue filter
LED Luxeon Royal Blue 455 nm
+
Blue filter (Intor 450-40)
+
Delvie 2422
Blue is still visible
Red filter OD not enough

## Slide 47: Testbed
X-Y-Z stage
Base plate
Imaging Source cameras:
BW and Color 1600x1200
Blue LED
White LED
Microscope objective
Lens tube

## Slide 48: Testbed: imaging scale
4x objective
1600x1200 pixel
1600 um wide = > 1 um / pixel
10x objective
1600x1200 pixel
650um wide = > 0.4 um / pixel
100 um optical glass target

## Slide 49: Testbed: BW imaging, slide
4x objective
1600x1200 pixel, 1 um / pixel
RG630 filter
BW camera: all pigments – one color
Fluorescence
“White” light
Samples courtesy of Holly Bowers
160 x 160 um region
160 x 160 um region

## Slide 50: Testbed: color, fixed sample
10x objective
1600x1200 pixel, 0.4 um / pixel
Delvie 2422 filter
Color camera: ID pigment group
Fluorescence
“White” light
Samples courtesy of Lisa Ziccarelly
65 x 65 um region
65 x 65 um region
Ceratium sp.

## Slide 51: Testbed: color, net tow sample
10x objective
1600x1200 pixel, 0.4 um / pixel
Delvie 2422 filter
Color camera: ID pigment group
Fluorescence
White light
Samples courtesy of Julio Harvey
Multiple cells per slide, each in its own bubble!

## Slide 52: Comparison
Includes processing platform
Limited control of camera
Better camera control
Access to image trigger signals
Multi-camera architecture
Use of microscope objectives
Compact, low power
Need processor board
Cell phone obsolescence
Single camera?
“Cell phone level”
Entry level
High end
microscope components

## Slide 53: Learned
2 channels:
Synchronous acquisition seems to be valuable
Fluorescence signal is weak
  Longer exposure
  Higher excitation power
  Fluorescence of materials
Depth of field is quite short at high resolution
“Cell phone microscope”: not bad!
Images of same object
organic vs. inorganic
Flow cell configuration
High dynamic range of detection

## Slide 54: Next steps
Energy: sensitivity, power, exposure
Architecture
  Software: image processing
  Flow cell design
  Processing platform
Capability  requirements
  Search for middle ground
  Multidisciplinary team
Have some idea
Running out of time this year
Very important

## Slide 55: Future project directions
Continue investigation into latest tech
Refine test-bed, test against science use cases
Investigate automated cell identification (e.g. as used by IFCB, Flowcam)
Trade-offs; Build vs buy, science vs cost
Continue collaborator engagement (JRF, Sosik, Swalwell, Garren...)
Implementation project in 2016

## Slide 56: Questions?

## Slide 57: Particle size differences: physiologic state
IFCB – round???
FlowCam
Zeiss Axioplan-2
Dunaliela (~10 microns)

## Slide 58: Cytometer (Gk kytos + metron)
Automatically measures individual cell characteristics; size, shape, morphology, taxonomy, physiologic state...
Takes many measurements (e.g. >1000s per second)
Several technical approaches, more kinds of “cytometer” today
Notes: Contrast with ‘bulk’ techniques, e.g. transmissometer, nephelometer...

## Slide 59: Flow Cytometry
laser light
objective
lens
field stop
pmt
Fwd scatter filter
fluor filter
pmt
sample core
sheath fluid
Infer cell type from whole cell scatter, fluorescence
Best for simple cell morphologies, ~0.5 – 10 microns. Complex morphologies more problematic

## Slide 60: Flow cytometry - cytogram
Swalwell et al, 2011
Notes: Cytogram, aka “dot plot” shows correlation between two parameters. Populations are identified through multi-variate training sets which define discriminant functions for different cell types.

## Slide 61: Additional cytometer techniques; sorting
Sort cells based on scatter/fluorescence; purifies rare subpopulations for further analysis
Courtesy of P. Lopez

## Slide 62: Some potential MBARI applications
Application | Target cells | Investigator
Oligotrophic population dynamics | prochlorococcus; 0.5-0.8 um, 10^4-10^5 cell/ml syncechococcus; 0.8-1.5 um, 10^3-10^4 cell/ml eukaryotic ultraphytoplankton; 2-20 um, 10^3-10^4 cell/ml heterotrophic bacteria; <5 um, 10^5-10^6 cells/ml; | Worden
Coastal population dynamics | cells > 2 um (Chavez) ~10^5 phytoplankton cell/ml | Chavez, Smith, Ryan
HAB survey | Pseudo-nitzschia; Width = 2-8um, Length = 40-175um Washington coast 1998 Pseudo-nitzschia bloom, 18000 cell/ml Washington coast 1997 Pseudo-nitzschia bloom, 13000 cell/ml Gulf of Mexico Karenia bloom: 20-40 um, ~700 cell/ml Gulf of Mexico Dinophysis bloom: 20-40 um, ~300 cell/ml Gulf of Maine Alexandrium bloom: 40-50 um, ~5 cell/ml | Scholin

## Slide 63
 | Focal length | Objective mag. | Resolution (micron/pxl)
Imaging Flow Cytobot |  | 10x | 0.32
FlowCam |  |  | 0.55
LISST Holo |  |  | 4.4
Submersible Holo Microscope |  |  |