## Slide 1: Cytometer technology for autonomous platforms: update
Tom O’Reilly and Denis Klimov
Jim Bellingham
Francisco Chavez
Alex Worden

## Slide 2: Agenda
2013 activities
  CANON use cases
  Literature review of existing and emerging technologies
  Visits and discussions
2014 plans
  Prototype development
  Instrument evaluations,
  visits, discussions

## Slide 3: CANON applications
Bulk versus individual particle measurements
Cytometry
  Worden
  Chavez

## Slide 4: Cytometer dynamic range, sensitivity trade-offs
Time to measure N cells, where p is cell number density (i.e. concentration) and f is flow rate through measurement volume:
T = N / (fp)
If N = 1000, coefficient of variation = 3.1% and SNR = 32.

## Slide 5: Count rate, time, energy to measure 1000 cells (σ = 3.2 %, SNR = 32)
 | Imaging Flow Cytobot | SeaLabel
Power | 35 W | 35 W
Flow rate thru measurement volume | 15 ml/hr | 0.9 ml/hr
100 cells/ml | 0.4 pulse/sec 2400 sec 8.4 x 104 Joule | 2.5 x 10-2 pulse/sec 40000 sec 1.4 x 103 Joule
105 cells/ml | 417 pulse/sec 2.4 sec 84 Joule | 2500 pulse/sec 40 sec 1400 Joule

## Slide 6: Single-volume detection trade-offs slide

## Slide 7: Imaging versus single-volume detection
Complex morphology, non-spherical particles, etc...

## Slide 8: Some key instrument characteristics
Size and mass
Power
Deployment duration
Unattended operation
Consumables (sheath fluid)
Data volume, in situ processing requirements
Resistance to shock, orientation change
Target platforms - LRAUV, Dorado, Wave Glider – varying capabilities, duration, power, comms, value to community (graphic images?)

## Slide 9: Existing oceanographic cytometers
SeaFlow
  Discussions with Swalwell; he will give MBARI seminar in 2014
Imaging Flow Cytobot
  Sosik visited in September 2013
CytoBuoy
  Dugdale and Chavez tested CytoBuoy aboard RV Carson in September 2013
  Publication for ASLO 2014

## Slide 10: Microscopes
FlowCam
  We might borrow instrument for CANON 2014
LISST Holo
Submersible microscope
  Hobson, Bellingham coordinating visit, demos

## Slide 11: Existing oceanographic cytometers
Comparison matrix goes here

## Slide 12: Technology overview
Denis will add

## Slide 13: Related MBARI activities
Bellingham lab spin-off
Johnson lab Coulter Counter
Hobson investigation of submersible microscope

## Slide 14: Instrument test-bed plan
Evaluate imaging, counting approaches that could be miniaturized
Hardware and electronics
Software