| Subconn Pin # |  | Colour Code |  | Ethernet (TIA 568B) |  | 1000Base-T Signal ID |  | Power |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  | Black |  |  |  |  |  | Not Used |
| 2 |  | Orange |  |  |  |  |  | Not Used |
| 3 |  | White |  |  |  |  |  | +12V |
| 4 |  | Brown |  | 8 |  | DD- |  |  |
| 5 |  | Brown/White |  | 7 |  | DD+ |  |  |
| 6 |  | Blue |  | 4 |  | DC+ |  |  |
| 7 |  | Blue/White |  | 5 |  | DC- |  |  |
| 8 |  | Orange |  | 2 |  | DA- |  |  |
| 9 |  | Orange/White |  | 1 |  | DA+ |  |  |
| 10 |  | Green |  | 6 |  | DB- |  |  |
| 11 |  | Green/White |  | 3 |  | DB+ |  |  |
| 12 |  | Red |  |  |  |  |  | Not Used |
| 13 |  | Green |  |  |  |  |  | GND |

From: <william@4-deep.com>

Subject: RE: Microscope evaluation: ideas?

Date: March 15, 2016 6:13:16 AM PDT

To: "'Tom O'Reilly'" <oreilly@mbari.org>

Cc: "'John Samson'" <john@4-deep.com>, "'Stephen Jones'" <stephen@4-deep.com>

Hi Tom,

Please see attached file for cable pinout. To interface with the submersible you will need a Subconn 13 pin Ethernet connector and locking ring. (male connector ‘DBH13M’ – ring ‘DLSA-M’) Note that we use Ethernet wiring scheme TIA568B and that +12V is fed to the unit via the white (+12V) and green (GND) 18 AWG wires. Since this connector is not mated to a cable you will need to mount it to a bulkhead. The threads on the connector are ½-20 UNF. There is another Subconn connector type you could use but you will need to mate it to a cable. The part number is: DOM13M

We do have a flow chamber you may use. John has it with the demonstration unit now.

If you have any other questions, please let me know.

Tks

Bill

OPERATIONS MANAGER

William MacLean

www.4-deep.com

T +1 (902) 444 4449

william@4-deep.com

6589 Chebucto Road

Halifax, Nova Scotia

Canada B3L 1L9

From: Tom O'Reilly [mailto:oreilly@mbari.org] 
Sent: March 14, 2016 8:32 PM
To: John Samson <john@4-deep.com>
Cc: Bill Mclean <william@4-deep.com>; Denis Klimov <klimov@mbari.org>
Subject: Re: Microscope evaluation: ideas?

Also, we plan to evaluate the instrument in the lab as well as at sea.

Does 4Deep have a 'flow chamber' that we can use in the lab? (Otherwise we can build our own)

Thanks!

Tom

On Mar 14, 2016, at 3:58 PM, Tom O'Reilly <oreilly@mbari.org> wrote:

Hi Bill

Can you please send specification/drawing/pinout for the cable?

We will likely lower the instrument from a boat to about 50-100 meters depth.

Thanks

Tom

On Mar 14, 2016, at 11:02 AM, John Samson <john@4-deep.com> wrote:

I do not see this as a problem, except that we have only 5 m of cable for the demo sub. If you are to deploy to an ROV, you will want to build your own comm/power cable. I include Bill, our operational manager, to confer with you on this, as he can make suggestions about the cable system for an ROV.

Thanks,

John

<image003.jpg>

Research Scientist

John Samson

www.4-deep.com

T +1 (902) 444 4449

john@4-deep.com

6589 Chebucto Road

Halifax, Nova Scotia

Canada B3L 1L9

On Mon, Mar 14, 2016 at 2:39 PM, Tom O'Reilly <oreilly@mbari.org> wrote:

Thanks for that info John.

We are meeting with MBARI scientists to discuss an evaluation plan - are we allowed to test the loaner instrument at sea, e.g. lowered from a cable or mounted on an ROV?

Thanks,

Tom

3/16/2016 call to Bill (William MacLean) www.4-deep.com   T +1 (902) 444 4449

McCartney

They build cable, they mold Subconn, mount RG45 connector,

4-deep add power

D-p4tp2#4c18

3/16/2016 email from Brock.

Hi Dennis:

Here are two cables that Falmat has in stock:

FMXCAT51205K24 = 164 ft @ $11.84/ft.  5 #12 awg

FMXCAT51806K12 = 164 ft @ $8.35/ft.   6 # 18awg

See attached specifications.

The DOM13M is $150.00.

Regards,

Brock Rosenthal

Ocean Innovations

brock@o-vations.com

(858) 454-4044

www.o-vations.com

3/17/2016

3/17/2016 10:40 call to Brock 858-454-4044

DOM13M is $150.00

DLSA-M $6

Boot mold $13

Fedex from Houston TX 77043 to 95039

Has one in stock in San Diego

$9 ground shipping

4808 0190 0255 210,  05/18,  Denis Klimov, 711