## Slide 1: DeepPIV Subastian Integration

## Slide 2: Project Goal
Goal: to deploy DeepPIV instrument on the ROV Subastian during 2 research cruises (10/19, 12/20)
Tasks: create interface can to operate DeepPIV.

## Slide 3: Electrical Interfaces
Subastian Instrument Interface
  Comms: Fiber
    9/125 um
    Connector: Teledyne Omicron 5070
  Power:
    240 VAC
    Connector: Subconn MCBH3M
  Science Port
    24 V Power, 8A fuse?
    RS232 Comms
DeepPIV Interface
  Laser Housing
    110 VAC, 2A
    Ethernet
  Camera
    24 V, <1A
    RS232 comm ports (2)
  Dye Injector
    24 V, 2A
    RS232 comm port

## Slide 4: System Block Diagram
Subastian ROV
LASER
Camera
Dye Injector
DeepPIV v.3
240 VAC
Interface Can
Fiber
24 V, RS232
240 VAC
Ethernet
24V
RS232 (2)

## Slide 5: System Block Diagram

## Slide 6: Camera Housing

## Slide 7: Laser Housing

## Slide 8: DeepPIV + Integration Can Power Consumption
 | Voltage (V) | Current (A) | Power (W)
Laser Housing |  |  | 
Laser + Controller | 12 | 10.5 | 126
DGH Modules | 12 | 0.17 | 2.04
Digi Serial Server | 12 | 1.5 | 18
temp/humidity | 12 | 0.05 | 0.6
Total Laser Housing Power |  |  | 146.64
Camera Housing |  |  | 0
Camera | 12 | 0.3 | 3.6
Twiga Int Board | 12 | 0.2 | 2.4
Temp/humidity board | 12 | 0.05 | 0.6
Total Camera Housing Power |  |  | 6.6
Dye Injector |  |  | 
Elmo +Motor | 24 | 2 | 48
Total Dye Injector Power |  |  | 48
Interface Can |  |  | 
Focal stack | 5 | 1.5 | 7.5

## Slide 9: Integration Can

## Slide 10: Focal Mux
Focal 914-HDE
  Supports:
    3G-SDI
    Ethernet (100Base-T)
    Serial Comms (2 Channels)
    Expansion interface (CWDM, Multiple Ethernet, Iso serial)
    Single or Dual single mode fibers 9/125 um

## Slide 11: Interface Can AC-DC
Meanwell NED-35B
  Input 85-264 VAC
  Dual Output:
    5V, 2.2 A
    24V, 1 A
    35 W total
  Easily replaceable with other models at higher power if needed
  Mechanical: 99 cm x 97 cm x 36 cm

## Slide 12: Top side
Focal 914-HDE
Console
HD Recorder
3G-SDI
Laptop
Ethernet
RS232 (Camera)
RS232 (Temp/Humid)
NTP Server
Fiber

## Slide 13: Cables

## Slide 14: Integration Can Housing
Former Tiburon motor control housing – not one that imploded
Used at 4000m
OD ~ 6.5”
ID ~ 5.25 “
Length ~ 18”
New Endcap

## Slide 15: Laser and Camera Housings
Weight (Total/Manip): Dry 84.8/25.4kg (187/56 lbs); In-water 20.4/9.1 kg (45/20 lbs)
Dimensions (OD/length): Camera (5.9/13.6 in), Laser Controller (8.4/29.4 in), Remote Laser (2.5/7.5)

## Slide 16: Schedule – lead times
Connectors
  Impulse: have used Omicron 5070, expecting new ones 6/19 - ∞
  Seacon: have enough in house to start (order expected 7/1/19 - ∞)
  Subconn: have in house
Focal mux ~ 4 weeks
AC-DC readily available
Endcap manufacturing: 4-6 weeks

## Slide 17: Schedule
Bench testing of DeepPIV + focal 914-HDE through MBARI ROV early July
Integration and testing of complete can in August (test with MBARI ROV using Lancer??)
Pieces ship to Hawaii late September
Cruise mid-October

## Slide 18: TBD
Budget

## Slide 19: Questions

## Slide 20: Coherent Laser Assmbly
Notes: PURPOSE – X
Text(.8x1.4)

## Slide 21: Laser Remote Head
Notes: Text(.8x.14) (.8x5.4)

## Slide 22: Focal 914-HDE

## Slide 23: Focal 914-HDE

## Slide 24: Alternate Plans