Spray Glider O2 Calibration:
1. Attach air pump to spray flow cell, turn on, and allow 30 minutes-1hr to stabilize.
2. Have the barometer in the same localized area as the air pump, use this to check for stability in relative humidity and temp.
3. When ready to record: ON COMPUTER
a. Press period twice to get out of low power mode (holding period also works)
b. Press “Z” for SBE base menu
c. Hit “C” to take CTD Sample.
d. 3rd line is: DO (mL/L), TEMP, RAW PHASE ---  Record these values!
4. Do this 3x and take an average of the three readings.
5. Press Q to quit the menu and go back to base program 
a. Glider will automatically go back to sleep
6. Record in the “Oxygen_air_calibration_log” in Y:\901805-Coastal Biogeochemical Sensing\Spray pH\Air calibrations
7. Report the Gain to Yui!


[bookmark: _GoBack]Mission ID: YYMXXXYY, where YY is year, M is month in hex, XXX is the spray serial number, and YY is the mission number that month




Glider Commands:
Period twice to get out of low power mode (hold it)
Base menu- Z (SBE)
C- CTD SAMPLE
Q to quit menu and go back to base
3rd line DO, Temp, Raw phase



