Spray Glider O2 Calibration:

- Attach air pump to spray flow cell, turn on, and allow 30 minutes-1hr to stabilize.

- Have the barometer in the same localized area as the air pump, use this to check for stability in relative humidity and temp.

- When ready to record: ON COMPUTER

  - Press period twice to get out of low power mode (holding period also works)

  - Press “Z” for SBE base menu

  - Hit “C” to take CTD Sample.

  - 3rd line is: DO (mL/L), TEMP, RAW PHASE ---  Record these values!

- Do this 3x and take an average of the three readings.

- Press Q to quit the menu and go back to base program

  - Glider will automatically go back to sleep

- Record in the “Oxygen_air_calibration_log” in Y:\901805-Coastal Biogeochemical Sensing\Spray pH\Air calibrations

- Report the Gain to Yui!

Mission ID: YYMXXXYY, where YY is year, M is month in hex, XXX is the spray serial number, and YY is the mission number that month

Glider Commands:

Period twice to get out of low power mode (hold it)

Base menu- Z (SBE)

C- CTD SAMPLE

Q to quit menu and go back to base

3rd line DO, Temp, Raw phase