Mission ID: YYMXXXYY, where YY is year, M is month in hex, XXX is the spray serial number, and YY is the mission number that month

SPRAY GLIDER TEST TANK SAMPLE: (If using Test Menus)

- Place Spray in test tank overnight so that sensors can saturate, and bubbles can be pushed out of the flow cell.

- After ~24 hours in the tank, bring Spray to surface (or just about sub-surface) so that the pumped flow path is able to still bring in water. Spray should be mostly submerged.

- Check O2 and CTD sample and get stable readings: ON COMPUTER in Terminal

  - Press period twice to get out of low power mode (holding period also works)

  - Press “Z” for SBE base menu

  - Press “H” for Hi pump

  - Hit “C” to take CTD Sample.

  - 3rd line is: DO (mL/L), TEMP, RAW PHASE ---  Record these values!

  - CTD data is BRENT

- Check pH sample data, go to test menu and Deploy in POLLED Mode (0s wait time)

  - TS to take sample, do this several times and record the stable reading data as AVG.

  - Data fields: SAMP NUM, Time/Date, Batt volts, Vext Ref, Vext Ref STD, Vcounter, Vcounter STD, Counter Current (Ik), Substrate current (Ib).

- While the data is stable and readings are being taken, a second person will deploy the Niskin for a side-by-side grab sample for Winkler O2 and Spectrophotometric pH.

  - O2 samples taken in duplicate will go first, followed by a Spec pH sample.

    - O2 Samples: Fixed via winkler reagents

    - pH: Either run immediately after thermal equilibration in 20C bath, or add 200ul of mercuric chloride and seal for storage.

- Data taken from the pH measurement and ancillary data (Vext, SAL, Temp) on the sensor will be combined with the spec pH measurement to get a calibration value (k0) utilizing CO2SYS and Matlab.

- K0 values and ancillary tank data will be logged in: YUI

pH:

- Periods to wake Spray into BAS menu

- ‘A’ to get into pH TEST

- ‘W’ to wake pH board

- ‘H’ for Hi speed pump

- ‘D’ for TS

- WHEN DONE: ‘2’ for pump off

- ‘Q’ to quit

- ‘Z’ for sleep

CTD:

- Press period twice to get out of low power mode (holding period also works)

- Press “Z” for SBE base menu

- ‘H’ for hi speed pump

- Hit “C” to take CTD Sample.

- 3rd line is: DO (mL/L), TEMP, RAW PHASE ---  Record these values!

- When done: ‘2’ to turn off pump

- ‘Q’ for base menu

- ‘Z’ to sleep