Lessons from MAR19 deployment

- On first few dives we didn’t parse the pH serial return string properly resulting in bad 20 bin returned sat data

- SBE63 and Optical cable failed after dive 3.

Lessons Learned from FEB19 deployment. Results in RED from MAR19 deployment

- Some Iridium sbd emails not reaching our shore server.

  - Held up in Mimecast.  Issue should be fixed by whitelisting iridium sbd sender email address – Fixed, no issues with missing sbd emails

- SBE 63 raw phase not recording.  Apparently the fix tested on the bench did not work.  BDJ will investigate further.  I installed new SBE63 drivers from SIO that now search for the prompt.  This should fix behavior.  Tested successfully on bench.  Fixed

- pH sensor leaked at ~500m dive. Replaced sensor and it worked fine to 200m

  - Will find out exactly what depth it started leaking

- Slight corrosive pitting discovered on Al anodized adapter.  Will monitor.  Nothing to note

Lessons Learned from DEC18 deployment.  Results in RED from FEB19 deployment

1.Spray ballasted a bit heavy. Will remove some pB

YT: How did you identify this? Is it in some engineering data?

See how long it takes to pump before it becomes buoyant. ~100 seconds to buoyant - > pretty good range.

- Good time to buoyant on this mission

2.CF card not returning last 3 voltages. Bug in code. Fixed, tested.

3.Light intrusion affecting shallow measurements (indicated by high, negative Ib). Machined a black delrin manifold and installed

- Rotate DSD so that ISFET faces the incoming flow stream

- Flew with new black delrin flowcell

4. SBE 63 raw phase offset not recording. This was an insidious one. The "5001" value returned on the CF card is an error code indicating SBE comms were messed up. Turns out, SBE inserted 2 additional non printable characters in the sensor output string. AWESOME! Love it when they do that. Anyway, fixed and tested

5. pH resetting and not starting at bottom of profile. I retooled my query routine to now issue a start deployment ( "2<cr/lf>") as the wakeup from sleep. The idea is, if the pH has been reset, this will immediately start it up. Since the wake up character doesnt matter, it also wakes up a controller that is deployed. So far, in bench testing I have not been able to recreate the resetting behavior. However, I need to test this on the bench more. I plan on doing this next week.

- Have we verified that the pH controller reset at the bottom of every profile?

- Multiple headers in pH file.

- Maybe we can add a column in the file of # of times it has reset? Makes parsing easier

- 2 is the deployment character

- Limited to 200m. New anodized aluminum fitting is being machined.

  - Flew with new anodized adapter to 750m.

7. Salinity was reading ~0.06 low.

8. Want Durafet information in header for glider data. What limitations are there for including this in the .sat file?

#----start calibration coefficients --------------

VN   40  5 2 0810  Press Temp Sal Opt DOx

VO FLUOR

VA  88371  Argos ID

CH00  -14   1  hcal188040.txt

CP12  -10.000   0.040    0.0000   1.0000

CT22   -5.000   0.001    0.0000   1.0000

CS32   -1.000   0.001    0.0000   1.0000

CO44    0.000   0.001  10     3.00   0.0000   1.0000

CX62    0.000   0.001    0.0000   1.0000 0.0000 0.0000

PX161  k0 k2 c1 c2 c3 c4 c5 c6

#----end calibration coefficients ---------------

Can we have some of this info in the .raw file as well?

- Powers on all instruments at the beginning of profile, and then shuts power at the top of the profile.

Goals

- Target deployment date Feb 11. Be at M1 on Feb 14th, recover Feb 18 (or early next week).

- Calibration cast on recovery.

- Profile to 1000 m, go to M1 and back.

Action Items before next deployment

- Loading pH sensor information on mFET controller (EEPROM or text file in SD card)

- Make checklist for pH sensor, include information/calibration coefficients into checklist

- Straighten up state machine (TM)

- Put voltage monitor on power line for mFET

- Put alignment marks for pH sensor

- Ability to turn on/off headers for O2 and SBE37

- Compare Spray salinity to recovery CTD

- Shore station, data visualization -> get that to MBARI.

- Talk to Ken about FloatViz, what it takes for GliderViz

- Brent: Put in cruise plan for Feb 11 and Feb 18.